## HCBLM_I001
Beam Loss Monitor: Ionization Chamber
Identification |  |  |  |  | Step 1 (): Assembly | [6 empty] | Step 1: Non-Conformities |  | Properties |  |  | Step 2 (): Vacuum stand | [6 empty] | Step 2: Non-Conformities |  | Properties | [7 empty] | Step 3 (): Leakage current measurement at IHEP by IHEP | [6 empty] | Step 3: Non-Conformities |  | Properties |  |  | Step 4 (): Leakage current measurement at IHEP by CERN | [6 empty] | Step 4: Non-Conformities |  | Properties |  | Step 5 (): Shipping | [6 empty] | Step 5: Non-Conformities |  | Properties |  |  | Step 6 (): Leakage current measurement at CERN | [6 empty] | Step 6: Non-Conformities |  | Properties |  | Step 7 (): Source test | [6 empty] | Step 7: Non-Conformities |  | Properties | [13 empty] | Step 8 (): Source test at tunnel | [6 empty] | Step 8: Non-Conformities |  | Properties |  |  | Equipment Non-Conformities |  | Additional document 1 | Additional document 2 | Additional document 3 | Plastic Box №
Equipment ID | Other ID | Manufacturer | Status | Comment | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Leakage Current (Assemb) | Head Producer | High Voltage (Assembly) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Batch Number | Gas Bottle Number | Spectrum (before filling) | Temperature (filling) | Pressure (filling) | Temperature (heating) | Duration (heating) | Spectrum (after filling) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Leakage Current_f | Leakage Current_v | High Voltage (Leak-I-I) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Leakage Current (IHEP) | High Voltage (Leak-I-C) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Lorry Number | IHEP Departure Date | CERN Arrival Date | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | High Voltage (Leak-C-C) | Leakage Current (CERN) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | High Voltage1 | Current1 | High Voltage2 | Current2 | High Voltage3 | Current3 | High Voltage4 | Current4 | High Voltage5 | Current5 | High Voltage6 | Current6 | Dose Rate | Source Activity (Surface) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Current | High Voltage (Tunnel) | Source Activity (Tunnel) | Doc. Number | Type | Doc. Number | Doc. Number | Doc. Number
[14 empty] | [pA] | [] | [V] | [9 empty] | [] | [] | [] | [°C] | [mbar] | [°C] | [h] | [] | [9 empty] | [pA] | [pA] | [V] | [9 empty] | [pA] | [V] | [9 empty] | [] | [] | [] | [9 empty] | [V] | [pA] | [9 empty] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [millisievert par heure] | [GBq] | [9 empty] | [pA] | [V] | [GBq]
(CERN ID) | [42 empty] | <1 pA
BL01 | HCBLM_I001 | Beam Loss Monitor: Ionization Chamber |  |  | MTF013476 | [7 empty] | 3 | P11483 | P11446 | P11448 | MTF013477 | [7 empty] | 8 | P11450 | P11468 | P11453 | P11451 | P11452 | P11454 | P11456 | P11455 | MTF013478 | [7 empty] | 3 | P11458 | P11457 | P11459 | MTF013479 | [7 empty] | 2 | P11449 | P11460 | MTF013480 | [7 empty] | 3 | P11461 | P11462 | P11463 | MTF013481 | [7 empty] | 2 | P11464 | P11465 | MTF013482 | [7 empty] | 14 | P11471 | P11477 | P11472 | P11478 | P11473 | P11479 | P11474 | P11480 | P11475 | P11481 | P11476 | P11482 | P11464.7 | P11469 | MTF013483 | [7 empty] | 3 | P11470 | P11466 | P11447
| 0 | 8 |  | 3
HCBLM_I001-05000008 | IHEP0008 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.4 | 0.1 | 2000 | Done | Ok | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.8 | 1800 | Done | Ok | 5 | 2006-09-28 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | -2.3 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54 | [10 empty] | 4.7 | 98 | Pending |  | 8
HCBLM_I001-05000009 | IHEP0009 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.7 | 0.9 | 2000 | Done | Ok | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.7 | 1800 | [12 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.6 | [10 empty] | 4.7 | 98 | Pending |  | 8
HCBLM_I001-05000010 | IHEP0010 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 150 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6
HCBLM_I001-05000011 | IHEP0011 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.2 | 0.9 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 1 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.8 | [10 empty] | 4.7 | 98
HCBLM_I001-05000012 | IHEP0012 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.7 | [10 empty] | 4.7 | 98
HCBLM_I001-05000013 | IHEP0013 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.8 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.8 | [10 empty] | 4.7 | 98
HCBLM_I001-05000014 | IHEP0014 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 2.5 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.48 | 1 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.5 | [10 empty] | 4.7 | 98
HCBLM_I001-05000015 | IHEP0015 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 3 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.1 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.7 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.6 | [10 empty] | 4.7 | 98
HCBLM_I001-05000016 | IHEP0016 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.4 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 11 | 1800 | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 |  |  | 2 | 25.07.2006 |  | Done | Ok | 6 | 2006-06-23 | Ferioli |  | prob in Protv |  |  | 2000 | 40 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.92 | [10 empty] | 4.7 | 98
HCBLM_I001-05000017 | IHEP0017 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.5 | [10 empty] | 4.7 | 98
HCBLM_I001-05000018 | IHEP0018 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 4 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.1 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | -0.4 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.3 | [10 empty] | 4.7 | 98
HCBLM_I001-05000019 | IHEP0019 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 5 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.4 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.9 | [10 empty] | 4.7 | 98
HCBLM_I001-05000020 | IHEP0020 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 7 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Tikhonov |  |  |  |  | 0.24 | 0.2 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.5 | [10 empty] | 4.7 | 98
HCBLM_I001-05000021 | IHEP0021 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.2 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.7 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.2 | [10 empty] | 4.7 | 98
HCBLM_I001-05000022 | IHEP0022 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 2.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 1.4 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 15
HCBLM_I001-05000023 | IHEP0023 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.4 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.4 | [10 empty] | 4.7 | 98
HCBLM_I001-05000024 | IHEP0024 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 |  |  |  | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.6
HCBLM_I001-05000025 | IHEP0025 | IHEP | Manufacturing |  | Done | Not Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 2.5 | IHEP | 2000 | Cancelled | Cancelled | [15 empty] | Cancelled | Cancelled | 3 | [11 empty] | 4 | 2006-06-01 | G.Ferioli | [18 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  | prob in Protv |  |  | 2000 | 0.3 | Done | Ok | 7 | [18 empty] | 4.7 | 98
HCBLM_I001-05000026 | IHEP0026 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 1 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.3 | [10 empty] | 4.7 | 98
HCBLM_I001-05000027 | IHEP0027 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Cancelled | Cancelled | 3 | [11 empty] | 4 | 2006-06-01 | G.Ferioli | [18 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.3
HCBLM_I001-05000028 | IHEP0028 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 4 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.3 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.9 | [10 empty] | 4.7 | 98
HCBLM_I001-05000029 | IHEP0029 | IHEP | Manufacturing |  | Done | Not Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.5 | IHEP | 2000 | Cancelled | Not Ok | [15 empty] | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.73 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.3 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.6 | [10 empty] | 4.7 | 98
HCBLM_I001-05000030 | IHEP0030 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.73 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli | [5 empty] | 1800 | Done | Ok | 5 | 2006-07-25 | Grishin | [21 empty] | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.182
HCBLM_I001-05000031 | IHEP0031 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.7 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.3 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.1 | [10 empty] | 4.7 | 98
HCBLM_I001-05000032 | IHEP0032 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.8 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.2 | [10 empty] | 4.7 | 98
HCBLM_I001-05000033 | IHEP0033 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.3 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.4 | [10 empty] | 4.7 | 98
HCBLM_I001-05000034 | IHEP0034 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.9 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | -1.3 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 55 | [10 empty] | 4.7 | 98
HCBLM_I001-05000035 | IHEP0035 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.2 | [10 empty] | 4.7 | 98
HCBLM_I001-05000036 | IHEP0036 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | -0.3 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.2 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54 | [10 empty] | 4.7 | 98
HCBLM_I001-05000037 | IHEP0037 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-27 | Mishagin |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.73 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 1 | 1800 | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 |  |  | 2 | 25.07.2006 |  | Done | Ok | 6 | 2006-06-23 | Ferioli |  | prob in Protv |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.052 | [10 empty] | 4.7 | 98
HCBLM_I001-05000038 | IHEP0038 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli | [5 empty] | 1800 | Done | Ok | 5 | 2006-07-25 | Grishin | [7 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8
HCBLM_I001-05000039 | IHEP0039 | IHEP | Manufacturing |  | Done | Not Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.4 | IHEP | 2000 | Cancelled | Cancelled | [15 empty] | Cancelled | Cancelled | 3 | [11 empty] | 4 | 2006-06-01 | G.Ferioli
HCBLM_I001-05000040 | IHEP0040 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.3 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54 | [10 empty] | 4.7 | 98
HCBLM_I001-05000041 | IHEP0041 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-27 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.73 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.2 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.1 | [10 empty] | 4.7 | 98
HCBLM_I001-05000042 | IHEP0042 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.6 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 53.5 | [10 empty] | 4.7 | 98
HCBLM_I001-05000043 | IHEP0043 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli | [5 empty] | 1800 | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 |  |  | 2 | 25.07.2006 |  | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.226 | [10 empty] | 4.7 | 98
HCBLM_I001-05000044 | IHEP0044 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.49 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.2 | [10 empty] | 4.7 | 98
HCBLM_I001-05000045 | IHEP0045 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.6 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.8 | [10 empty] | 4.7 | 98
HCBLM_I001-05000046 | IHEP0046 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.9 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.6 | [10 empty] | 4.7 | 98
HCBLM_I001-05000047 | IHEP0047 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-27 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.73 | 0.74 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.6 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.5 | [10 empty] | 4.7 | 98
HCBLM_I001-05000048 | IHEP0048 | IHEP | Manufacturing |  | Done | Not Ok | 1 | 2006-04-27 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Cancelled | Cancelled | [15 empty] | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.74 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.6 | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.5 | [10 empty] | 4.7 | 98
HCBLM_I001-05000049 | IHEP0049 | IHEP | Manufacturing |  | Done | Not Ok | 1 | 2006-04-27 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Cancelled | Cancelled | [15 empty] | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | 1800 | Done | Ok | 5 | 2006-07-25 | Grishin | [7 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3
HCBLM_I001-05000050 | IHEP0050 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-18 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli | [5 empty] | 1800 | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.3 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.4 | [10 empty] | 4.7 | 98
HCBLM_I001-05000051 | IHEP0051 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-18 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 |  | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.3 | [10 empty] | 4.7 | 98
HCBLM_I001-05000052 | IHEP0052 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-18 | Tikhonov |  |  |  |  | 0.73 | 0.97 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli | [6 empty] | Done | Ok | 5 | 2006-06-16 | Grishin |  | 1 |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.3 | [10 empty] | 4.7 | 98
HCBLM_I001-05000053 | IHEP0053 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.764 | [10 empty] | 4.7 | 98
HCBLM_I001-05000054 | IHEP0054 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.226 | [10 empty] | 4.7 | 98
HCBLM_I001-05000055 | IHEP0055 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.098 | [10 empty] | 4.7 | 98 | [17 empty] | 16
HCBLM_I001-05000056 | IHEP0056 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.898 | [10 empty] | 4.7 | 98
HCBLM_I001-05000057 | IHEP0057 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.74 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.344 | [10 empty] | 4.7 | 98
HCBLM_I001-05000058 | IHEP0058 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Seleznev |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.97 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.686 | [10 empty] | 4.7 | 98
HCBLM_I001-05000059 | IHEP0059 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.524 | [10 empty] | 4.7 | 98
HCBLM_I001-05000060 | IHEP0060 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.914 | [10 empty] | 4.7 | 98 | [17 empty] | 1
HCBLM_I001-05000061 | IHEP0061 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Seleznev |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.24 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 68.768 | [10 empty] | 4.7 | 98
HCBLM_I001-05000062 | IHEP0062 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Seleznev |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.712 | [10 empty] | 4.7 | 98 | [17 empty] | 2
HCBLM_I001-05000063 | IHEP0063 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Seleznev |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.344 | [10 empty] | 4.7 | 98 | [17 empty] | 1
HCBLM_I001-05000064 | IHEP0064 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.9 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.726 | [10 empty] | 4.7 | 98
HCBLM_I001-05000065 | IHEP0065 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Cancelled | Not Ok | [8 empty] | 1 | [6 empty] | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.22 | [10 empty] | 4.7 | 98 | [17 empty] | 16
HCBLM_I001-05000066 | IHEP0066 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 68.146 | [10 empty] | 4.7 | 98
HCBLM_I001-05000067 | IHEP0067 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.742 | [10 empty] | 4.7 | 98 | [17 empty] | 4
HCBLM_I001-05000068 | IHEP0068 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.532 | [10 empty] | 4.7 | 98
HCBLM_I001-05000069 | IHEP0069 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Koshelev |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | [18 empty] | 4.7 | 98
HCBLM_I001-05000070 | IHEP0070 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Koshelev |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-06-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.072 | [10 empty] | 4.7 | 98
HCBLM_I001-05000071 | IHEP0071 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Koshelev |  |  |  |  | 0.97 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  | 2nA in Protv |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.758 | [10 empty] | 4.7 | 98
HCBLM_I001-05000072 | IHEP0072 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Koshelev |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | [18 empty] | 4.7 | 98
HCBLM_I001-05000073 | IHEP0073 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.178 | [10 empty] | 4.7 | 98
HCBLM_I001-05000074 | IHEP0074 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.148 | [10 empty] | 4.7 | 98 | [17 empty] | 14
HCBLM_I001-05000075 | IHEP0075 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-22 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.756 | [10 empty] | 4.7 | 98
HCBLM_I001-05000076 | IHEP0076 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-15 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 72 | [10 empty] | 4.7 | 98
HCBLM_I001-05000077 | IHEP0077 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Cancelled | Not Ok | 2 | [7 empty] | 1 | [6 empty] | Done | Ok | 3 | 2006-05-25 | Koshelev |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.708 | [10 empty] | 4.7 | 98 | [17 empty] | 3
HCBLM_I001-05000078 | IHEP0078 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 1.1 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.97 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.736 | [10 empty] | 4.7 | 98 | [17 empty] | 1
HCBLM_I001-05000079 | IHEP0079 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Koshelev |  |  |  |  | 0.48 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 68.748 | [10 empty] | 4.7 | 98
HCBLM_I001-05000080 | IHEP0080 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.394 | [10 empty] | 4.7 | 98
HCBLM_I001-05000081 | IHEP0081 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-19 | Tikhonov |  |  |  |  | 0.78 | IHEP | 2000 | Cancelled | Not Ok | [8 empty] | 1 | [6 empty] | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.068 | [10 empty] | 4.7 | 98
HCBLM_I001-05000082 | IHEP0082 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | [18 empty] | 4.7 | 98
HCBLM_I001-05000083 | IHEP0083 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.454 | [10 empty] | 4.7 | 98 | [17 empty] | 15
HCBLM_I001-05000084 | IHEP0084 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-22 | Tikhonov |  |  |  |  | 0.97 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 68.684 | [10 empty] | 4.7 | 98
HCBLM_I001-05000085 | IHEP0085 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-19 | Tikhonov |  |  |  |  | 0.28 | IHEP | 2000 | Cancelled | Not Ok | [8 empty] | 1 | [6 empty] | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.982 | [10 empty] | 4.7 | 98
HCBLM_I001-05000086 | IHEP0086 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-22 | Tikhonov |  |  |  |  | 0.73 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.338 | [10 empty] | 4.7 | 98
HCBLM_I001-05000087 | IHEP0087 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-22 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.904 | [10 empty] | 4.7 | 98
HCBLM_I001-05000088 | IHEP0088 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.97 | IHEP | 2000 | Cancelled | Not Ok | [8 empty] | 1 | [6 empty] | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 68.524 | [10 empty] | 4.7 | 98 | [17 empty] | 3
HCBLM_I001-05000089 | IHEP0089 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-02 | Tikhonov |  |  |  |  | 0.73 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.84 | [10 empty] | 4.7 | 98 | [17 empty] | 4
HCBLM_I001-05000090 | IHEP0090 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar004_8 | 19 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 72.148 | [10 empty] | 4.7 | 98
HCBLM_I001-05000091 | IHEP0091 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.244 | [10 empty] | 4.7 | 98 | [17 empty] | 3
HCBLM_I001-05000092 | IHEP0092 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 0.97 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.15 | [10 empty] | 4.7 | 98
HCBLM_I001-05000093 | IHEP0093 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.74 | [10 empty] | 4.7 | 98 | [17 empty] | 12
HCBLM_I001-05000094 | IHEP0094 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 9.8 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.628 | [10 empty] | 4.7 | 98 | [17 empty] | 2
HCBLM_I001-05000095 | IHEP0095 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.73 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 70.74 | [10 empty] | 4.7 | 98
HCBLM_I001-05000096 | IHEP0096 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 72.018 | [10 empty] | 4.7 | 98
HCBLM_I001-05000097 | IHEP0097 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.208 | [10 empty] | 4.7 | 98
HCBLM_I001-05000098 | IHEP0098 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 1.1 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.156 | [10 empty] | 4.7 | 98 | [17 empty] | 4
HCBLM_I001-05000099 | IHEP0099 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 69.2 | [10 empty] | 4.7 | 98
HCBLM_I001-05000100 | IHEP0100 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 71.236 | [10 empty] | 4.7 | 98
HCBLM_I001-05000101 | IHEP0101 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-19 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-29 | Tikhonov |  |  |  |  | 0.73 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.57 | [10 empty] | 4.7 | 98
HCBLM_I001-05000102 | IHEP0102 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.372 | [10 empty] | 4.7 | 98
HCBLM_I001-05000103 | IHEP0103 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.192 | [10 empty] | 4.7 | 98
HCBLM_I001-05000104 | IHEP0104 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 1.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.01 | [10 empty] | 4.7 | 98
HCBLM_I001-05000105 | IHEP0105 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.48 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.196 | [10 empty] | 4.7 | 98
HCBLM_I001-05000106 | IHEP0106 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.204 | [10 empty] | 4.7 | 98 | [17 empty] | 9
HCBLM_I001-05000107 | IHEP0107 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.814 | [10 empty] | 4.7 | 98
HCBLM_I001-05000108 | IHEP0108 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.308 | [10 empty] | 4.7 | 98
HCBLM_I001-05000109 | IHEP0109 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.218 | [10 empty] | 4.7 | 98
HCBLM_I001-05000110 | IHEP0110 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.866 | [10 empty] | 4.7 | 98 | [17 empty] | 5
HCBLM_I001-05000111 | IHEP0111 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.726 | [10 empty] | 4.7 | 98
HCBLM_I001-05000112 | IHEP0112 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 68.944 | [10 empty] | 4.7 | 98
HCBLM_I001-05000113 | IHEP0113 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Koshelev |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 68.958 | [10 empty] | 4.7 | 98
HCBLM_I001-05000114 | IHEP0114 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 68.482 | [10 empty] | 4.7 | 98
HCBLM_I001-05000115 | IHEP0115 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-24 | Seleznev |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.902 | [10 empty] | 4.7 | 98
HCBLM_I001-05000116 | IHEP0116 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.684 | [10 empty] | 4.7 | 98
HCBLM_I001-05000117 | IHEP0117 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.386 | [10 empty] | 4.7 | 98
HCBLM_I001-05000118 | IHEP0118 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Seleznev |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.9 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.858 | [10 empty] | 4.7 | 98
HCBLM_I001-05000119 | IHEP0119 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.48 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2
HCBLM_I001-05000120 | IHEP0120 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.97 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 68.354 | [10 empty] | 4.7 | 98
HCBLM_I001-05000121 | IHEP0121 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.96 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.506 | [10 empty] | 4.7 | 98
HCBLM_I001-05000122 | IHEP0122 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.48 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3
HCBLM_I001-05000123 | IHEP0123 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.97 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.004 | [10 empty] | 4.7 | 98
HCBLM_I001-05000124 | IHEP0124 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.78 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.668 | [10 empty] | 4.7 | 98
HCBLM_I001-05000125 | IHEP0125 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.132 | [10 empty] | 4.7 | 98 | [17 empty] | 4
HCBLM_I001-05000126 | IHEP0126 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2206-05-02 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.862 | [10 empty] | 4.7 | 98
HCBLM_I001-05000127 | IHEP0127 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.78 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.026 | [10 empty] | 4.7 | 98
HCBLM_I001-05000128 | IHEP0128 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.504 | [10 empty] | 4.7 | 98
HCBLM_I001-05000129 | IHEP0129 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.456 | [10 empty] | 4.7 | 98
HCBLM_I001-05000130 | IHEP0130 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-05-24 | Larionov |  |  |  |  | 5 | 1 | lar005_1;lar005_2 | 21 | 1111 | 240 | 6 | lar005_3;lar005_4 | Done | Ok | 3 | 2006-05-25 | Tikhonov |  |  |  |  | 0.24 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.124 | [10 empty] | 4.7 | 98
HCBLM_I001-05000131 | IHEP0131 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-07 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.002 | [10 empty] | 4.7 | 98
HCBLM_I001-05000132 | IHEP0132 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6
HCBLM_I001-05000133 | IHEP0133 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.038 | [10 empty] | 4.7 | 98
HCBLM_I001-05000134 | IHEP0134 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.5 | [10 empty] | 4.7 | 98
HCBLM_I001-05000135 | IHEP0135 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.192 | [10 empty] | 4.7 | 98
HCBLM_I001-05000136 | IHEP0136 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 |  | Mishagin | [16 empty] | 4.7 | 98
HCBLM_I001-05000137 | IHEP0137 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-01 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 75.802 | [10 empty] | 4.7 | 98
HCBLM_I001-05000138 | IHEP0138 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-01 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.8 | [10 empty] | 4.7 | 98
HCBLM_I001-05000139 | IHEP0139 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.726 | [10 empty] | 4.7 | 98 | [17 empty] | 7
HCBLM_I001-05000140 | IHEP0140 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 67.21 | [10 empty] | 4.7 | 98
HCBLM_I001-05000141 | IHEP0141 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-07 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.154 | [10 empty] | 4.7 | 98
HCBLM_I001-05000142 | IHEP0142 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.806 | [10 empty] | 4.7 | 98
HCBLM_I001-05000143 | IHEP0143 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8
HCBLM_I001-05000144 | IHEP0144 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.75 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.7 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.696 | [10 empty] | 4.7 | 98
HCBLM_I001-05000145 | IHEP0145 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  | >2nA at cern |  |  | 2000 | 0.6
HCBLM_I001-05000146 | IHEP0146 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-01 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.5 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.968 | [10 empty] | 4.7 | 98
HCBLM_I001-05000147 | IHEP0147 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.882 | [10 empty] | 4.7 | 98
HCBLM_I001-05000148 | IHEP0148 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.964 | [10 empty] | 4.7 | 98
HCBLM_I001-05000149 | IHEP0149 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-01 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 68.696 | [10 empty] | 4.7 | 98
HCBLM_I001-05000150 | IHEP0150 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2
HCBLM_I001-05000151 | IHEP0151 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 72.99 | [10 empty] | 4.7 | 98
HCBLM_I001-05000152 | IHEP0152 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-07 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.188 | [10 empty] | 4.7 | 98
HCBLM_I001-05000153 | IHEP0153 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 74.254 | [10 empty] | 4.7 | 98
HCBLM_I001-05000154 | IHEP0154 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-07 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.314 | [10 empty] | 4.7 | 98
HCBLM_I001-05000155 | IHEP0155 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.706 | [10 empty] | 4.7 | 98
HCBLM_I001-05000156 | IHEP0156 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-01 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.986 | [10 empty] | 4.7 | 98
HCBLM_I001-05000157 | IHEP0157 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.644 | [10 empty] | 4.7 | 98
HCBLM_I001-05000158 | IHEP0158 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-01 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 72.328 | [10 empty] | 4.7 | 98
HCBLM_I001-05000159 | IHEP0159 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 76.264 | [10 empty] | 4.7 | 98 | [17 empty] | 5
HCBLM_I001-05000160 | IHEP0160 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.344 | [10 empty] | 4.7 | 98 | [17 empty] | 1
HCBLM_I001-05000161 | IHEP0161 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 68.234 | [10 empty] | 4.7 | 98 | [17 empty] | 9
HCBLM_I001-05000162 | IHEP0162 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.578 | [10 empty] | 4.7 | 98
HCBLM_I001-05000163 | IHEP0163 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-07 | Tikhonov |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.3 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.37 | [10 empty] | 4.7 | 98
HCBLM_I001-05000164 | IHEP0164 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.6 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.456 | [10 empty] | 4.7 | 98
HCBLM_I001-05000165 | IHEP0165 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 1.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.404 | [10 empty] | 4.7 | 98
HCBLM_I001-05000166 | IHEP0166 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 68.844 | [10 empty] | 4.7 | 98
HCBLM_I001-05000167 | IHEP0167 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.78 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.84 | [10 empty] | 4.7 | 98
HCBLM_I001-05000168 | IHEP0168 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 1.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 72.124 | [10 empty] | 4.7 | 98
HCBLM_I001-05000169 | IHEP0169 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.486 | [10 empty] | 4.7 | 98
HCBLM_I001-05000170 | IHEP0170 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.85 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.08 | [10 empty] | 4.7 | 98
HCBLM_I001-05000171 | IHEP0171 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 0.24 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.562 | [10 empty] | 4.7 | 98
HCBLM_I001-05000172 | IHEP0172 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.96 | [10 empty] | 4.7 | 98
HCBLM_I001-05000173 | IHEP0173 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-02 | Tikhonov |  |  |  |  | 0.1 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.004 | [10 empty] | 4.7 | 98
HCBLM_I001-05000174 | IHEP0174 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.53 | [10 empty] | 4.7 | 98
HCBLM_I001-05000175 | IHEP0175 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.306 | [10 empty] | 4.7 | 98
HCBLM_I001-05000176 | IHEP0176 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | [40 empty] | 4
HCBLM_I001-05000177 | IHEP0177 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-02 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.224 | [10 empty] | 4.7 | 98
HCBLM_I001-05000178 | IHEP0178 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.002 | [10 empty] | 4.7 | 98 | [17 empty] | 7
HCBLM_I001-05000179 | IHEP0179 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-01 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 67.372 | [10 empty] | 4.7 | 98 | [17 empty] | 6
HCBLM_I001-05000180 | IHEP0180 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  | 10p in protv |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.158 | [10 empty] | 4.7 | 98
HCBLM_I001-05000181 | IHEP0181 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-02 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.916 | [10 empty] | 4.7 | 98
HCBLM_I001-05000182 | IHEP0182 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 1.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 1.7 | 1.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.006 | [10 empty] | 4.7 | 98
HCBLM_I001-05000183 | IHEP0183 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-02 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.798 | [10 empty] | 4.7 | 98
HCBLM_I001-05000184 | IHEP0184 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-02 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.178 | [10 empty] | 4.7 | 98
HCBLM_I001-05000185 | IHEP0185 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 1.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 69.168 | [10 empty] | 4.7 | 98
HCBLM_I001-05000186 | IHEP0186 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.97 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.304 | [10 empty] | 4.7 | 98 | [17 empty] | 7
HCBLM_I001-05000187 | IHEP0187 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-02 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.232 | [10 empty] | 4.7 | 98
HCBLM_I001-05000188 | IHEP0188 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.95 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-02 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.43 | [10 empty] | 4.7 | 98
HCBLM_I001-05000189 | IHEP0189 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.8 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 72.732 | [10 empty] | 4.7 | 98
HCBLM_I001-05000190 | IHEP0190 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 71.17 | [10 empty] | 4.7 | 98
HCBLM_I001-05000191 | IHEP0191 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 70.656 | [10 empty] | 4.7 | 98
HCBLM_I001-05000192 | IHEP0192 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 68.95 | [10 empty] | 4.7 | 98
HCBLM_I001-05000193 | IHEP0193 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-01 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 68.39 | [10 empty] | 4.7 | 98
HCBLM_I001-05000194 | IHEP0194 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.73 | 0.73 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-18 | Mishagin |  |  |  |  | 1500 | 68.464 | [10 empty] | 4.7 | 98
HCBLM_I001-05000195 | IHEP0195 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-06 | Tikhonov |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-06 | Tereshin |  |  |  |  | 7 | 1 | ter007_1 | 22 | 1110 | 240 | 6 | no | Done | Ok | 3 | 2006-06-09 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 73.81 | [10 empty] | 4.7 | 98
HCBLM_I001-05000196 | IHEP0196 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-30 | Tikhonov |  |  |  |  | 0.97 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-02 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.876 | [10 empty] | 4.7 | 98 | [17 empty] | 7
HCBLM_I001-05000197 | IHEP0197 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.758 | [10 empty] | 4.7 | 98
HCBLM_I001-05000198 | IHEP0198 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-06 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.934 | [10 empty] | 4.7 | 98 | [17 empty] | 7
HCBLM_I001-05000199 | IHEP0199 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  | >2nA in protv |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 74.248 | [10 empty] | 4.7 | 98
HCBLM_I001-05000200 | IHEP0200 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-29 | Tikhonov |  |  |  |  | 0.48 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.768 | [10 empty] | 4.7 | 98 | [17 empty] | 12
HCBLM_I001-05000201 | IHEP0201 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 2.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.17 | [10 empty] | 4.7 | 98 | [17 empty] | 13
HCBLM_I001-05000202 | IHEP0202 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-22 | Tikhonov |  |  |  |  | 0.73 | IHEP | 2000 | Done | Ok | 2 | 2006-06-01 | Larionov |  |  |  |  | 6 | 1 | lar006_1;lar006_2 | 20 | 1106 | 240 | 6 | lar006_3;lar006_4 | Done | Ok | 3 | 2006-06-05 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.086 | [10 empty] | 4.7 | 98
HCBLM_I001-05000203 | IHEP0203 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.934 | [10 empty] | 4.7 | 98
HCBLM_I001-05000204 | IHEP0204 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.4 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.716 | [10 empty] | 4.7 | 98 | [17 empty] | 8
HCBLM_I001-05000205 | IHEP0205 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.87 | [10 empty] | 4.7 | 98
HCBLM_I001-05000206 | IHEP0206 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 80 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.938 | [10 empty] | 4.7 | 98
HCBLM_I001-05000207 | IHEP0207 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 0.6 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.042 | [10 empty] | 4.7 | 98
HCBLM_I001-05000208 | IHEP0208 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 310 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.6 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.098 | [10 empty] | 4.7 | 98
HCBLM_I001-05000209 | IHEP0209 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.998 | [10 empty] | 4.7 | 98
HCBLM_I001-05000210 | IHEP0210 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 2.9 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.8 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.76 | [10 empty] | 4.7 | 98
HCBLM_I001-05000211 | IHEP0211 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-14 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.9 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  | >2nA at cern |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 72.12 | [10 empty] | 4.7 | 98
HCBLM_I001-05000212 | IHEP0212 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.3 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.448 | [10 empty] | 4.7 | 98
HCBLM_I001-05000213 | IHEP0213 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.872 | [10 empty] | 4.7 | 98
HCBLM_I001-05000214 | IHEP0214 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 1.2 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.084 | [10 empty] | 4.7 | 98 | [17 empty] | 1
HCBLM_I001-05000215 | IHEP0215 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.54 | [10 empty] | 4.7 | 98
HCBLM_I001-05000216 | IHEP0216 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 63 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.3 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.354 | [10 empty] | 4.7 | 98 | [17 empty] | 5
HCBLM_I001-05000217 | IHEP0217 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 1.8 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.768 | [10 empty] | 4.7 | 98
HCBLM_I001-05000218 | IHEP0218 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 0.7 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.41 | [10 empty] | 4.7 | 98
HCBLM_I001-05000219 | IHEP0219 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Koshelev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.2 | [10 empty] | 4.7 | 98
HCBLM_I001-05000220 | IHEP0220 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 12 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.5 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.236 | [10 empty] | 4.7 | 98
HCBLM_I001-05000221 | IHEP0221 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Koshelev |  |  |  |  | 6.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.7 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.1 | [10 empty] | 4.7 | 98
HCBLM_I001-05000222 | IHEP0222 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.8 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.332 | [10 empty] | 4.7 | 98
HCBLM_I001-05000223 | IHEP0223 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.8 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.426 | [10 empty] | 4.7 | 98
HCBLM_I001-05000224 | IHEP0224 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 20 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-20 | Tikhonov |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.276 | [10 empty] | 4.7 | 98
HCBLM_I001-05000225 | IHEP0225 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.6 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.084 | [10 empty] | 4.7 | 98
HCBLM_I001-05000226 | IHEP0226 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 1500 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.9 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.21 | [10 empty] | 4.7 | 98
HCBLM_I001-05000227 | IHEP0227 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.8 | [10 empty] | 4.7 | 98
HCBLM_I001-05000228 | IHEP0228 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-14 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.3 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.558 | [10 empty] | 4.7 | 98 | [17 empty] | 6
HCBLM_I001-05000229 | IHEP0229 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.34 | [10 empty] | 4.7 | 98
HCBLM_I001-05000230 | IHEP0230 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.422 | [10 empty] | 4.7 | 98 | [17 empty] | 14
HCBLM_I001-05000231 | IHEP0231 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.134 | [10 empty] | 4.7 | 98 | [17 empty] | 8
HCBLM_I001-05000232 | IHEP0232 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 3.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.7 | [10 empty] | 4.7 | 98
HCBLM_I001-05000233 | IHEP0233 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.216 | [10 empty] | 4.7 | 98 | [17 empty] | 11
HCBLM_I001-05000234 | IHEP0234 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.4 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.5 | [10 empty] | 4.7 | 98
HCBLM_I001-05000235 | IHEP0235 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-14 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.284 | [10 empty] | 4.7 | 98
HCBLM_I001-05000236 | IHEP0236 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-14 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.708 | [10 empty] | 4.7 | 98
HCBLM_I001-05000237 | IHEP0237 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 15 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.238 | [10 empty] | 4.7 | 98
HCBLM_I001-05000238 | IHEP0238 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 3 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.946 | [10 empty] | 4.7 | 98
HCBLM_I001-05000239 | IHEP0239 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.566 | [10 empty] | 4.7 | 98
HCBLM_I001-05000240 | IHEP0240 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 0.6 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.104 | [10 empty] | 4.7 | 98
HCBLM_I001-05000241 | IHEP0241 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 0.7 | IHEP | 1800 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 72.048 | [10 empty] | 4.7 | 98
HCBLM_I001-05000242 | IHEP0242 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 3.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.42 | [10 empty] | 4.7 | 98 | [17 empty] | 6
HCBLM_I001-05000243 | IHEP0243 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Tikhonov |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.776 | [10 empty] | 4.7 | 98 | [17 empty] | 9
HCBLM_I001-05000244 | IHEP0244 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-26 | Koshelev |  |  |  |  | 2.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.1 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.774 | [10 empty] | 4.7 | 98
HCBLM_I001-05000245 | IHEP0245 | IHEP | Manufacturing |  | Done | Ok | 1 | [6 empty] | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.944 | [10 empty] | 4.7 | 98 | [17 empty] | 8
HCBLM_I001-05000246 | IHEP0246 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  | >2nA in protv |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.986 | [10 empty] | 4.7 | 98
HCBLM_I001-05000247 | IHEP0247 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Koshelev |  |  |  |  | 5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.4 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.338 | [10 empty] | 4.7 | 98
HCBLM_I001-05000248 | IHEP0248 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-15 | Tikhonov |  |  |  |  | 10 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.7 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.764 | [10 empty] | 4.7 | 98 | [17 empty] | 5
HCBLM_I001-05000249 | IHEP0249 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 38 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.6 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.35 | [10 empty] | 4.7 | 98
HCBLM_I001-05000250 | IHEP0250 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-16 | Tikhonov |  |  |  |  | 9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-20 | Larionov |  |  |  |  | 9 | 1 | lar009_1;lar009_2 | 24 | 1121 | 240 | 6 | lar009_3;lar009_4 | Done | Ok | 3 | 2006-06-21 | Tikhonov |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.604 | [10 empty] | 4.7 | 98
HCBLM_I001-05000251 | IHEP0251 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 3.1 | IHEP | 2000 | Cancelled | Not Ok | [8 empty] | 1 | [6 empty] | Cancelled | not Ok | 3 | [32 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000252 | IHEP0252 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | [40 empty] | 12
HCBLM_I001-05000253 | IHEP0253 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 21.1 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.753 | [10 empty] | 4.7 | 98
HCBLM_I001-05000254 | IHEP0254 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 14.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.3 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.342 | [10 empty] | 4.7 | 98 | [17 empty] | 11
HCBLM_I001-05000255 | IHEP0255 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 1454.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.1 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.132 | [10 empty] | 4.7 | 98 | [17 empty] | 3
HCBLM_I001-05000256 | IHEP0256 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.21 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.406 | [10 empty] | 4.7 | 98 | [17 empty] | 15
HCBLM_I001-05000257 | IHEP0257 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.866 | [10 empty] | 4.7 | 98 | [17 empty] | 14
HCBLM_I001-05000258 | IHEP0258 | IHEP | Manufacturing |  | Done | Ok | 1 | [8 empty] | 2000 | Cancelled | Not Ok | [8 empty] | 1 | [6 empty] | Cancelled | Not Ok | 3
HCBLM_I001-05000259 | IHEP0259 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.6 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2
HCBLM_I001-05000260 | IHEP0260 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.469 | [10 empty] | 4.7 | 98 | [17 empty] | 13
HCBLM_I001-05000261 | IHEP0261 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.28 | [10 empty] | 4.7 | 98 | [17 empty] | 6
HCBLM_I001-05000262 | IHEP0262 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.1
HCBLM_I001-05000263 | IHEP0263 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.238 | [10 empty] | 4.7 | 98 | [17 empty] | 14
HCBLM_I001-05000264 | IHEP0264 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 183.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.5 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.212 | [10 empty] | 4.7 | 98 | [17 empty] | 5
HCBLM_I001-05000265 | IHEP0265 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 11.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.77 | [10 empty] | 4.7 | 98
HCBLM_I001-05000266 | IHEP0266 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 6.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-23 | Tikhonov |  |  |  |  | 0.7 | 0.2 | 2000 | [23 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.3 | [10 empty] | 4.7 | 98
HCBLM_I001-05000267 | IHEP0267 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.192 | [10 empty] | 4.7 | 98
HCBLM_I001-05000268 | IHEP0268 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.558 | [10 empty] | 4.7 | 98
HCBLM_I001-05000269 | IHEP0269 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 1245.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-23 | Tikhonov |  |  |  |  | 0.3 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.646 | [10 empty] | 4.7 | 98
HCBLM_I001-05000270 | IHEP0270 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-23 | Tikhonov |  |  |  |  | 0.6 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.844 | [10 empty] | 4.7 | 98 | [17 empty] | 2
HCBLM_I001-05000271 | IHEP0271 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.5 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.436 | [10 empty] | 4.7 | 98
HCBLM_I001-05000272 | IHEP0272 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.574 | [10 empty] | 4.7 | 98
HCBLM_I001-05000273 | IHEP0273 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.884 | [10 empty] | 4.7 | 98
HCBLM_I001-05000274 | IHEP0274 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.5 | IHEP | 2000 | Cancelled | Not Ok | 2 | [7 empty] | 1 | [6 empty] | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.22 | [10 empty] | 4.7 | 98
HCBLM_I001-05000275 | IHEP0275 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-26 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | [40 empty] | 16
HCBLM_I001-05000276 | IHEP0276 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.328 | [10 empty] | 4.7 | 98
HCBLM_I001-05000277 | IHEP0277 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.412 | [10 empty] | 4.7 | 98
HCBLM_I001-05000278 | IHEP0278 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.996 | [10 empty] | 4.7 | 98
HCBLM_I001-05000279 | IHEP0279 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.714 | [10 empty] | 4.7 | 98 | [17 empty] | 4
HCBLM_I001-05000280 | IHEP0280 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.152 | [10 empty] | 4.7 | 98
HCBLM_I001-05000281 | IHEP0281 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.3 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.538 | [10 empty] | 4.7 | 98
HCBLM_I001-05000282 | IHEP0282 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 1599.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-23 | Tikhonov |  |  |  |  | 0.3 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.002 | [10 empty] | 4.7 | 98 | [17 empty] | 10
HCBLM_I001-05000283 | IHEP0283 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Koshelev |  |  |  |  | 1408.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.14 | [10 empty] | 4.7 | 98 | [17 empty] | 15
HCBLM_I001-05000284 | IHEP0284 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.026 | [10 empty] | 4.7 | 98 | [17 empty] | 2
HCBLM_I001-05000285 | IHEP0285 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.778 | [10 empty] | 4.7 | 98
HCBLM_I001-05000286 | IHEP0286 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.9 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.388 | [10 empty] | 4.7 | 98
HCBLM_I001-05000287 | IHEP0287 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-23 | Tikhonov |  |  |  |  | 0.3 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3
HCBLM_I001-05000288 | IHEP0288 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.6 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.456 | [10 empty] | 4.7 | 98
HCBLM_I001-05000289 | IHEP0289 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.636 | [10 empty] | 4.7 | 98 | [17 empty] | 8
HCBLM_I001-05000290 | IHEP0290 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.1 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.278 | [10 empty] | 4.7 | 98 | [17 empty] | 10
HCBLM_I001-05000291 | IHEP0291 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.694 | [10 empty] | 4.7 | 98
HCBLM_I001-05000292 | IHEP0292 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.516 | [10 empty] | 4.7 | 98
HCBLM_I001-05000293 | IHEP0293 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.7 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.246 | [10 empty] | 4.7 | 98
HCBLM_I001-05000294 | IHEP0294 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.3 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.662 | [10 empty] | 4.7 | 98
HCBLM_I001-05000295 | IHEP0295 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.982 | [10 empty] | 4.7 | 98
HCBLM_I001-05000296 | IHEP0296 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.69 | [10 empty] | 4.7 | 98
HCBLM_I001-05000297 | IHEP0297 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-09 | Tikhonov |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.45 | [10 empty] | 4.7 | 98
HCBLM_I001-05000298 | IHEP0298 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-23 | Tikhonov |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.548 | [10 empty] | 4.7 | 98
HCBLM_I001-05000299 | IHEP0299 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.538 | [10 empty] | 4.7 | 98
HCBLM_I001-05000300 | IHEP0300 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.082 | [10 empty] | 4.7 | 98
HCBLM_I001-05000301 | IHEP0301 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-15 | Tereshin |  |  |  |  | 8 | 1 | ter008_1;ter008_2 | 23 | 1107 | 240 | 6 | ter008_3;ter008_4 | Done | Ok | 3 | 2006-06-16 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.538 | [10 empty] | 4.7 | 98 | [17 empty] | 1
HCBLM_I001-05000302 | IHEP0302 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-23 | Tikhonov |  |  |  |  | 0.5 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.958 | [10 empty] | 4.7 | 98
HCBLM_I001-05000303 | IHEP0303 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Tikhonov |  |  |  |  | 0.1 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.498 | [10 empty] | 4.7 | 98
HCBLM_I001-05000304 | IHEP0304 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.044 | [10 empty] | 4.7 | 98 | [17 empty] | 15
HCBLM_I001-05000305 | IHEP0305 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | >1000 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.84 | [10 empty] | 4.7 | 98 | [17 empty] | 12
HCBLM_I001-05000306 | IHEP0306 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 4.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.72 | [10 empty] | 4.7 | 98 | [17 empty] | 11
HCBLM_I001-05000307 | IHEP0307 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Tikhonov |  |  |  |  | 0.9 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.924 | [10 empty] | 4.7 | 98 | [17 empty] | 16
HCBLM_I001-05000308 | IHEP0308 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Tikhonov |  |  |  |  | 0.3 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.118 | [10 empty] | 4.7 | 98 | [17 empty] | 11
HCBLM_I001-05000309 | IHEP0309 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.772 | [10 empty] | 4.7 | 98 | [17 empty] | 9
HCBLM_I001-05000310 | IHEP0310 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-28 | Koshelev |  |  |  |  | 0.3 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 75.76 | [10 empty] | 4.7 | 98 | [17 empty] | 7
HCBLM_I001-05000311 | IHEP0311 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-28 | Koshelev |  |  |  |  | 0.1 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000312 | IHEP0312 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 5.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.026 | [10 empty] | 4.7 | 98
HCBLM_I001-05000313 | IHEP0313 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Mishagin |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5
HCBLM_I001-05000314 | IHEP0314 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 6.3 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Mishagin |  |  |  |  | 0.4 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.05 | [10 empty] | 4.7 | 98
HCBLM_I001-05000315 | IHEP0315 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Mishagin |  |  |  |  | 0.6 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000316 | IHEP0316 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.41 | [10 empty] | 4.7 | 98
HCBLM_I001-05000317 | IHEP0317 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.2 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000318 | IHEP0318 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.1 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.04 | [10 empty] | 4.7 | 98
HCBLM_I001-05000319 | IHEP0319 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.2 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.81 | [10 empty] | 4.7 | 98
HCBLM_I001-05000320 | IHEP0320 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-28 | Mishagin |  |  |  |  | 0.1 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000321 | IHEP0321 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Tikhonov |  |  |  |  | 1.1 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-28 | Koshelev |  |  |  |  | 0.4 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.726 | [10 empty] | 4.7 | 98 | [17 empty] | 9
HCBLM_I001-05000322 | IHEP0322 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-28 | Koshelev |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6
HCBLM_I001-05000323 | IHEP0323 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 1659.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2
HCBLM_I001-05000324 | IHEP0324 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 9.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Mishagin |  |  |  |  | 0.9 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.842 | [10 empty] | 4.7 | 98 | [17 empty] | 15
HCBLM_I001-05000325 | IHEP0325 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Mishagin |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.298 | [10 empty] | 4.7 | 98
HCBLM_I001-05000326 | IHEP0326 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Mishagin |  |  |  |  | 0.8 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.056 | [10 empty] | 4.7 | 98
HCBLM_I001-05000327 | IHEP0327 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 1396 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.152 | [10 empty] | 4.7 | 98
HCBLM_I001-05000328 | IHEP0328 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.098 | [10 empty] | 4.7 | 98
HCBLM_I001-05000329 | IHEP0329 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5
HCBLM_I001-05000330 | IHEP0330 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 23.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-04 | Seleznev |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.02 | [10 empty] | 4.7 | 98 | [17 empty] | 13
HCBLM_I001-05000331 | IHEP0331 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 280 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Mishagin |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6
HCBLM_I001-05000332 | IHEP0332 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Tikhonov |  |  |  |  | 15 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9
HCBLM_I001-05000333 | IHEP0333 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Tikhonov |  |  |  |  | 2.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.058 | [10 empty] | 4.7 | 98
HCBLM_I001-05000334 | IHEP0334 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Tikhonov |  |  |  |  | 2.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-07-13 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.632 | [10 empty] | 4.7 | 98
HCBLM_I001-05000335 | IHEP0335 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Tikhonov |  |  |  |  | 20 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Mishagin |  |  |  |  | 0.9 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1
HCBLM_I001-05000336 | IHEP0336 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1
HCBLM_I001-05000337 | IHEP0337 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.82 | [10 empty] | 4.7 | 98 | [17 empty] | 2
HCBLM_I001-05000338 | IHEP0338 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000339 | IHEP0339 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-20 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-22 | Larionov |  |  |  |  | 10 | 1 | lar010_1;lar010_2 | 27 | 1125 | 220 | 6 | lar010_3;lar010_4 | Done | Ok | 3 | 2006-06-26 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.1 | [10 empty] | 4.7 | 98
HCBLM_I001-05000340 | IHEP0340 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Seleznev |  |  |  |  | 0.8 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.546 | [10 empty] | 4.7 | 98 | [17 empty] | 16
HCBLM_I001-05000341 | IHEP0341 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 1446 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-28 | Mishagin |  |  |  |  | 0.1 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.468 | [10 empty] | 4.7 | 98 | [17 empty] | 3
HCBLM_I001-05000342 | IHEP0342 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.1 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.068 | [10 empty] | 4.7 | 98 | [17 empty] | 8
HCBLM_I001-05000343 | IHEP0343 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 1440 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-28 | Koshelev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.502 | [10 empty] | 4.7 | 98 | [17 empty] | 6
HCBLM_I001-05000344 | IHEP0344 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Koshelev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.744 | [10 empty] | 4.7 | 98 | [17 empty] | 10
HCBLM_I001-05000345 | IHEP0345 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-27 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.834 | [10 empty] | 4.7 | 98 | [17 empty] | 8
HCBLM_I001-05000346 | IHEP0346 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 4.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Seleznev |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.7122 | [10 empty] | 4.7 | 98 | [17 empty] | 16
HCBLM_I001-05000347 | IHEP0347 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Koshelev |  |  |  |  | 0.5 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.934 | [10 empty] | 4.7 | 98
HCBLM_I001-05000348 | IHEP0348 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 1.5 | IHEP | 2000 | Cancelled | Not Ok | 2 | [7 empty] | 1 | [6 empty] | Done | Ok | 3 | 2006-06-28 | Mishagin |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.544 | [10 empty] | 4.7 | 98
HCBLM_I001-05000349 | IHEP0349 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 1.4 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Koshelev |  |  |  |  | 0.2 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.788 | [10 empty] | 4.7 | 98
HCBLM_I001-05000350 | IHEP0350 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Tikhonov |  |  |  |  | 124 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-07-13 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.55 | [10 empty] | 4.7 | 98
HCBLM_I001-05000351 | IHEP0351 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Seleznev |  |  |  |  | 0.4 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.28 | [10 empty] | 4.7 | 98
HCBLM_I001-05000352 | IHEP0352 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-28 | Seleznev |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.912 | [10 empty] | 4.7 | 98 | [17 empty] | 2
HCBLM_I001-05000353 | IHEP0353 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Koshelev |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.3 | [10 empty] | 4.7 | 98 | [17 empty] | 11
HCBLM_I001-05000354 | IHEP0354 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.912 | [10 empty] | 4.7 | 98
HCBLM_I001-05000355 | IHEP0355 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Koshelev |  |  |  |  | 0.9 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.998 | [10 empty] | 4.7 | 98 | [17 empty] | 9
HCBLM_I001-05000356 | IHEP0356 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 9.2 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-28 | Koshelev |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | [40 empty] | 12
HCBLM_I001-05000357 | IHEP0357 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 2.3 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Koshelev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.11 | [10 empty] | 4.7 | 98
HCBLM_I001-05000358 | IHEP0358 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-23 | Koshelev |  |  |  |  | 1.4 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Koshelev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.002 | [10 empty] | 4.7 | 98 | [17 empty] | 10
HCBLM_I001-05000359 | IHEP0359 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-22 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-06-26 | Larionov |  |  |  |  | 11 | 1 | lar011_1;lar011_2 | 27 | 1125 | 220 | 6 | lar011_3;lar011_4 | Done | Ok | 3 | 2006-06-29 | Seleznev |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.2591 | [10 empty] | 4.7 | 98
HCBLM_I001-05000360 | IHEP0360 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Mishagin |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.538 | [10 empty] | 4.7 | 98 | [17 empty] | 10
HCBLM_I001-05000361 | IHEP0361 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.44 | [10 empty] | 4.7 | 98
HCBLM_I001-05000362 | IHEP0362 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.1 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.814 | [10 empty] | 4.7 | 98 | [17 empty] | 5
HCBLM_I001-05000363 | IHEP0363 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-30 | Koshelev |  |  |  |  | 1219.8 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.214 | [10 empty] | 4.7 | 98
HCBLM_I001-05000364 | IHEP0364 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 1202.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.028 | [10 empty] | 4.7 | 98
HCBLM_I001-05000365 | IHEP0365 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.602 | [10 empty] | 4.7 | 98
HCBLM_I001-05000366 | IHEP0366 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.076 | [10 empty] | 4.7 | 98 | [17 empty] | 13
HCBLM_I001-05000367 | IHEP0367 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.404 | [10 empty] | 4.7 | 98
HCBLM_I001-05000368 | IHEP0368 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.166 | [10 empty] | 4.7 | 98
HCBLM_I001-05000369 | IHEP0369 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Koshelev |  |  |  |  | 25.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.708 | [10 empty] | 4.7 | 98
HCBLM_I001-05000370 | IHEP0370 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | >1000 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.257 | [10 empty] | 4.7 | 98 | [17 empty] | 14
HCBLM_I001-05000371 | IHEP0371 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Koshelev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.624 | [10 empty] | 4.7 | 98
HCBLM_I001-05000372 | IHEP0372 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Seleznev |  |  |  |  | 250 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.6 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.751 | [10 empty] | 4.7 | 98
HCBLM_I001-05000373 | IHEP0373 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Seleznev |  |  |  |  | 0.3 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.632 | [10 empty] | 4.7 | 98
HCBLM_I001-05000374 | IHEP0374 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.228 | [10 empty] | 4.7 | 98
HCBLM_I001-05000375 | IHEP0375 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.706 | [10 empty] | 4.7 | 98
HCBLM_I001-05000376 | IHEP0376 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.218 | [10 empty] | 4.7 | 98
HCBLM_I001-05000377 | IHEP0377 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 1394.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.322 | [10 empty] | 4.7 | 98 | [17 empty] | 12
HCBLM_I001-05000378 | IHEP0378 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 1380.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.7 | [10 empty] | 4.7 | 98
HCBLM_I001-05000379 | IHEP0379 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.57 | [10 empty] | 4.7 | 98
HCBLM_I001-05000380 | IHEP0380 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.818 | [10 empty] | 4.7 | 98
HCBLM_I001-05000381 | IHEP0381 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.577 | [10 empty] | 4.7 | 98 | [17 empty] | 13
HCBLM_I001-05000382 | IHEP0382 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.262 | [10 empty] | 4.7 | 98
HCBLM_I001-05000383 | IHEP0383 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.5 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.764 | [10 empty] | 4.7 | 98 | [17 empty] | 15
HCBLM_I001-05000384 | IHEP0384 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.44 | [10 empty] | 4.7 | 98
HCBLM_I001-05000385 | IHEP0385 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Koshelev |  |  |  |  | 292.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Tikhonov |  |  |  |  | 0.3 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.528 | [10 empty] | 4.7 | 98
HCBLM_I001-05000386 | IHEP0386 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Koshelev |  |  |  |  | 245.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Seleznev |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.532 | [10 empty] | 4.7 | 98
HCBLM_I001-05000387 | IHEP0387 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.918 | [10 empty] | 4.7 | 98
HCBLM_I001-05000388 | IHEP0388 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.554 | [10 empty] | 4.7 | 98
HCBLM_I001-05000389 | IHEP0389 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.95 | [10 empty] | 4.7 | 98
HCBLM_I001-05000390 | IHEP0390 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.612 | [10 empty] | 4.7 | 98
HCBLM_I001-05000391 | IHEP0391 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-10 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.84 | [10 empty] | 4.7 | 98
HCBLM_I001-05000392 | IHEP0392 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.3 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.332 | [10 empty] | 4.7 | 98
HCBLM_I001-05000393 | IHEP0393 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 24 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.65 | [10 empty] | 4.7 | 98
HCBLM_I001-05000394 | IHEP0394 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.4 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.41 | [10 empty] | 4.7 | 98
HCBLM_I001-05000395 | IHEP0395 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Koshelev |  |  |  |  | 1424.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.994 | [10 empty] | 4.7 | 98
HCBLM_I001-05000396 | IHEP0396 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-28 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-30 | Larionov |  |  |  |  | 12 | 1 | lar012_1;lar012_2 | 28 | 1127 | 220 | 6 | lar012_3;lar012_4 | Done | Ok | 3 | 2006-07-03 | Koshelev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.766 | [10 empty] | 4.7 | 98
HCBLM_I001-05000397 | IHEP0397 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-30 | Koshelev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.492 | [10 empty] | 4.7 | 98
HCBLM_I001-05000398 | IHEP0398 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.73 | [10 empty] | 4.7 | 98
HCBLM_I001-05000399 | IHEP0399 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-06-29 | Koshelev |  |  |  |  | 1103.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.658 | [10 empty] | 4.7 | 98
HCBLM_I001-05000400 | IHEP0400 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.364 | [10 empty] | 4.7 | 98
HCBLM_I001-05000401 | IHEP0401 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.3 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.06 | [10 empty] | 4.7 | 98
HCBLM_I001-05000402 | IHEP0402 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.882 | [10 empty] | 4.7 | 98
HCBLM_I001-05000403 | IHEP0403 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.588 | [10 empty] | 4.7 | 98
HCBLM_I001-05000404 | IHEP0404 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.3 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.82 | [10 empty] | 4.7 | 98
HCBLM_I001-05000405 | IHEP0405 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.266 | [10 empty] | 4.7 | 98
HCBLM_I001-05000406 | IHEP0406 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.868 | [10 empty] | 4.7 | 98
HCBLM_I001-05000407 | IHEP0407 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 72.773 | [10 empty] | 4.7 | 98
HCBLM_I001-05000408 | IHEP0408 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.878 | [10 empty] | 4.7 | 98
HCBLM_I001-05000409 | IHEP0409 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.592 | [10 empty] | 4.7 | 98
HCBLM_I001-05000410 | IHEP0410 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Koshelev |  |  |  |  | 133.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.428 | [10 empty] | 4.7 | 98
HCBLM_I001-05000411 | IHEP0411 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.824 | [10 empty] | 4.7 | 98
HCBLM_I001-05000412 | IHEP0412 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.01 | [10 empty] | 4.7 | 98
HCBLM_I001-05000413 | IHEP0413 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.774 | [10 empty] | 4.7 | 98
HCBLM_I001-05000414 | IHEP0414 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.41 | [10 empty] | 4.7 | 98 | [17 empty] | 14
HCBLM_I001-05000415 | IHEP0415 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 4.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.4 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.034 | [10 empty] | 4.7 | 98
HCBLM_I001-05000416 | IHEP0416 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.1 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.786 | [10 empty] | 4.7 | 98
HCBLM_I001-05000417 | IHEP0417 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 0.1 | UK | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 73.065 | [10 empty] | 4.7 | 98
HCBLM_I001-05000418 | IHEP0418 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.1 | [10 empty] | 4.7 | 98
HCBLM_I001-05000419 | IHEP0419 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.2 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.52 | [10 empty] | 4.7 | 98
HCBLM_I001-05000420 | IHEP0420 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.2 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.798 | [10 empty] | 4.7 | 98
HCBLM_I001-05000421 | IHEP0421 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.328 | [10 empty] | 4.7 | 98
HCBLM_I001-05000422 | IHEP0422 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 0.2 | UK | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.72 | [10 empty] | 4.7 | 98
HCBLM_I001-05000423 | IHEP0423 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.7 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.9 | [10 empty] | 4.7 | 98
HCBLM_I001-05000424 | IHEP0424 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.3 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.676 | [10 empty] | 4.7 | 98
HCBLM_I001-05000425 | IHEP0425 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.036 | [10 empty] | 4.7 | 98
HCBLM_I001-05000426 | IHEP0426 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.3 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.548 | [10 empty] | 4.7 | 98
HCBLM_I001-05000427 | IHEP0427 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.7 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.454 | [10 empty] | 4.7 | 98
HCBLM_I001-05000428 | IHEP0428 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.544 | [10 empty] | 4.7 | 98
HCBLM_I001-05000429 | IHEP0429 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.4 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.92 | [10 empty] | 4.7 | 98
HCBLM_I001-05000430 | IHEP0430 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 4.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000431 | IHEP0431 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.338 | [10 empty] | 4.7 | 98
HCBLM_I001-05000432 | IHEP0432 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.58 | [10 empty] | 4.7 | 98
HCBLM_I001-05000433 | IHEP0433 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.1 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.486 | [10 empty] | 4.7 | 98
HCBLM_I001-05000434 | IHEP0434 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.7 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.564 | [10 empty] | 4.7 | 98
HCBLM_I001-05000435 | IHEP0435 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.4 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.426 | [10 empty] | 4.7 | 98
HCBLM_I001-05000436 | IHEP0436 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.264 | [10 empty] | 4.7 | 98
HCBLM_I001-05000437 | IHEP0437 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.3 | [10 empty] | 4.7 | 98
HCBLM_I001-05000438 | IHEP0438 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Koshelev |  |  |  |  | 132.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-07 | Tereshin |  |  |  |  | 13 | 1 | ter013_1;ter013_2 | 23 | 1116 | 220 | 6 | ter013_3;ter013_4 | Done | Ok | 3 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.852 | [10 empty] | 4.7 | 98
HCBLM_I001-05000439 | IHEP0439 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-03 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.2 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.336 | [10 empty] | 4.7 | 98
HCBLM_I001-05000440 | IHEP0440 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.482 | [10 empty] | 4.7 | 98 | [17 empty] | 13
HCBLM_I001-05000441 | IHEP0441 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.904 | [10 empty] | 4.7 | 98
HCBLM_I001-05000442 | IHEP0442 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.164 | [10 empty] | 4.7 | 98
HCBLM_I001-05000443 | IHEP0443 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.338 | [10 empty] | 4.7 | 98
HCBLM_I001-05000444 | IHEP0444 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-10 | Tikhonov |  |  |  |  | 62 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.042 | [10 empty] | 4.7 | 98
HCBLM_I001-05000445 | IHEP0445 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.408 | [10 empty] | 4.7 | 98
HCBLM_I001-05000446 | IHEP0446 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.4 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 72.12 | [10 empty] | 4.7 | 98
HCBLM_I001-05000447 | IHEP0447 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Seleznev |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.64 | [10 empty] | 4.7 | 98
HCBLM_I001-05000448 | IHEP0448 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 34.1 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.4 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.072 | [10 empty] | 4.7 | 98
HCBLM_I001-05000449 | IHEP0449 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.984 | [10 empty] | 4.7 | 98
HCBLM_I001-05000450 | IHEP0450 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.144 | [10 empty] | 4.7 | 98
HCBLM_I001-05000451 | IHEP0451 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-10 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.264 | [10 empty] | 4.7 | 98
HCBLM_I001-05000452 | IHEP0452 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Mishagin |  |  |  |  | 1200.8 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.38 | [10 empty] | 4.7 | 98
HCBLM_I001-05000453 | IHEP0453 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-10 | Koshelev |  |  |  |  | 9.2 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.1 | [10 empty] | 4.7 | 98
HCBLM_I001-05000454 | IHEP0454 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.54 | [10 empty] | 4.7 | 98
HCBLM_I001-05000455 | IHEP0455 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-10 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.1 | [10 empty] | 4.7 | 98
HCBLM_I001-05000456 | IHEP0456 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Tikhonov |  |  |  |  | 0.5 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.816 | [10 empty] | 4.7 | 98
HCBLM_I001-05000457 | IHEP0457 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 13 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.6 | [10 empty] | 4.7 | 98
HCBLM_I001-05000458 | IHEP0458 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Mishagin |  |  |  |  | 20 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.528 | [10 empty] | 4.7 | 98
HCBLM_I001-05000459 | IHEP0459 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.614 | [10 empty] | 4.7 | 98
HCBLM_I001-05000460 | IHEP0460 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-06 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.818 | [10 empty] | 4.7 | 98 | [17 empty] | 3
HCBLM_I001-05000461 | IHEP0461 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.2 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.952 | [10 empty] | 4.7 | 98
HCBLM_I001-05000462 | IHEP0462 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Koshelev |  |  |  |  | 15.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.21 | [10 empty] | 4.7 | 98
HCBLM_I001-05000463 | IHEP0463 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Koshelev |  |  |  |  | 775.2 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.648 | [10 empty] | 4.7 | 98
HCBLM_I001-05000464 | IHEP0464 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.1 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.456 | [10 empty] | 4.7 | 98
HCBLM_I001-05000465 | IHEP0465 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.83 | [10 empty] | 4.7 | 98
HCBLM_I001-05000466 | IHEP0466 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.46 | [10 empty] | 4.7 | 98
HCBLM_I001-05000467 | IHEP0467 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.832 | [10 empty] | 4.7 | 98
HCBLM_I001-05000468 | IHEP0468 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.2 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.998 | [10 empty] | 4.7 | 98 | [17 empty] | 10
HCBLM_I001-05000469 | IHEP0469 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.434 | [10 empty] | 4.7 | 98
HCBLM_I001-05000470 | IHEP0470 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Koshelev |  |  |  |  | 32.1 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.526 | [10 empty] | 4.7 | 98
HCBLM_I001-05000471 | IHEP0471 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 10.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.5 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05000472 | IHEP0472 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.176 | [10 empty] | 4.7 | 98
HCBLM_I001-05000473 | IHEP0473 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.828 | [10 empty] | 4.7 | 98
HCBLM_I001-05000474 | IHEP0474 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-10 | Koshelev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.884 | [10 empty] | 4.7 | 98
HCBLM_I001-05000475 | IHEP0475 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Koshelev |  |  |  |  | 37.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.136 | [10 empty] | 4.7 | 98
HCBLM_I001-05000476 | IHEP0476 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-10 | Seleznev |  |  |  |  | 0.4 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.502 | [10 empty] | 4.7 | 98
HCBLM_I001-05000477 | IHEP0477 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.212 | [10 empty] | 4.7 | 98
HCBLM_I001-05000478 | IHEP0478 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-10 | Koshelev |  |  |  |  | 1233.3 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-12 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.624 | [10 empty] | 4.7 | 98 | [17 empty] | 6
HCBLM_I001-05000479 | IHEP0479 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-05 | Seleznev |  |  |  |  | 3.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-07 | Tereshin |  |  |  |  | 14 | 1 | ter014_1;ter014_2 | 24 | 1116 | 220 | 6 | ter014_3;ter014_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.856 | [10 empty] | 4.7 | 98
HCBLM_I001-05000480 | IHEP0480 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 41.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.008 | [10 empty] | 4.7 | 98
HCBLM_I001-05000481 | IHEP0481 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Mishagin |  |  |  |  | 1.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.126 | [10 empty] | 4.7 | 98
HCBLM_I001-05000482 | IHEP0482 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 1.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.4 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.832 | [10 empty] | 4.7 | 98
HCBLM_I001-05000483 | IHEP0483 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.152 | [10 empty] | 4.7 | 98
HCBLM_I001-05000484 | IHEP0484 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.632 | [10 empty] | 4.7 | 98
HCBLM_I001-05000485 | IHEP0485 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Mishagin |  |  |  |  | 25.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.202 | [10 empty] | 4.7 | 98
HCBLM_I001-05000486 | IHEP0486 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 1214 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.8 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.9 | [10 empty] | 4.7 | 98
HCBLM_I001-05000487 | IHEP0487 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Koshelev |  |  |  |  | 322.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.846 | [10 empty] | 4.7 | 98
HCBLM_I001-05000488 | IHEP0488 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Koshelev |  |  |  |  | 45.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.272 | [10 empty] | 4.7 | 98
HCBLM_I001-05000489 | IHEP0489 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 15.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.8 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.522 | [10 empty] | 4.7 | 98
HCBLM_I001-05000490 | IHEP0490 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 65.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.7 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.056 | [10 empty] | 4.7 | 98
HCBLM_I001-05000491 | IHEP0491 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 59 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.75 | [10 empty] | 4.7 | 98
HCBLM_I001-05000492 | IHEP0492 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Koshelev |  |  |  |  | 62.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.18 | [10 empty] | 4.7 | 98
HCBLM_I001-05000493 | IHEP0493 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Koshelev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.02 | [10 empty] | 4.7 | 98
HCBLM_I001-05000494 | IHEP0494 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 19 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.694 | [10 empty] | 4.7 | 98
HCBLM_I001-05000495 | IHEP0495 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 55 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.908 | [10 empty] | 4.7 | 98
HCBLM_I001-05000496 | IHEP0496 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 28.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.16 | [10 empty] | 4.7 | 98
HCBLM_I001-05000497 | IHEP0497 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 32 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 67.842 | [10 empty] | 4.7 | 98
HCBLM_I001-05000498 | IHEP0498 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Koshelev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.4 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.298 | [10 empty] | 4.7 | 98
HCBLM_I001-05000499 | IHEP0499 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 1.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.946 | [10 empty] | 4.7 | 98
HCBLM_I001-05000500 | IHEP0500 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 78 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.138 | [10 empty] | 4.7 | 98
HCBLM_I001-05000501 | IHEP0501 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 687 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.4 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.688 | [10 empty] | 4.7 | 98
HCBLM_I001-05000502 | IHEP0502 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 19 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.182 | [10 empty] | 4.7 | 98 | [17 empty] | 11
HCBLM_I001-05000503 | IHEP0503 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 48 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.4 | [10 empty] | 4.7 | 98
HCBLM_I001-05000504 | IHEP0504 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Koshelev |  |  |  |  | 275.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.18 | [10 empty] | 4.7 | 98
HCBLM_I001-05000505 | IHEP0505 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-12 | Tikhonov |  |  |  |  | 111.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 72.492 | [10 empty] | 4.7 | 98
HCBLM_I001-05000506 | IHEP0506 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Koshelev |  |  |  |  | 32 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.44 | [10 empty] | 4.7 | 98
HCBLM_I001-05000507 | IHEP0507 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Seleznev |  |  |  |  | 185 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.39 | [10 empty] | 4.7 | 98
HCBLM_I001-05000508 | IHEP0508 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.092 | [10 empty] | 4.7 | 98
HCBLM_I001-05000509 | IHEP0509 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Seleznev |  |  |  |  | 139 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.004 | [10 empty] | 4.7 | 98
HCBLM_I001-05000510 | IHEP0510 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.3 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 68.474 | [10 empty] | 4.7 | 98
HCBLM_I001-05000511 | IHEP0511 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Koshelev |  |  |  |  | 22.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.8 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.258 | [10 empty] | 4.7 | 98
HCBLM_I001-05000512 | IHEP0512 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.394 | [10 empty] | 4.7 | 98
HCBLM_I001-05000513 | IHEP0513 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Seleznev |  |  |  |  | 216 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.7 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.9
HCBLM_I001-05000514 | IHEP0514 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.272 | [10 empty] | 4.7 | 98
HCBLM_I001-05000515 | IHEP0515 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.642 | [10 empty] | 4.7 | 98
HCBLM_I001-05000516 | IHEP0516 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2
HCBLM_I001-05000517 | IHEP0517 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 1.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.824 | [10 empty] | 4.7 | 98
HCBLM_I001-05000518 | IHEP0518 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 70.666 | [10 empty] | 4.7 | 98
HCBLM_I001-05000519 | IHEP0519 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Tikhonov |  |  |  |  | 2.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.186 | [10 empty] | 4.7 | 98
HCBLM_I001-05000520 | IHEP0520 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.1 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 71.592 | [10 empty] | 4.7 | 98
HCBLM_I001-05000521 | IHEP0521 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-19 | Mishagin |  |  |  |  | 1500 | 69.286 | [10 empty] | 4.7 | 98
HCBLM_I001-05000522 | IHEP0522 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Seleznev |  |  |  |  | 3.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.1 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.914 | [10 empty] | 4.7 | 98
HCBLM_I001-05000523 | IHEP0523 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.404 | [10 empty] | 4.7 | 98
HCBLM_I001-05000524 | IHEP0524 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.842 | [10 empty] | 4.7 | 98
HCBLM_I001-05000525 | IHEP0525 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Koshelev |  |  |  |  | 23.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.1 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.724 | [10 empty] | 4.7 | 98
HCBLM_I001-05000526 | IHEP0526 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Koshelev |  |  |  |  | 2.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.602 | [10 empty] | 4.7 | 98
HCBLM_I001-05000527 | IHEP0527 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Tikhonov |  |  |  |  | 0.1 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.454 | [10 empty] | 4.7 | 98
HCBLM_I001-05000528 | IHEP0528 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Koshelev |  |  |  |  | 32.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Tikhonov |  |  |  |  | 0.1 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.952 | [10 empty] | 4.7 | 98
HCBLM_I001-05000529 | IHEP0529 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Seleznev |  |  |  |  | 40 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.2 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.248 | [10 empty] | 4.7 | 98
HCBLM_I001-05000530 | IHEP0530 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Mishagin |  |  |  |  | 50 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.4 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.656 | [10 empty] | 4.7 | 98
HCBLM_I001-05000531 | IHEP0531 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Mishagin |  |  |  |  | 25 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-06-23 | Ferioli |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.54 | [10 empty] | 4.7 | 98
HCBLM_I001-05000532 | IHEP0532 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Seleznev |  |  |  |  | 30 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.102 | [10 empty] | 4.7 | 98
HCBLM_I001-05000533 | IHEP0533 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Seleznev |  |  |  |  | 39 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.176 | [10 empty] | 4.7 | 98
HCBLM_I001-05000534 | IHEP0534 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 1.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.89 | [10 empty] | 4.7 | 98
HCBLM_I001-05000535 | IHEP0535 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Mishagin |  |  |  |  | 1.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.952 | [10 empty] | 4.7 | 98
HCBLM_I001-05000536 | IHEP0536 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 1.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.508 | [10 empty] | 4.7 | 98
HCBLM_I001-05000537 | IHEP0537 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-14 | Tikhonov |  |  |  |  | 1.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-18 | Tereshin |  |  |  |  | 17 | 1 | ter017_1;ter017_2 | 26 | 1117 | 220 | 6 | ter017_3;ter017_4 | Done | Ok | 3 | 2006-07-17 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.638 | [10 empty] | 4.7 | 98
HCBLM_I001-05000538 | IHEP0538 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Mishagin | [5 empty] | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.272 | [10 empty] | 4.7 | 98
HCBLM_I001-05000539 | IHEP0539 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Seleznev |  |  |  |  | 1700 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.76 | [10 empty] | 4.7 | 98
HCBLM_I001-05000540 | IHEP0540 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Seleznev |  |  |  |  | 43 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.214 | [10 empty] | 4.7 | 98
HCBLM_I001-05000541 | IHEP0541 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-13 | Seleznev |  |  |  |  | 650 | IHEP | 2000 | Done | Ok | 2 | 2006-07-16 | Tereshin |  |  |  |  | 16 | 1 | ter016_1;ter016_2 | 25 | 1117 | 220 | 6 | ter016_3;ter016_4 | Done | Ok | 3 | 2006-07-17 | Mishagin |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.336 | [10 empty] | 4.7 | 98
HCBLM_I001-05000542 | IHEP0542 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.508 | [10 empty] | 4.7 | 98
HCBLM_I001-05000543 | IHEP0543 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.274 | [10 empty] | 4.7 | 98
HCBLM_I001-05000544 | IHEP0544 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.228 | [10 empty] | 4.7 | 98
HCBLM_I001-05000545 | IHEP0545 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 1.8 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.5 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.35 | [10 empty] | 4.7 | 98
HCBLM_I001-05000546 | IHEP0546 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 12140 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.888 | [10 empty] | 4.7 | 98
HCBLM_I001-05000547 | IHEP0547 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.3 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.83 | [10 empty] | 4.7 | 98
HCBLM_I001-05000548 | IHEP0548 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.842 | [10 empty] | 4.7 | 98
HCBLM_I001-05000549 | IHEP0549 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 1.4 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.494 | [10 empty] | 4.7 | 98
HCBLM_I001-05000550 | IHEP0550 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.044 | [10 empty] | 4.7 | 98
HCBLM_I001-05000551 | IHEP0551 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.33 | [10 empty] | 4.7 | 98
HCBLM_I001-05000552 | IHEP0552 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 2.1 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.7 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.024 | [10 empty] | 4.7 | 98
HCBLM_I001-05000553 | IHEP0553 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.91 | [10 empty] | 4.7 | 98
HCBLM_I001-05000554 | IHEP0554 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-18 | Tikhonov |  |  |  |  | 1.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.61 | [10 empty] | 4.7 | 98
HCBLM_I001-05000555 | IHEP0555 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Larionov |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.632 | [10 empty] | 4.7 | 98
HCBLM_I001-05000556 | IHEP0556 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.628 | [10 empty] | 4.7 | 98
HCBLM_I001-05000557 | IHEP0557 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.86 | [10 empty] | 4.7 | 98
HCBLM_I001-05000558 | IHEP0558 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.578 | [10 empty] | 4.7 | 98
HCBLM_I001-05000559 | IHEP0559 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-18 | Tikhonov |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.402 | [10 empty] | 4.7 | 98
HCBLM_I001-05000560 | IHEP0560 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.926 | [10 empty] | 4.7 | 98
HCBLM_I001-05000561 | IHEP0561 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.726 | [10 empty] | 4.7 | 98
HCBLM_I001-05000562 | IHEP0562 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.8 | [10 empty] | 4.7 | 98
HCBLM_I001-05000563 | IHEP0563 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-18 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.8 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.744 | [10 empty] | 4.7 | 98
HCBLM_I001-05000564 | IHEP0564 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.564 | [10 empty] | 4.7 | 98
HCBLM_I001-05000565 | IHEP0565 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05000566 | IHEP0566 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.314 | [10 empty] | 4.7 | 98
HCBLM_I001-05000567 | IHEP0567 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 12 | UK | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.348 | [10 empty] | 4.7 | 98
HCBLM_I001-05000568 | IHEP0568 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Tikhonov |  |  |  |  | 1.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.036 | [10 empty] | 4.7 | 98
HCBLM_I001-05000569 | IHEP0569 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-18 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.592 | [10 empty] | 4.7 | 98
HCBLM_I001-05000570 | IHEP0570 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Koshelev |  |  |  |  | 1297.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.208 | [10 empty] | 4.7 | 98
HCBLM_I001-05000571 | IHEP0571 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.982 | [10 empty] | 4.7 | 98
HCBLM_I001-05000572 | IHEP0572 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Tikhonov |  |  |  |  | 1.1 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.746 | [10 empty] | 4.7 | 98
HCBLM_I001-05000573 | IHEP0573 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Tikhonov |  |  |  |  | 1.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.312 | [10 empty] | 4.7 | 98
HCBLM_I001-05000574 | IHEP0574 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.772 | [10 empty] | 4.7 | 98
HCBLM_I001-05000575 | IHEP0575 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Tikhonov |  |  |  |  | 1.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.8 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.072 | [10 empty] | 4.7 | 98
HCBLM_I001-05000576 | IHEP0576 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 1 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.388 | [10 empty] | 4.7 | 98
HCBLM_I001-05000577 | IHEP0577 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.506 | [10 empty] | 4.7 | 98
HCBLM_I001-05000578 | IHEP0578 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-28 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.128 | [10 empty] | 4.7 | 98
HCBLM_I001-05000579 | IHEP0579 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.14 | [10 empty] | 4.7 | 98
HCBLM_I001-05000580 | IHEP0580 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.608 | [10 empty] | 4.7 | 98
HCBLM_I001-05000581 | IHEP0581 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.572 | [10 empty] | 4.7 | 98
HCBLM_I001-05000582 | IHEP0582 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.742 | [10 empty] | 4.7 | 98
HCBLM_I001-05000583 | IHEP0583 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.666 | [10 empty] | 4.7 | 98
HCBLM_I001-05000584 | IHEP0584 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.846 | [10 empty] | 4.7 | 98
HCBLM_I001-05000585 | IHEP0585 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.63 | [10 empty] | 4.7 | 98
HCBLM_I001-05000586 | IHEP0586 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.934 | [10 empty] | 4.7 | 98
HCBLM_I001-05000587 | IHEP0587 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.31 | [10 empty] | 4.7 | 98
HCBLM_I001-05000588 | IHEP0588 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.178 | [10 empty] | 4.7 | 98
HCBLM_I001-05000589 | IHEP0589 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 1239 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.026 | [10 empty] | 4.7 | 98
HCBLM_I001-05000590 | IHEP0590 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Tikhonov |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.052 | [10 empty] | 4.7 | 98
HCBLM_I001-05000591 | IHEP0591 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 4.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.318 | [10 empty] | 4.7 | 98
HCBLM_I001-05000592 | IHEP0592 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.862 | [10 empty] | 4.7 | 98
HCBLM_I001-05000593 | IHEP0593 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Tikhonov |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.708 | [10 empty] | 4.7 | 98
HCBLM_I001-05000594 | IHEP0594 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.276 | [10 empty] | 4.7 | 98
HCBLM_I001-05000595 | IHEP0595 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.6 | 0.6 | 2000 | [23 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.228 | [10 empty] | 4.7 | 98
HCBLM_I001-05000596 | IHEP0596 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.356 | [10 empty] | 4.7 | 98
HCBLM_I001-05000597 | IHEP0597 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.48 | [10 empty] | 4.7 | 98
HCBLM_I001-05000598 | IHEP0598 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.942 | [10 empty] | 4.7 | 98
HCBLM_I001-05000599 | IHEP0599 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-25 | Seleznev |  |  |  |  | 0.5 | 0.3 | 2000 | [23 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.602 | [10 empty] | 4.7 | 98
HCBLM_I001-05000600 | IHEP0600 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 74.478 | [10 empty] | 4.7 | 98
HCBLM_I001-05000601 | IHEP0601 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.642 | [10 empty] | 4.7 | 98
HCBLM_I001-05000602 | IHEP0602 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.434 | [10 empty] | 4.7 | 98
HCBLM_I001-05000603 | IHEP0603 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.16 | [10 empty] | 4.7 | 98
HCBLM_I001-05000604 | IHEP0604 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.164 | [10 empty] | 4.7 | 98
HCBLM_I001-05000605 | IHEP0605 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Tikhonov |  |  |  |  | 1.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.962 | [10 empty] | 4.7 | 98
HCBLM_I001-05000606 | IHEP0606 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.022 | [10 empty] | 4.7 | 98
HCBLM_I001-05000607 | IHEP0607 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.1 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.746 | [10 empty] | 4.7 | 98
HCBLM_I001-05000608 | IHEP0608 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Tikhonov |  |  |  |  | 1.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.418 | [10 empty] | 4.7 | 98
HCBLM_I001-05000609 | IHEP0609 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.622 | [10 empty] | 4.7 | 98
HCBLM_I001-05000610 | IHEP0610 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-21 | Tereshin |  |  |  |  | 18 | 1 | ter018_1;ter018_2 | 24 | 1116 | 220 | 6 | ter018_3;ter018_4 | Done | Ok | 3 | 2006-07-24 | Koshelev |  |  |  |  | 0.1 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.618 | [10 empty] | 4.7 | 98
HCBLM_I001-05000611 | IHEP0611 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.256 | [10 empty] | 4.7 | 98
HCBLM_I001-05000612 | IHEP0612 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Tikhonov |  |  |  |  | 1.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.156 | [10 empty] | 4.7 | 98
HCBLM_I001-05000613 | IHEP0613 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-20 | Tikhonov |  |  |  |  | 1232 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.1 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.9 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.186 | [10 empty] | 4.7 | 98
HCBLM_I001-05000614 | IHEP0614 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-19 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.956 | [10 empty] | 4.7 | 98
HCBLM_I001-05000615 | IHEP0615 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.3 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.49 | [10 empty] | 4.7 | 98
HCBLM_I001-05000616 | IHEP0616 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.2 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.2 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.704 | [10 empty] | 4.7 | 98
HCBLM_I001-05000617 | IHEP0617 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.4 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.638 | [10 empty] | 4.7 | 98
HCBLM_I001-05000618 | IHEP0618 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.01 | [10 empty] | 4.7 | 98
HCBLM_I001-05000619 | IHEP0619 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.658 | [10 empty] | 4.7 | 98
HCBLM_I001-05000620 | IHEP0620 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.446 | [10 empty] | 4.7 | 98
HCBLM_I001-05000621 | IHEP0621 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.5 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.07 | [10 empty] | 4.7 | 98
HCBLM_I001-05000622 | IHEP0622 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.228 | [10 empty] | 4.7 | 98
HCBLM_I001-05000623 | IHEP0623 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.792 | [10 empty] | 4.7 | 98
HCBLM_I001-05000624 | IHEP0624 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.768 | [10 empty] | 4.7 | 98
HCBLM_I001-05000625 | IHEP0625 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 196 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.684 | [10 empty] | 4.7 | 98
HCBLM_I001-05000626 | IHEP0626 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.11 | [10 empty] | 4.7 | 98
HCBLM_I001-05000627 | IHEP0627 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.666 | [10 empty] | 4.7 | 98
HCBLM_I001-05000628 | IHEP0628 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.03 | [10 empty] | 4.7 | 98
HCBLM_I001-05000629 | IHEP0629 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.076 | [10 empty] | 4.7 | 98
HCBLM_I001-05000630 | IHEP0630 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 12 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.878 | [10 empty] | 4.7 | 98
HCBLM_I001-05000631 | IHEP0631 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.09 | [10 empty] | 4.7 | 98
HCBLM_I001-05000632 | IHEP0632 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.76 | [10 empty] | 4.7 | 98
HCBLM_I001-05000633 | IHEP0633 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Koshelev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Seleznev |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.584 | [10 empty] | 4.7 | 98
HCBLM_I001-05000634 | IHEP0634 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Mishagin |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.344 | [10 empty] | 4.7 | 98
HCBLM_I001-05000635 | IHEP0635 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.2 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.655 | [10 empty] | 4.7 | 98
HCBLM_I001-05000636 | IHEP0636 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Larionov |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.43 | [10 empty] | 4.7 | 98
HCBLM_I001-05000637 | IHEP0637 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.04 | [10 empty] | 4.7 | 98
HCBLM_I001-05000638 | IHEP0638 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.764 | [10 empty] | 4.7 | 98
HCBLM_I001-05000639 | IHEP0639 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.4 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.03 | [10 empty] | 4.7 | 98
HCBLM_I001-05000640 | IHEP0640 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 4.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 67.864 | [10 empty] | 4.7 | 98
HCBLM_I001-05000641 | IHEP0641 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.248 | [10 empty] | 4.7 | 98
HCBLM_I001-05000642 | IHEP0642 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.324 | [10 empty] | 4.7 | 98
HCBLM_I001-05000643 | IHEP0643 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.648 | [10 empty] | 4.7 | 98
HCBLM_I001-05000644 | IHEP0644 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.204 | [10 empty] | 4.7 | 98
HCBLM_I001-05000645 | IHEP0645 | IHEP | Manufacturing |  | Done | Ok | 1 |  | Mishagin |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-28 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.97 | [10 empty] | 4.7 | 98
HCBLM_I001-05000646 | IHEP0646 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-21 | Koshelev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-23 | Tereshin |  |  |  |  | 19 | 1 | ter019_1;ter019_2 | 23 | 1116 | 220 | 6 | ter019_3;ter019_4 | Done | Ok | 3 | 2006-07-24 | Tikhonov |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.87 | [10 empty] | 4.7 | 98
HCBLM_I001-05000647 | IHEP0647 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.416 | [10 empty] | 4.7 | 98
HCBLM_I001-05000648 | IHEP0648 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.488 | [10 empty] | 4.7 | 98
HCBLM_I001-05000649 | IHEP0649 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.294 | [10 empty] | 4.7 | 98
HCBLM_I001-05000650 | IHEP0650 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 77.569 | [10 empty] | 4.7 | 98
HCBLM_I001-05000651 | IHEP0651 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-24 | Seleznev |  |  |  |  | 1 | UK | 2000 | Done | Ok | 2 | 2006-07-27 | Tereshin |  |  |  |  | 20 | 2 | ter020_1;ter020_2 | 24 | 1115 | 220 | 6 | ter020_3;ter020_4 | Done | Ok | 3 | 2006-07-27 | Mishagin |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.924 | [10 empty] | 4.7 | 98
HCBLM_I001-05000652 | IHEP0652 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.242 | [10 empty] | 4.7 | 98
HCBLM_I001-05000653 | IHEP0653 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.9 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.188 | [10 empty] | 4.7 | 98
HCBLM_I001-05000654 | IHEP0654 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Koshelev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.5 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.164 | [10 empty] | 4.7 | 98
HCBLM_I001-05000655 | IHEP0655 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-25 | Seleznev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.29 | [10 empty] | 4.7 | 98
HCBLM_I001-05000656 | IHEP0656 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.6 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.088 | [10 empty] | 4.7 | 98
HCBLM_I001-05000657 | IHEP0657 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Mishagin |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 73 | [10 empty] | 4.7 | 98
HCBLM_I001-05000658 | IHEP0658 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Tikhonov |  |  |  |  | 2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.1 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.046 | [10 empty] | 4.7 | 98
HCBLM_I001-05000659 | IHEP0659 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Tikhonov |  |  |  |  | 12 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.4 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.324 | [10 empty] | 4.7 | 98
HCBLM_I001-05000660 | IHEP0660 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.03 | [10 empty] | 4.7 | 98
HCBLM_I001-05000661 | IHEP0661 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-27 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.446 | [10 empty] | 4.7 | 98
HCBLM_I001-05000662 | IHEP0662 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.8 | UK | 2000 | Cancelled | Not Ok | 2 | 2006-02-08 | [5 empty] | 22 | 2 | [6 empty] | Done | Ok | 3 | 2006-08-03 | Seleznev |  |  |  |  | 0.8 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.13 | [10 empty] | 4.7 | 98
HCBLM_I001-05000663 | IHEP0663 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 44 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71 | [10 empty] | 4.7 | 98
HCBLM_I001-05000664 | IHEP0664 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.888 | [10 empty] | 4.7 | 98
HCBLM_I001-05000665 | IHEP0665 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.88 | [10 empty] | 4.7 | 98
HCBLM_I001-05000666 | IHEP0666 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Mishagin |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.946 | [10 empty] | 4.7 | 98
HCBLM_I001-05000667 | IHEP0667 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Mishagin |  |  |  |  | 2.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.122 | [10 empty] | 4.7 | 98
HCBLM_I001-05000668 | IHEP0668 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Tikhonov |  |  |  |  | 250 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.192 | [10 empty] | 4.7 | 98
HCBLM_I001-05000669 | IHEP0669 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Tikhonov |  |  |  |  | 1200 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.5 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.33 | [10 empty] | 4.7 | 98
HCBLM_I001-05000670 | IHEP0670 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Tikhonov |  |  |  |  | 27 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 73.36 | [10 empty] | 4.7 | 98
HCBLM_I001-05000671 | IHEP0671 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 56 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.2 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.174 | [10 empty] | 4.7 | 98
HCBLM_I001-05000672 | IHEP0672 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.218 | [10 empty] | 4.7 | 98
HCBLM_I001-05000673 | IHEP0673 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.374 | [10 empty] | 4.7 | 98
HCBLM_I001-05000674 | IHEP0674 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Mishagin |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.98 | [10 empty] | 4.7 | 98
HCBLM_I001-05000675 | IHEP0675 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.086 | [10 empty] | 4.7 | 98
HCBLM_I001-05000676 | IHEP0676 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Mishagin |  |  |  |  | 180 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.6 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.488 | [10 empty] | 4.7 | 98
HCBLM_I001-05000677 | IHEP0677 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.684 | [10 empty] | 4.7 | 98
HCBLM_I001-05000678 | IHEP0678 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-27 | Seleznev |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.016 | [10 empty] | 4.7 | 98
HCBLM_I001-05000679 | IHEP0679 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.188 | [10 empty] | 4.7 | 98
HCBLM_I001-05000680 | IHEP0680 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.5 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.752 | [10 empty] | 4.7 | 98
HCBLM_I001-05000681 | IHEP0681 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Mishagin |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.5 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.582 | [10 empty] | 4.7 | 98
HCBLM_I001-05000682 | IHEP0682 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Tikhonov |  |  |  |  | 25 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 73.426 | [10 empty] | 4.7 | 98
HCBLM_I001-05000683 | IHEP0683 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.182 | [10 empty] | 4.7 | 98
HCBLM_I001-05000684 | IHEP0684 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05000685 | IHEP0685 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Tikhonov |  |  |  |  | 12 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.5 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.512 | [10 empty] | 4.7 | 98
HCBLM_I001-05000686 | IHEP0686 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Seleznev |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.47 | [10 empty] | 4.7 | 98
HCBLM_I001-05000687 | IHEP0687 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 10 | IHEP | 2000 | Cancelled | Not Ok | 2 | 2006-02-08 | [5 empty] | 22 | [7 empty] | Done | Ok | 3 | 2006-08-03 | Seleznev |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.32 | [10 empty] | 4.7 | 98
HCBLM_I001-05000688 | IHEP0688 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-31 | Larionov |  |  |  |  | 21 | 2 | lar021_1;lar021_2 | 20 | 1107 | 220 | 6 | lar021_3;lar021_4 | Done | Ok | 3 | 2006-07-31 | Tikhonov |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.088 | [10 empty] | 4.7 | 98
HCBLM_I001-05000689 | IHEP0689 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.8 | UK | 2000 | Cancelled | Not Ok | 2 | 2006-02-08 | [5 empty] | 22 | [7 empty] | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.8 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.64 | [10 empty] | 4.7 | 98
HCBLM_I001-05000690 | IHEP0690 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.7 | UK | 2000 | Cancelled | Not Ok | 2 | 2006-02-08 | [5 empty] | 22 | [7 empty] | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.77 | [10 empty] | 4.7 | 98
HCBLM_I001-05000691 | IHEP0691 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Seleznev |  |  |  |  | 1231 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 2 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Tikhonov |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.408 | [10 empty] | 4.7 | 98
HCBLM_I001-05000692 | IHEP0692 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Seleznev |  |  |  |  | 1442 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 2 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.442 | [10 empty] | 4.7 | 98
HCBLM_I001-05000693 | IHEP0693 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Seleznev |  |  |  |  | 26 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 2 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.148 | [10 empty] | 4.7 | 98
HCBLM_I001-05000694 | IHEP0694 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Seleznev |  |  |  |  | 13.5 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 2 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.5 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.8 | [10 empty] | 4.7 | 98
HCBLM_I001-05000695 | IHEP0695 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Seleznev |  |  |  |  | 1440 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 2 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.366 | [10 empty] | 4.7 | 98
HCBLM_I001-05000696 | IHEP0696 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Seleznev |  |  |  |  | 1472 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 2 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.124 | [10 empty] | 4.7 | 98
HCBLM_I001-05000697 | IHEP0697 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 2 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.424 | [10 empty] | 4.7 | 98
HCBLM_I001-05000698 | IHEP0698 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 2 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.534 | [10 empty] | 4.7 | 98
HCBLM_I001-05000699 | IHEP0699 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-07 | Seleznev |  |  |  |  | 2.8 | IHEP | 2000 | Done | Ok | 2 | 2006-12-07 | Tereshin |  |  |  |  | 15 | 2 | ter015_1;ter015_2 | 27 | 1126 | 220 | 6 | ter015_3;ter015_4 | Done | Ok | 3 | 2006-07-13 | Mishagin |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-07-25 | Grishin |  | 2 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.176 | [10 empty] | 4.7 | 98
HCBLM_I001-05000700 | IHEP0700 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 4.5 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.936 | [10 empty] | 4.7 | 98
HCBLM_I001-05000701 | IHEP0701 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.5 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.316 | [10 empty] | 4.7 | 98
HCBLM_I001-05000702 | IHEP0702 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.818 | [10 empty] | 4.7 | 98
HCBLM_I001-05000703 | IHEP0703 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Mishagin |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.58 | [10 empty] | 4.7 | 98
HCBLM_I001-05000704 | IHEP0704 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.912 | [10 empty] | 4.7 | 98
HCBLM_I001-05000705 | IHEP0705 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.384 | [10 empty] | 4.7 | 98
HCBLM_I001-05000706 | IHEP0706 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Koshelev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.25 | [10 empty] | 4.7 | 98
HCBLM_I001-05000707 | IHEP0707 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.754 | [10 empty] | 4.7 | 98
HCBLM_I001-05000708 | IHEP0708 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.912 | [10 empty] | 4.7 | 98
HCBLM_I001-05000709 | IHEP0709 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Mishagin |  |  |  |  | 1.2 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.614 | [10 empty] | 4.7 | 98
HCBLM_I001-05000710 | IHEP0710 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.876 | [10 empty] | 4.7 | 98
HCBLM_I001-05000711 | IHEP0711 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.376 | [10 empty] | 4.7 | 98
HCBLM_I001-05000712 | IHEP0712 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Mishagin |  |  |  |  | 0.3 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.806 | [10 empty] | 4.7 | 98
HCBLM_I001-05000713 | IHEP0713 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Koshelev |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.868 | [10 empty] | 4.7 | 98
HCBLM_I001-05000714 | IHEP0714 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Koshelev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Seleznev |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.422 | [10 empty] | 4.7 | 98
HCBLM_I001-05000715 | IHEP0715 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Mishagin |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.194 | [10 empty] | 4.7 | 98
HCBLM_I001-05000716 | IHEP0716 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.994 | [10 empty] | 4.7 | 98
HCBLM_I001-05000717 | IHEP0717 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Koshelev |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.914 | [10 empty] | 4.7 | 98
HCBLM_I001-05000718 | IHEP0718 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Mishagin |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-04 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.696 | [10 empty] | 4.7 | 98
HCBLM_I001-05000719 | IHEP0719 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Mishagin |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-04 | Tikhonov |  |  |  |  | 0.9 | 0.4 | 2000 | [23 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 72.11 | [10 empty] | 4.7 | 98
HCBLM_I001-05000720 | IHEP0720 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-04 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.714 | [10 empty] | 4.7 | 98
HCBLM_I001-05000721 | IHEP0721 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Mishagin |  |  |  |  | 12 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-04 | Tikhonov |  |  |  |  | 0.7 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.92 | [10 empty] | 4.7 | 98
HCBLM_I001-05000722 | IHEP0722 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Mishagin |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-04 | Tikhonov |  |  |  |  | 0.8 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.874 | [10 empty] | 4.7 | 98
HCBLM_I001-05000723 | IHEP0723 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Mishagin |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.306 | [10 empty] | 4.7 | 98
HCBLM_I001-05000724 | IHEP0724 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-04 | Tikhonov |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.708 | [10 empty] | 4.7 | 98
HCBLM_I001-05000725 | IHEP0725 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 3.1 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-04 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.244 | [10 empty] | 4.7 | 98
HCBLM_I001-05000726 | IHEP0726 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-04 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.994 | [10 empty] | 4.7 | 98
HCBLM_I001-05000727 | IHEP0727 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-04 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.606 | [10 empty] | 4.7 | 98
HCBLM_I001-05000728 | IHEP0728 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.286 | [10 empty] | 4.7 | 98
HCBLM_I001-05000729 | IHEP0729 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Koshelev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.89 | [10 empty] | 4.7 | 98
HCBLM_I001-05000730 | IHEP0730 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-26 | Mishagin |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.032 | [10 empty] | 4.7 | 98
HCBLM_I001-05000731 | IHEP0731 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Mishagin |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.476 | [10 empty] | 4.7 | 98
HCBLM_I001-05000732 | IHEP0732 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-07-27 | Seleznev |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-02-08 | Larionov |  |  |  |  | 22 | 2 | lar022_1;lar022_2 | 21 | 1111 | 220 | 6 | lar022_3;lar022_4 | Done | Ok | 3 | 2006-08-03 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [23 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.658 | [10 empty] | 4.7 | 98
HCBLM_I001-05000733 | IHEP0733 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.15 | [10 empty] | 4.7 | 98
HCBLM_I001-05000734 | IHEP0734 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.3 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.818 | [10 empty] | 4.7 | 98
HCBLM_I001-05000735 | IHEP0735 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 1218 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.6 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.438 | [10 empty] | 4.7 | 98
HCBLM_I001-05000736 | IHEP0736 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.3 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.688 | [10 empty] | 4.7 | 98
HCBLM_I001-05000737 | IHEP0737 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.756 | [10 empty] | 4.7 | 98
HCBLM_I001-05000738 | IHEP0738 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 1.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.812 | [10 empty] | 4.7 | 98
HCBLM_I001-05000739 | IHEP0739 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [14 empty] | 2006-09-19 |  |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.642 | [10 empty] | 4.7 | 98
HCBLM_I001-05000740 | IHEP0740 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.2 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.2 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.15 | [10 empty] | 4.7 | 98
HCBLM_I001-05000741 | IHEP0741 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 6.7 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.3 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.49 | [10 empty] | 4.7 | 98
HCBLM_I001-05000742 | IHEP0742 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 1240 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.432 | [10 empty] | 4.7 | 98
HCBLM_I001-05000743 | IHEP0743 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05000744 | IHEP0744 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.084 | [10 empty] | 4.7 | 98
HCBLM_I001-05000745 | IHEP0745 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 1.1 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.028 | [10 empty] | 4.7 | 98
HCBLM_I001-05000746 | IHEP0746 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 1.4 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.462 | [10 empty] | 4.7 | 98
HCBLM_I001-05000747 | IHEP0747 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 1.2 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.2 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.668 | [10 empty] | 4.7 | 98
HCBLM_I001-05000748 | IHEP0748 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.3 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.8 | [10 empty] | 4.7 | 98
HCBLM_I001-05000749 | IHEP0749 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.1 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.748 | [10 empty] | 4.7 | 98
HCBLM_I001-05000750 | IHEP0750 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.3 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.2 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.918 | [10 empty] | 4.7 | 98
HCBLM_I001-05000751 | IHEP0751 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.38 | [10 empty] | 4.7 | 98
HCBLM_I001-05000752 | IHEP0752 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 17 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.542 | [10 empty] | 4.7 | 98
HCBLM_I001-05000753 | IHEP0753 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Seleznev |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.022 | [10 empty] | 4.7 | 98
HCBLM_I001-05000754 | IHEP0754 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 18 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.3 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.98 | [10 empty] | 4.7 | 98
HCBLM_I001-05000755 | IHEP0755 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [14 empty] | 2006-09-19 |  |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.8 | [10 empty] | 4.7 | 98
HCBLM_I001-05000756 | IHEP0756 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 69.976 | [10 empty] | 4.7 | 98
HCBLM_I001-05000757 | IHEP0757 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 12 | IHEP | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.4 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.906 | [10 empty] | 4.7 | 98
HCBLM_I001-05000758 | IHEP0758 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.84 | [10 empty] | 4.7 | 98
HCBLM_I001-05000759 | IHEP0759 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.3 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.3 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.654 | [10 empty] | 4.7 | 98
HCBLM_I001-05000760 | IHEP0760 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [14 empty] | 2006-09-19 |  |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.374 | [10 empty] | 4.7 | 98
HCBLM_I001-05000761 | IHEP0761 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.2 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [14 empty] | 2006-09-19 |  |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.066 | [10 empty] | 4.7 | 98
HCBLM_I001-05000762 | IHEP0762 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 1.3 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.9 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.506 | [10 empty] | 4.7 | 98
HCBLM_I001-05000763 | IHEP0763 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.562 | [10 empty] | 4.7 | 98
HCBLM_I001-05000764 | IHEP0764 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [14 empty] | 2006-09-19 |  |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.08 | [10 empty] | 4.7 | 98
HCBLM_I001-05000765 | IHEP0765 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.6 | 0.4 | 2000 | [14 empty] | 2006-09-19 |  |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.114 | [10 empty] | 4.7 | 98
HCBLM_I001-05000766 | IHEP0766 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.294 | [10 empty] | 4.7 | 98
HCBLM_I001-05000767 | IHEP0767 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.62 | [10 empty] | 4.7 | 98
HCBLM_I001-05000768 | IHEP0768 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-08 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.406 | [10 empty] | 4.7 | 98
HCBLM_I001-05000769 | IHEP0769 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.42 | [10 empty] | 4.7 | 98
HCBLM_I001-05000770 | IHEP0770 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Seleznev |  |  |  |  | 1.6 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.9 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.262 | [10 empty] | 4.7 | 98
HCBLM_I001-05000771 | IHEP0771 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.268 | [10 empty] | 4.7 | 98
HCBLM_I001-05000772 | IHEP0772 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.658 | [10 empty] | 4.7 | 98
HCBLM_I001-05000773 | IHEP0773 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.07 | [10 empty] | 4.7 | 98
HCBLM_I001-05000774 | IHEP0774 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.8 | [10 empty] | 4.7 | 98
HCBLM_I001-05000775 | IHEP0775 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.534 | [10 empty] | 4.7 | 98
HCBLM_I001-05000776 | IHEP0776 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.49 | [10 empty] | 4.7 | 98
HCBLM_I001-05000777 | IHEP0777 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Seleznev |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-11 | Tikhonov |  |  |  |  | 0.9 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.394 | [10 empty] | 4.7 | 98
HCBLM_I001-05000778 | IHEP0778 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-11 | Mishagin |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.702 | [10 empty] | 4.7 | 98
HCBLM_I001-05000779 | IHEP0779 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.6 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.164 | [10 empty] | 4.7 | 98
HCBLM_I001-05000780 | IHEP0780 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Seleznev |  |  |  |  | 49 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.9 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.138 | [10 empty] | 4.7 | 98
HCBLM_I001-05000781 | IHEP0781 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.8 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.259 | [10 empty] | 4.7 | 98
HCBLM_I001-05000782 | IHEP0782 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.21 | [10 empty] | 4.7 | 98
HCBLM_I001-05000783 | IHEP0783 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.3 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.068 | [10 empty] | 4.7 | 98
HCBLM_I001-05000784 | IHEP0784 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 1.3 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.404 | [10 empty] | 4.7 | 98
HCBLM_I001-05000785 | IHEP0785 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.8 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.95 | [10 empty] | 4.7 | 98
HCBLM_I001-05000786 | IHEP0786 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 450 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.398 | [10 empty] | 4.7 | 98
HCBLM_I001-05000787 | IHEP0787 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-11 | Tikhonov |  |  |  |  | 0.3 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.624 | [10 empty] | 4.7 | 98
HCBLM_I001-05000788 | IHEP0788 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.872 | [10 empty] | 4.7 | 98
HCBLM_I001-05000789 | IHEP0789 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.37 | [10 empty] | 4.7 | 98
HCBLM_I001-05000790 | IHEP0790 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.766 | [10 empty] | 4.7 | 98
HCBLM_I001-05000791 | IHEP0791 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 15 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.238 | [10 empty] | 4.7 | 98
HCBLM_I001-05000792 | IHEP0792 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.394 | [10 empty] | 4.7 | 98
HCBLM_I001-05000793 | IHEP0793 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.026 | [10 empty] | 4.7 | 98
HCBLM_I001-05000794 | IHEP0794 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 14 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.8 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.754 | [10 empty] | 4.7 | 98
HCBLM_I001-05000795 | IHEP0795 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.788 | [10 empty] | 4.7 | 98
HCBLM_I001-05000796 | IHEP0796 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 68.83 | [10 empty] | 4.7 | 98
HCBLM_I001-05000797 | IHEP0797 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.102 | [10 empty] | 4.7 | 98
HCBLM_I001-05000798 | IHEP0798 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Seleznev |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-11 | Tikhonov |  |  |  |  | 1.1 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.278 | [10 empty] | 4.7 | 98
HCBLM_I001-05000799 | IHEP0799 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Seleznev |  |  |  |  | 44 | IHEP | 2000 | Done | Ok | 2 | 2006-09-08 | Larionov |  |  |  |  | 23 | 2 | lar023_1;lar023_2 | 21 | 1110 | 220 | 6 | lar023_3;lar023_4 | Done | Ok | 3 | 2006-08-10 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.936 | [10 empty] | 4.7 | 98
HCBLM_I001-05000800 | IHEP0800 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-11 | Tikhonov |  |  |  |  | 0.3 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.09 | [10 empty] | 4.7 | 98
HCBLM_I001-05000801 | IHEP0801 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 70.936 | [10 empty] | 4.7 | 98
HCBLM_I001-05000802 | IHEP0802 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-11 | Tikhonov |  |  |  |  | 1.1 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.282 | [10 empty] | 4.7 | 98
HCBLM_I001-05000803 | IHEP0803 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-12 | Seleznev |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-08-15 | Grishin |  | 3 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-09-29 | Mishagin |  |  |  |  | 1500 | 71.498 | [10 empty] | 4.7 | 98
HCBLM_I001-05000804 | IHEP0804 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-11 | Tikhonov |  |  |  |  | 1.1 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.564 | [10 empty] | 4.7 | 98
HCBLM_I001-05000805 | IHEP0805 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.506 | [10 empty] | 4.7 | 98
HCBLM_I001-05000806 | IHEP0806 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.9 | [10 empty] | 4.7 | 98
HCBLM_I001-05000807 | IHEP0807 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.66 | [10 empty] | 4.7 | 98
HCBLM_I001-05000808 | IHEP0808 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Tikhonov |  |  |  |  | 1245 | IHEP | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.986 | [10 empty] | 4.7 | 98
HCBLM_I001-05000809 | IHEP0809 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Tikhonov |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.088 | [10 empty] | 4.7 | 98
HCBLM_I001-05000810 | IHEP0810 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Seleznev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.282 | [10 empty] | 4.7 | 98
HCBLM_I001-05000811 | IHEP0811 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Seleznev |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.314 | [10 empty] | 4.7 | 98
HCBLM_I001-05000812 | IHEP0812 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Tikhonov |  |  |  |  | 84 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.6 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.6 | [10 empty] | 4.7 | 98
HCBLM_I001-05000813 | IHEP0813 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Seleznev |  |  |  |  | 1240 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.6 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.774 | [10 empty] | 4.7 | 98
HCBLM_I001-05000814 | IHEP0814 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Seleznev |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.016 | [10 empty] | 4.7 | 98
HCBLM_I001-05000815 | IHEP0815 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Tikhonov |  |  |  |  | 16 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.126 | [10 empty] | 4.7 | 98
HCBLM_I001-05000816 | IHEP0816 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Seleznev |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.6 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.876 | [10 empty] | 4.7 | 98
HCBLM_I001-05000817 | IHEP0817 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Seleznev |  |  |  |  | 0.7 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.82 | [10 empty] | 4.7 | 98
HCBLM_I001-05000818 | IHEP0818 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Seleznev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.83 | [10 empty] | 4.7 | 98
HCBLM_I001-05000819 | IHEP0819 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 1.4 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-11 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.016 | [10 empty] | 4.7 | 98
HCBLM_I001-05000820 | IHEP0820 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-16 | Tikhonov |  |  |  |  | 18 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-22 | Seleznev |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.59 | [10 empty] | 4.7 | 98
HCBLM_I001-05000821 | IHEP0821 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Seleznev |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.6 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 67.698 | [10 empty] | 4.7 | 98
HCBLM_I001-05000822 | IHEP0822 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Seleznev |  |  |  |  | 0.5 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.366 | [10 empty] | 4.7 | 98
HCBLM_I001-05000823 | IHEP0823 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-12-08 | Larionov |  |  |  |  | 24 | 2 | lar024_1;lar024_2 | 20 | 1100 | 220 | 6 | lar024_3;lar024_4 | Done | Ok | 3 | 2006-08-11 | Tikhonov |  |  |  |  | 1.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.422 | [10 empty] | 4.7 | 98
HCBLM_I001-05000824 | IHEP0824 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Seleznev |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.744 | [10 empty] | 4.7 | 98
HCBLM_I001-05000825 | IHEP0825 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 1240 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 9 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 83.174 | [10 empty] | 4.7 | 98
HCBLM_I001-05000826 | IHEP0826 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 1.3 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.9 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.551 | [10 empty] | 4.7 | 98
HCBLM_I001-05000827 | IHEP0827 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.338 | [10 empty] | 4.7 | 98
HCBLM_I001-05000828 | IHEP0828 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.3 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.3 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.438 | [10 empty] | 4.7 | 98
HCBLM_I001-05000829 | IHEP0829 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.3 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Mishagin |  |  |  |  | 0.3 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.583 | [10 empty] | 4.7 | 98
HCBLM_I001-05000830 | IHEP0830 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Mishagin |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.512 | [10 empty] | 4.7 | 98
HCBLM_I001-05000831 | IHEP0831 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.234 | [10 empty] | 4.7 | 98
HCBLM_I001-05000832 | IHEP0832 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.2 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.146 | [10 empty] | 4.7 | 98
HCBLM_I001-05000833 | IHEP0833 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.2 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.69 | [10 empty] | 4.7 | 98
HCBLM_I001-05000834 | IHEP0834 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.786 | [10 empty] | 4.7 | 98
HCBLM_I001-05000835 | IHEP0835 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.4 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.236 | [10 empty] | 4.7 | 98
HCBLM_I001-05000836 | IHEP0836 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-11 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.755 | [10 empty] | 4.7 | 98
HCBLM_I001-05000837 | IHEP0837 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-10 | Tikhonov |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.9 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.516 | [10 empty] | 4.7 | 98
HCBLM_I001-05000838 | IHEP0838 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.2 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.3 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.72 | [10 empty] | 4.7 | 98
HCBLM_I001-05000839 | IHEP0839 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Mishagin |  |  |  |  | 0.2 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.696 | [10 empty] | 4.7 | 98
HCBLM_I001-05000840 | IHEP0840 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-09 | Tikhonov |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-08-16 | Larionov |  |  |  |  | 25 | 2 | lar025_1;lar025_2 | 24 | 1112 | 220 | 6 | lar025_3;lar025_4 | Done | Ok | 3 | 2006-08-16 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 74.336 | [10 empty] | 4.7 | 98
HCBLM_I001-05000841 | IHEP0841 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.974 | [10 empty] | 4.7 | 98
HCBLM_I001-05000842 | IHEP0842 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.394 | [10 empty] | 4.7 | 98
HCBLM_I001-05000843 | IHEP0843 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.7 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.048 | [10 empty] | 4.7 | 98
HCBLM_I001-05000844 | IHEP0844 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 1230 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.9 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.436 | [10 empty] | 4.7 | 98
HCBLM_I001-05000845 | IHEP0845 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.364 | [10 empty] | 4.7 | 98
HCBLM_I001-05000846 | IHEP0846 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 1231 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.488 | [10 empty] | 4.7 | 98
HCBLM_I001-05000847 | IHEP0847 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.552 | [10 empty] | 4.7 | 98
HCBLM_I001-05000848 | IHEP0848 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Galyaev |  |  |  |  | 1.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.134 | [10 empty] | 4.7 | 98
HCBLM_I001-05000849 | IHEP0849 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.88 | [10 empty] | 4.7 | 98
HCBLM_I001-05000850 | IHEP0850 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-16 | Tikhonov |  |  |  |  | 19 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.84 | [10 empty] | 4.7 | 98
HCBLM_I001-05000851 | IHEP0851 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.882 | [10 empty] | 4.7 | 98
HCBLM_I001-05000852 | IHEP0852 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.656 | [10 empty] | 4.7 | 98
HCBLM_I001-05000853 | IHEP0853 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 371 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 1.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.522 | [10 empty] | 4.7 | 98
HCBLM_I001-05000854 | IHEP0854 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.1 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.332 | [10 empty] | 4.7 | 98
HCBLM_I001-05000855 | IHEP0855 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 1.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.202 | [10 empty] | 4.7 | 98
HCBLM_I001-05000856 | IHEP0856 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.767 | [10 empty] | 4.7 | 98
HCBLM_I001-05000857 | IHEP0857 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.2 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.1102 | [10 empty] | 4.7 | 98
HCBLM_I001-05000858 | IHEP0858 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.824 | [10 empty] | 4.7 | 98
HCBLM_I001-05000859 | IHEP0859 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 67.462 | [10 empty] | 4.7 | 98
HCBLM_I001-05000860 | IHEP0860 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-16 | Tikhonov |  |  |  |  | 1.1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.494 | [10 empty] | 4.7 | 98
HCBLM_I001-05000861 | IHEP0861 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 4 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.412 | [10 empty] | 4.7 | 98
HCBLM_I001-05000862 | IHEP0862 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.898 | [10 empty] | 4.7 | 98
HCBLM_I001-05000863 | IHEP0863 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Mishagin |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.208 | [10 empty] | 4.7 | 98
HCBLM_I001-05000864 | IHEP0864 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.736 | [10 empty] | 4.7 | 98
HCBLM_I001-05000865 | IHEP0865 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.304 | [10 empty] | 4.7 | 98
HCBLM_I001-05000866 | IHEP0866 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 1.8 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.2469 | [10 empty] | 4.7 | 98
HCBLM_I001-05000867 | IHEP0867 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-16 | Tikhonov |  |  |  |  | 12 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.5 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.026 | [10 empty] | 4.7 | 98
HCBLM_I001-05000868 | IHEP0868 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Galyaev |  |  |  |  | 0.6 | IHEP | 2000 | Cancelled | Not Ok | 2 | [14 empty] | Cancelled | Cancelled | 3 | [20 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05000869 | IHEP0869 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.2 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.504 | [10 empty] | 4.7 | 98
HCBLM_I001-05000870 | IHEP0870 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-16 | Tikhonov |  |  |  |  | 25 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.9 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.298 | [10 empty] | 4.7 | 98
HCBLM_I001-05000871 | IHEP0871 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.858 | [10 empty] | 4.7 | 98
HCBLM_I001-05000872 | IHEP0872 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.416 | [10 empty] | 4.7 | 98
HCBLM_I001-05000873 | IHEP0873 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-17 | Mishagin |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.6 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.956 | [10 empty] | 4.7 | 98
HCBLM_I001-05000874 | IHEP0874 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-16 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.9 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.382 | [10 empty] | 4.7 | 98
HCBLM_I001-05000875 | IHEP0875 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.2 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.38 | [10 empty] | 4.7 | 98
HCBLM_I001-05000876 | IHEP0876 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 1800 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.1469 | [10 empty] | 4.7 | 98
HCBLM_I001-05000877 | IHEP0877 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 1229 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.962 | [10 empty] | 4.7 | 98
HCBLM_I001-05000878 | IHEP0878 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.7 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.06 | [10 empty] | 4.7 | 98
HCBLM_I001-05000879 | IHEP0879 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.186 | [10 empty] | 4.7 | 98
HCBLM_I001-05000880 | IHEP0880 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.638 | [10 empty] | 4.7 | 98
HCBLM_I001-05000881 | IHEP0881 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 233 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000882 | IHEP0882 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Seleznev |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.246 | [10 empty] | 4.7 | 98
HCBLM_I001-05000883 | IHEP0883 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.8 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.55 | [10 empty] | 4.7 | 98
HCBLM_I001-05000884 | IHEP0884 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.05 | [10 empty] | 4.7 | 98
HCBLM_I001-05000885 | IHEP0885 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.586 | [10 empty] | 4.7 | 98
HCBLM_I001-05000886 | IHEP0886 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.356 | [10 empty] | 4.7 | 98
HCBLM_I001-05000887 | IHEP0887 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.456 | [10 empty] | 4.7 | 98
HCBLM_I001-05000888 | IHEP0888 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.84 | [10 empty] | 4.7 | 98
HCBLM_I001-05000889 | IHEP0889 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.076 | [10 empty] | 4.7 | 98
HCBLM_I001-05000890 | IHEP0890 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Tikhonov |  |  |  |  | 12 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.414 | [10 empty] | 4.7 | 98
HCBLM_I001-05000891 | IHEP0891 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 1275 | IHEP | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 229 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-01 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.35 | [10 empty] | 4.7 | 98
HCBLM_I001-05000892 | IHEP0892 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Tikhonov |  |  |  |  | 1223 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-28 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.694 | [10 empty] | 4.7 | 98
HCBLM_I001-05000893 | IHEP0893 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-24 | Tikhonov |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.422 | [10 empty] | 4.7 | 98
HCBLM_I001-05000894 | IHEP0894 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.026 | [10 empty] | 4.7 | 98
HCBLM_I001-05000895 | IHEP0895 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.22 | [10 empty] | 4.7 | 98
HCBLM_I001-05000896 | IHEP0896 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Mishagin |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-28 | Tikhonov |  |  |  |  | 0.6 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.028 | [10 empty] | 4.7 | 98
HCBLM_I001-05000897 | IHEP0897 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Galyaev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.396 | [10 empty] | 4.7 | 98
HCBLM_I001-05000898 | IHEP0898 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Mishagin |  |  |  |  | 1223 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.1 | [10 empty] | 4.7 | 98
HCBLM_I001-05000899 | IHEP0899 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-18 | Tikhonov |  |  |  |  | 2.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-22 | Larionov |  |  |  |  | 26 | 2 | lar026_1;lar026_2 | 23 | 1112 | 220 | 6 | lar026_3;lar026_4 | Done | Ok | 3 | 2006-08-23 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.04 | [10 empty] | 4.7 | 98
HCBLM_I001-05000900 | IHEP0900 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 52 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 67.328 | [10 empty] | 4.7 | 98
HCBLM_I001-05000901 | IHEP0901 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 230 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.424 | [10 empty] | 4.7 | 98
HCBLM_I001-05000902 | IHEP0902 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 231 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.278 | [10 empty] | 4.7 | 98
HCBLM_I001-05000903 | IHEP0903 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.81 | [10 empty] | 4.7 | 98
HCBLM_I001-05000904 | IHEP0904 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.72 | [10 empty] | 4.7 | 98
HCBLM_I001-05000905 | IHEP0905 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 229 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.886 | [10 empty] | 4.7 | 98
HCBLM_I001-05000906 | IHEP0906 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.704 | [10 empty] | 4.7 | 98
HCBLM_I001-05000907 | IHEP0907 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 231 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.272 | [10 empty] | 4.7 | 98
HCBLM_I001-05000908 | IHEP0908 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 226 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-31 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.268 | [10 empty] | 4.7 | 98
HCBLM_I001-05000909 | IHEP0909 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-21 | Tikhonov |  |  |  |  | 15 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.334 | [10 empty] | 4.7 | 98
HCBLM_I001-05000910 | IHEP0910 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 241 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-31 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 72.038 | [10 empty] | 4.7 | 98
HCBLM_I001-05000911 | IHEP0911 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.346 | [10 empty] | 4.7 | 98
HCBLM_I001-05000912 | IHEP0912 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 6.4 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 236 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.67 | [10 empty] | 4.7 | 98
HCBLM_I001-05000913 | IHEP0913 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 6.2 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.238 | [10 empty] | 4.7 | 98
HCBLM_I001-05000914 | IHEP0914 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-28 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.532 | [10 empty] | 4.7 | 98
HCBLM_I001-05000915 | IHEP0915 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 1.4 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.134 | [10 empty] | 4.7 | 98
HCBLM_I001-05000916 | IHEP0916 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 1241 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Not Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 1235 |  | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05000917 | IHEP0917 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 62 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-28 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.414 | [10 empty] | 4.7 | 98
HCBLM_I001-05000918 | IHEP0918 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 227 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.834 | [10 empty] | 4.7 | 98
HCBLM_I001-05000919 | IHEP0919 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 222 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 67.934 | [10 empty] | 4.7 | 98
HCBLM_I001-05000920 | IHEP0920 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 1238 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.948 | [10 empty] | 4.7 | 98
HCBLM_I001-05000921 | IHEP0921 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 227 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.634 | [10 empty] | 4.7 | 98
HCBLM_I001-05000922 | IHEP0922 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.612 | [10 empty] | 4.7 | 98
HCBLM_I001-05000923 | IHEP0923 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 232 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-31 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.054 | [10 empty] | 4.7 | 98
HCBLM_I001-05000924 | IHEP0924 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 1235 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 1.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.888 | [10 empty] | 4.7 | 98
HCBLM_I001-05000925 | IHEP0925 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 247 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.256 | [10 empty] | 4.7 | 98
HCBLM_I001-05000926 | IHEP0926 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 1.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.686 | [10 empty] | 4.7 | 98
HCBLM_I001-05000927 | IHEP0927 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.883 | [10 empty] | 4.7 | 98
HCBLM_I001-05000928 | IHEP0928 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.63 | [10 empty] | 4.7 | 98
HCBLM_I001-05000929 | IHEP0929 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 1.7 | IHEP | 2000 | Done | Ok | 2 |  | Larionov |  |  |  |  | 27 | 2 | lar027_1;lar027_2 | 23 | 1110 | 220 | 6 | lar027_3;lar027_4 | Done | Ok | 3 | 2006-08-25 | Tikhonov |  |  |  |  | 0.9 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.93 | [10 empty] | 4.7 | 98
HCBLM_I001-05000930 | IHEP0930 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 72.81 | [10 empty] | 4.7 | 98
HCBLM_I001-05000931 | IHEP0931 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 225 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.7 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.778 | [10 empty] | 4.7 | 98
HCBLM_I001-05000932 | IHEP0932 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 226 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.7 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05000933 | IHEP0933 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 229 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000934 | IHEP0934 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Mishagin |  |  |  |  | 1237 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 224 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.306 | [10 empty] | 4.7 | 98
HCBLM_I001-05000935 | IHEP0935 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 238 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.858 | [10 empty] | 4.7 | 98
HCBLM_I001-05000936 | IHEP0936 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Mishagin |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 1.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.526 | [10 empty] | 4.7 | 98
HCBLM_I001-05000937 | IHEP0937 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 221 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.7163 | [10 empty] | 4.7 | 98
HCBLM_I001-05000938 | IHEP0938 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 235 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 72.012 | [10 empty] | 4.7 | 98
HCBLM_I001-05000939 | IHEP0939 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 228 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.066 | [10 empty] | 4.7 | 98
HCBLM_I001-05000940 | IHEP0940 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 256 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 1.6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.428 | [10 empty] | 4.7 | 98
HCBLM_I001-05000941 | IHEP0941 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-22 | Galyaev |  |  |  |  | 4.2 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 235 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.608 | [10 empty] | 4.7 | 98
HCBLM_I001-05000942 | IHEP0942 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Mishagin |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 247 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.756 | [10 empty] | 4.7 | 98
HCBLM_I001-05000943 | IHEP0943 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Mishagin |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 224 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.55 | [10 empty] | 4.7 | 98
HCBLM_I001-05000944 | IHEP0944 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-31 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.59 | [10 empty] | 4.7 | 98
HCBLM_I001-05000945 | IHEP0945 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 241 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.942 | [10 empty] | 4.7 | 98
HCBLM_I001-05000946 | IHEP0946 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Mishagin |  |  |  |  | 1227 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | -6 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.998 | [10 empty] | 4.7 | 98
HCBLM_I001-05000947 | IHEP0947 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.324 | [10 empty] | 4.7 | 98
HCBLM_I001-05000948 | IHEP0948 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 226 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 71.44 | [10 empty] | 4.7 | 98
HCBLM_I001-05000949 | IHEP0949 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 236 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05000950 | IHEP0950 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05000951 | IHEP0951 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 229 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8
HCBLM_I001-05000952 | IHEP0952 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000953 | IHEP0953 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.6 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05000954 | IHEP0954 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 222 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.6 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05000955 | IHEP0955 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 252 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05000956 | IHEP0956 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000957 | IHEP0957 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05000958 | IHEP0958 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05000959 | IHEP0959 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 221 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000960 | IHEP0960 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05000961 | IHEP0961 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000962 | IHEP0962 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 237 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05000963 | IHEP0963 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 240 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8
HCBLM_I001-05000964 | IHEP0964 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 240 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05000965 | IHEP0965 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 256 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8
HCBLM_I001-05000966 | IHEP0966 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 1320 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 69.104 | [10 empty] | 4.7 | 98
HCBLM_I001-05000967 | IHEP0967 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000968 | IHEP0968 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 239 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.8 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000969 | IHEP0969 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 231 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05000970 | IHEP0970 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 240 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000971 | IHEP0971 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-08-29 | Larionov |  |  |  |  | 28 | 2 | lar028_1;lar028_2 | 23 | 1101 | 220 | 6 | lar028_3;lar028_4 | Done | Ok | 3 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 1.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.536 | [10 empty] | 4.7 | 98
HCBLM_I001-05000972 | IHEP0972 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-24 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 226 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05000973 | IHEP0973 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 232 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000974 | IHEP0974 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 224 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000975 | IHEP0975 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 241 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000976 | IHEP0976 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 247 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000977 | IHEP0977 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 227 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05000978 | IHEP0978 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 228 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05000979 | IHEP0979 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 237 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.7 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05000980 | IHEP0980 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 2.1 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 234 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000981 | IHEP0981 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000982 | IHEP0982 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 227 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000983 | IHEP0983 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05000984 | IHEP0984 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05000985 | IHEP0985 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 257 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8
HCBLM_I001-05000986 | IHEP0986 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 1230 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 233 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-01 | Tikhonov |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 68.91 | [10 empty] | 4.7 | 98
HCBLM_I001-05000987 | IHEP0987 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | -1.3
HCBLM_I001-05000988 | IHEP0988 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 1240 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 221 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05000989 | IHEP0989 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 1200 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 252 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-01 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 4.5 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 78.27 | [10 empty] | 4.7 | 98
HCBLM_I001-05000990 | IHEP0990 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 230 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000991 | IHEP0991 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 1240 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000992 | IHEP0992 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 224 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000993 | IHEP0993 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 300 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 233 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05000994 | IHEP0994 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 1241 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000995 | IHEP0995 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 29 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 227 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-01 | Tikhonov |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 70.158 | [10 empty] | 4.7 | 98
HCBLM_I001-05000996 | IHEP0996 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 1231 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 224 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | [21 empty] | 4.7
HCBLM_I001-05000997 | IHEP0997 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 149 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05000998 | IHEP0998 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05000999 | IHEP0999 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 1215 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 234 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001000 | IHEP1000 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 252 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001001 | IHEP1001 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-04 | Tikhonov |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001002 | IHEP1002 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 222 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001003 | IHEP1003 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Seleznev |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 232 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Seleznev |  |  |  |  | 0.3 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 1.4
HCBLM_I001-05001004 | IHEP1004 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 230 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001005 | IHEP1005 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Seleznev |  |  |  |  | 0.4 | UK | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 229 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.4 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001006 | IHEP1006 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 1600 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001007 | IHEP1007 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 221 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001008 | IHEP1008 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-30 | Tikhonov |  |  |  |  | 1200 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.7 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001009 | IHEP1009 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 200 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 229 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001010 | IHEP1010 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.9 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8
HCBLM_I001-05001011 | IHEP1011 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 227 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-06 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001012 | IHEP1012 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 224 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.7 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001013 | IHEP1013 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-01 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001014 | IHEP1014 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 223 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001015 | IHEP1015 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.7 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001016 | IHEP1016 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 231 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.9 | 0.1 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001017 | IHEP1017 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-01 | Seleznev |  |  |  |  | 1.8 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001018 | IHEP1018 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 233 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001019 | IHEP1019 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-01 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 230 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001020 | IHEP1020 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 221 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.8 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.9
HCBLM_I001-05001021 | IHEP1021 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 227 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001022 | IHEP1022 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 228 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001023 | IHEP1023 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Mishagin |  |  |  |  | 13.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 224 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0
HCBLM_I001-05001024 | IHEP1024 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 224 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001025 | IHEP1025 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-29 | Tikhonov |  |  |  |  | 1231 | UK | 2000 | Done | Ok | 2 | 2006-01-09 | Larionov |  |  |  |  | 29 | 2 | lar029_1;lar029_2 | 20 | 1101 | 220 | 6 | lar029_3;lar029_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001026 | IHEP1026 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-01 | Seleznev |  |  |  |  | 1.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 240 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001027 | IHEP1027 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Mishagin |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 221 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0
HCBLM_I001-05001028 | IHEP1028 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 232 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.3 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001029 | IHEP1029 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Mishagin |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 237 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001030 | IHEP1030 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 1200 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 232 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001031 | IHEP1031 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 223 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001032 | IHEP1032 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 19 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.6 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001033 | IHEP1033 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001034 | IHEP1034 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8
HCBLM_I001-05001035 | IHEP1035 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Mishagin |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 232 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001036 | IHEP1036 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.4 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001037 | IHEP1037 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 20 | 1094 | 222 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.4 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001038 | IHEP1038 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 230 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001039 | IHEP1039 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Mishagin |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 256 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001040 | IHEP1040 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001041 | IHEP1041 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001042 | IHEP1042 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 2.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001043 | IHEP1043 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 230 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001044 | IHEP1044 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 225 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001045 | IHEP1045 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 234 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001046 | IHEP1046 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Seleznev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001047 | IHEP1047 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-07 | Seleznev |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001048 | IHEP1048 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-31 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-09 | Larionov |  |  |  |  | 30 | 2 | lar030_1;lar030_2 | 21 | 1103 | 220 | 6 | lar030_3;lar030_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001049 | IHEP1049 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Mishagin |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001050 | IHEP1050 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001051 | IHEP1051 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 230 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.7 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001052 | IHEP1052 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 2.1 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 234 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8
HCBLM_I001-05001053 | IHEP1053 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 1.3 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 240 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001054 | IHEP1054 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 224 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001055 | IHEP1055 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 20 | 1094 | 234 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.7 | 0.3 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001056 | IHEP1056 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Seleznev |  |  |  |  | 0.6 | UK | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 20 | 1094 | 220 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001057 | IHEP1057 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001058 | IHEP1058 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 225 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0
HCBLM_I001-05001059 | IHEP1059 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 21 | 1102 | 228 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001060 | IHEP1060 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 234 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001061 | IHEP1061 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 25.4 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001062 | IHEP1062 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 1.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 223 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001063 | IHEP1063 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 240 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0
HCBLM_I001-05001064 | IHEP1064 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 19 | 1093 | 224 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.7 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001065 | IHEP1065 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 221 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001066 | IHEP1066 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 19 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001067 | IHEP1067 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 240 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001068 | IHEP1068 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.6 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.8
HCBLM_I001-05001069 | IHEP1069 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 232 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001070 | IHEP1070 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | UK | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 229 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001071 | IHEP1071 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001072 | IHEP1072 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 1235 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 240 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001073 | IHEP1073 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 21 | 1102 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001074 | IHEP1074 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 234 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001075 | IHEP1075 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 2.1 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0
HCBLM_I001-05001076 | IHEP1076 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 2.2 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001077 | IHEP1077 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 225 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.9 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001078 | IHEP1078 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 230 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001079 | IHEP1079 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001080 | IHEP1080 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 228 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001081 | IHEP1081 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 2.5 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 19 | 1093 | 229 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001082 | IHEP1082 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 19 | 1093 | 229 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001083 | IHEP1083 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 19 | 1093 | 247 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.7 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001084 | IHEP1084 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 229 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001085 | IHEP1085 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 227 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001086 | IHEP1086 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001087 | IHEP1087 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 231 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001088 | IHEP1088 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 21 | 1102 | 232 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001089 | IHEP1089 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 229 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001090 | IHEP1090 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Seleznev |  |  |  |  | 0.8 | UK | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 237 | 6 | no | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001091 | IHEP1091 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001092 | IHEP1092 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 237 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001093 | IHEP1093 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 21 | 1102 | 230 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | -122
HCBLM_I001-05001094 | IHEP1094 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1100 | 229 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001095 | IHEP1095 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001096 | IHEP1096 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 30 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001097 | IHEP1097 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-04 | Galyaev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-07-09 | Larionov |  |  |  |  | 31 | 2 | lar031_1;lar031_2 | 21 | 1102 | 220 | 6 | lar031_3;lar031_4 | Done | Ok | 3 | 2006-09-08 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001098 | IHEP1098 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 240 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.9 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001099 | IHEP1099 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 1200 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 225 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001100 | IHEP1100 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 240 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001101 | IHEP1101 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 221 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | -0.2
HCBLM_I001-05001102 | IHEP1102 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 234 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.4 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001103 | IHEP1103 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 240 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001104 | IHEP1104 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 5 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 224 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001105 | IHEP1105 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.7 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001106 | IHEP1106 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001107 | IHEP1107 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 233 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001108 | IHEP1108 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 232 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001109 | IHEP1109 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 21 | 1102 | 222 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001110 | IHEP1110 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 247 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001111 | IHEP1111 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001112 | IHEP1112 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 240 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001113 | IHEP1113 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1100 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001114 | IHEP1114 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-09 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1094 | 221 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001115 | IHEP1115 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 230 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001116 | IHEP1116 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001117 | IHEP1117 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 224 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001118 | IHEP1118 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 1225 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1100 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001119 | IHEP1119 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 228 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001120 | IHEP1120 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 224 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001121 | IHEP1121 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 227 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001122 | IHEP1122 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.3 | 0.5 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001123 | IHEP1123 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001124 | IHEP1124 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 230 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001125 | IHEP1125 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001126 | IHEP1126 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001127 | IHEP1127 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 240 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.5 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001128 | IHEP1128 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001129 | IHEP1129 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Seleznev |  |  |  |  | 2.2 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 224 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7
HCBLM_I001-05001130 | IHEP1130 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 256 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001131 | IHEP1131 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001132 | IHEP1132 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001133 | IHEP1133 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001134 | IHEP1134 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 232 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001135 | IHEP1135 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001136 | IHEP1136 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 234 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.4 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001137 | IHEP1137 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1100 | 227 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001138 | IHEP1138 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 223 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001139 | IHEP1139 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 229 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001140 | IHEP1140 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 229 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001141 | IHEP1141 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 234 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001142 | IHEP1142 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 21 | 1101 | 234 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001143 | IHEP1143 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 38 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 252 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.9 | 0.6 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001144 | IHEP1144 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 225 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.7 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001145 | IHEP1145 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 21 | 1102 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001146 | IHEP1146 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 240 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.2 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6
HCBLM_I001-05001147 | IHEP1147 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 10 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5
HCBLM_I001-05001148 | IHEP1148 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | [11 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001149 | [16 empty] | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 227 | 6 | lar034_3;lar034_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001150 | [16 empty] | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 21 | 1101 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001151 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 241 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001152 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001153 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 223 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001154 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001155 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 227 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001156 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001157 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 221 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001158 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001159 | [16 empty] | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | [21 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 |  | Mishagin |  |  |  |  | 2000 | 0.2
HCBLM_I001-05001160 | [16 empty] | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 224 | 6 | ter033_3;ter033_4 | Done | Ok | [21 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 |  | Mishagin |  |  |  |  | 2000 | 0.1
HCBLM_I001-05001161 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 239 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001162 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 226 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001163 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 225 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001164 | [16 empty] | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 221 | 6 | lar034_3;lar034_4 | Done | Ok | [21 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 |  | Mishagin |  |  |  |  | 2000 | 0.4
HCBLM_I001-05001165 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 233 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001166 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001167 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 224 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001168 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 226 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001169 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001170 | [16 empty] | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 238 | 6 | ter033_3;ter033_4 | Done | Ok | [21 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4
HCBLM_I001-05001171 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001172 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 234 | 6 | lar035_3;lar035_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001173 | [16 empty] | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 21 | 1101 | 222 | 6 | lar034_3;lar034_4 | Done | Ok | [21 empty] | Done | Ok | 5 |  | Grishin
HCBLM_I001-05001174 | [16 empty] | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 257 | 6 | ter033_3;ter033_4 | Done | Ok | [21 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  | 4 | [5 empty] | Done | Ok | 6 |  | Mishagin |  |  |  |  | 2000 | 0.3
HCBLM_I001-05001175 | [16 empty] | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 237 | 6 | lar034_3;lar034_4 | Done | Ok
HCBLM_I001-05001176 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 221 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001177 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001178 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 255 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001179 | [16 empty] | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok
HCBLM_I001-05001180 | [16 empty] | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 234 | 6 | lar034_3;lar034_4 | Done | Ok
HCBLM_I001-05001181 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001182 | [16 empty] | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 223 | 6 | lar034_3;lar034_4 | Done | Ok
HCBLM_I001-05001183 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001184 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 222 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001185 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 241 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001186 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001187 | [16 empty] | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok
HCBLM_I001-05001188 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 233 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001189 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 233 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001190 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 237 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001191 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 226 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001192 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 229 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001193 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 228 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001194 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 220 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001195 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 230 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001196 | [16 empty] | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 256 | 6 | lar034_3;lar034_4 | Done | Ok
HCBLM_I001-05001197 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 230 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001198 | [16 empty] | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 |  |  | 251 | 6 | lar035_3;lar035_4 | Done | Ok
HCBLM_I001-05001199
HCBLM_I001-05001200
HCBLM_I001-05001201
HCBLM_I001-05001202
HCBLM_I001-05001203
HCBLM_I001-05001204
HCBLM_I001-05001205
HCBLM_I001-05001206
HCBLM_I001-05001207
HCBLM_I001-05001208
HCBLM_I001-05001209
HCBLM_I001-05001210
HCBLM_I001-05001211
HCBLM_I001-05001212
HCBLM_I001-05001213
HCBLM_I001-05001214
HCBLM_I001-05001215
HCBLM_I001-05001216
HCBLM_I001-05001217
HCBLM_I001-05001218
HCBLM_I001-05001219
HCBLM_I001-05001220
HCBLM_I001-05001221
HCBLM_I001-05001222
HCBLM_I001-05001223
HCBLM_I001-05001224
HCBLM_I001-05001225
HCBLM_I001-05001226
HCBLM_I001-05001227
HCBLM_I001-05001228
HCBLM_I001-05001229
HCBLM_I001-05001230
HCBLM_I001-05001231
HCBLM_I001-05001232
HCBLM_I001-05001233
HCBLM_I001-05001234
HCBLM_I001-05001235
HCBLM_I001-05001236
HCBLM_I001-05001237
HCBLM_I001-05001238
HCBLM_I001-05001239
HCBLM_I001-05001240
HCBLM_I001-05001241
HCBLM_I001-05001242
HCBLM_I001-05001243
HCBLM_I001-05001244
HCBLM_I001-05001245
HCBLM_I001-05001246
HCBLM_I001-05001247
HCBLM_I001-05001248
HCBLM_I001-05001249
HCBLM_I001-05001250
HCBLM_I001-05001251
HCBLM_I001-05001252
HCBLM_I001-05001253
HCBLM_I001-05001254
HCBLM_I001-05001255
HCBLM_I001-05001256
HCBLM_I001-05001257
HCBLM_I001-05001258
HCBLM_I001-05001259
HCBLM_I001-05001260
HCBLM_I001-05001261
HCBLM_I001-05001262
HCBLM_I001-05001263
HCBLM_I001-05001264
HCBLM_I001-05001265
HCBLM_I001-05001266
HCBLM_I001-05001267
HCBLM_I001-05001268
HCBLM_I001-05001269
HCBLM_I001-05001270
HCBLM_I001-05001271
HCBLM_I001-05001272
HCBLM_I001-05001273
HCBLM_I001-05001274
HCBLM_I001-05001275
HCBLM_I001-05001276
HCBLM_I001-05001277
HCBLM_I001-05001278
HCBLM_I001-05001279
HCBLM_I001-05001280
HCBLM_I001-05001281
HCBLM_I001-05001282
HCBLM_I001-05001283
HCBLM_I001-05001284
HCBLM_I001-05001285
HCBLM_I001-05001286
HCBLM_I001-05001287
HCBLM_I001-05001288
HCBLM_I001-05001289
HCBLM_I001-05001290
HCBLM_I001-05001291
HCBLM_I001-05001292
HCBLM_I001-05001293
HCBLM_I001-05001294
HCBLM_I001-05001295
HCBLM_I001-05001296
HCBLM_I001-05001297
HCBLM_I001-05001298
HCBLM_I001-05001299
HCBLM_I001-05001300
HCBLM_I001-05001301
HCBLM_I001-05001302
HCBLM_I001-05001303
HCBLM_I001-05001304
HCBLM_I001-05001305
HCBLM_I001-05001306
HCBLM_I001-05001307
HCBLM_I001-05001308
HCBLM_I001-05001309
HCBLM_I001-05001310
HCBLM_I001-05001311
HCBLM_I001-05001312
HCBLM_I001-05001313
HCBLM_I001-05001314
HCBLM_I001-05001315
HCBLM_I001-05001316
HCBLM_I001-05001317
HCBLM_I001-05001318
HCBLM_I001-05001319
HCBLM_I001-05001320
HCBLM_I001-05001321
HCBLM_I001-05001322
HCBLM_I001-05001323
HCBLM_I001-05001324
HCBLM_I001-05001325
HCBLM_I001-05001326
HCBLM_I001-05001327
HCBLM_I001-05001328
HCBLM_I001-05001329
HCBLM_I001-05001330
HCBLM_I001-05001331
HCBLM_I001-05001332
HCBLM_I001-05001333
HCBLM_I001-05001334
HCBLM_I001-05001335
HCBLM_I001-05001336
HCBLM_I001-05001337
HCBLM_I001-05001338
HCBLM_I001-05001339
HCBLM_I001-05001340
HCBLM_I001-05001341
HCBLM_I001-05001342
HCBLM_I001-05001343
HCBLM_I001-05001344
HCBLM_I001-05001345
HCBLM_I001-05001346
HCBLM_I001-05001347
HCBLM_I001-05001348
HCBLM_I001-05001349
HCBLM_I001-05001350
HCBLM_I001-05001351
HCBLM_I001-05001352
HCBLM_I001-05001353
HCBLM_I001-05001354
HCBLM_I001-05001355
HCBLM_I001-05001356
HCBLM_I001-05001357
HCBLM_I001-05001358
HCBLM_I001-05001359
HCBLM_I001-05001360
HCBLM_I001-05001361
HCBLM_I001-05001362
HCBLM_I001-05001363
HCBLM_I001-05001364
HCBLM_I001-05001365
HCBLM_I001-05001366
HCBLM_I001-05001367
HCBLM_I001-05001368
HCBLM_I001-05001369
HCBLM_I001-05001370
HCBLM_I001-05001371
HCBLM_I001-05001372
HCBLM_I001-05001373
HCBLM_I001-05001374
HCBLM_I001-05001375
HCBLM_I001-05001376
HCBLM_I001-05001377
HCBLM_I001-05001378
HCBLM_I001-05001379
HCBLM_I001-05001380
HCBLM_I001-05001381
HCBLM_I001-05001382
HCBLM_I001-05001383
HCBLM_I001-05001384
HCBLM_I001-05001385
HCBLM_I001-05001386
HCBLM_I001-05001387
HCBLM_I001-05001388
HCBLM_I001-05001389
HCBLM_I001-05001390
HCBLM_I001-05001391
HCBLM_I001-05001392
HCBLM_I001-05001393
HCBLM_I001-05001394
HCBLM_I001-05001395
HCBLM_I001-05001396
HCBLM_I001-05001397
HCBLM_I001-05001398
HCBLM_I001-05001399
HCBLM_I001-05001400
HCBLM_I001-05001401
HCBLM_I001-05001402
HCBLM_I001-05001403
HCBLM_I001-05001404
HCBLM_I001-05001405
HCBLM_I001-05001406
HCBLM_I001-05001407
HCBLM_I001-05001408
HCBLM_I001-05001409
HCBLM_I001-05001410
HCBLM_I001-05001411
HCBLM_I001-05001412
HCBLM_I001-05001413
HCBLM_I001-05001414
HCBLM_I001-05001415
HCBLM_I001-05001416
HCBLM_I001-05001417
HCBLM_I001-05001418
HCBLM_I001-05001419
HCBLM_I001-05001420
HCBLM_I001-05001421
HCBLM_I001-05001422
HCBLM_I001-05001423
HCBLM_I001-05001424
HCBLM_I001-05001425
HCBLM_I001-05001426
HCBLM_I001-05001427
HCBLM_I001-05001428
HCBLM_I001-05001429
HCBLM_I001-05001430
HCBLM_I001-05001431
HCBLM_I001-05001432
HCBLM_I001-05001433
HCBLM_I001-05001434
HCBLM_I001-05001435
HCBLM_I001-05001436
HCBLM_I001-05001437
HCBLM_I001-05001438
HCBLM_I001-05001439
HCBLM_I001-05001440
HCBLM_I001-05001441
HCBLM_I001-05001442
HCBLM_I001-05001443
HCBLM_I001-05001444
HCBLM_I001-05001445
HCBLM_I001-05001446
HCBLM_I001-05001447
HCBLM_I001-05001448
HCBLM_I001-05001449
HCBLM_I001-05001450
HCBLM_I001-05001451
HCBLM_I001-05001452
HCBLM_I001-05001453
HCBLM_I001-05001454
HCBLM_I001-05001455
HCBLM_I001-05001456
HCBLM_I001-05001457
HCBLM_I001-05001458
HCBLM_I001-05001459
HCBLM_I001-05001460
HCBLM_I001-05001461
HCBLM_I001-05001462
HCBLM_I001-05001463
HCBLM_I001-05001464
HCBLM_I001-05001465
HCBLM_I001-05001466
HCBLM_I001-05001467
HCBLM_I001-05001468
HCBLM_I001-05001469
HCBLM_I001-05001470
HCBLM_I001-05001471
HCBLM_I001-05001472
HCBLM_I001-05001473
HCBLM_I001-05001474
HCBLM_I001-05001475
HCBLM_I001-05001476
HCBLM_I001-05001477
HCBLM_I001-05001478
HCBLM_I001-05001479
HCBLM_I001-05001480
HCBLM_I001-05001481
HCBLM_I001-05001482
HCBLM_I001-05001483
HCBLM_I001-05001484
HCBLM_I001-05001485
HCBLM_I001-05001486
HCBLM_I001-05001487
HCBLM_I001-05001488
HCBLM_I001-05001489
HCBLM_I001-05001490
HCBLM_I001-05001491
HCBLM_I001-05001492
HCBLM_I001-05001493
HCBLM_I001-05001494
HCBLM_I001-05001495
HCBLM_I001-05001496
HCBLM_I001-05001497
HCBLM_I001-05001498
HCBLM_I001-05001499
HCBLM_I001-05001500

## 2.Standard Steps
Standard Steps
Beam Loss Monitor: Ionization Chamber
Standard Step ID | Other identifier | Sequence number | Description | Provided by | Expected by | Responsible | Aplicable standard | Chapter | Comment
MTF013476 |  | 1 | Assembly | IHEP | CERN | Asanov,Seleznev |  |  | Head's welding procedure; Cleaning;Assembly;Current measurement
MTF013477 |  | 2 | Vacuum stand | IHEP | CERN | Larionov |  |  | Vacuum procedure; 36 IC per batch
MTF013478 |  | 3 | Leakage current measurement at IHEP by IHEP | IHEP | CERN | Seleznev |  |  | CERN Picometer noice +/- 0.3 pA
MTF013479 |  | 4 | Leakage current measurement at IHEP by CERN | CERN | CERN | Ferioli |  |  | CERN Picometer noice +/- 0.3 pA
MTF013480 |  | 5 | Shipping | IHEP,CERN | CERN | Seleznev,Grishin |  |  | 360 IC per month
MTF013481 |  | 6 | Leakage current measurement at CERN | CERN,IHEP | CERN | Ferioli,Sleptsov |  |  | CERN Picometer noice +/- 0.3 pA
MTF013482 |  | 7 | Source test | CERN,IHEP | CERN | Ferioli,Sleptsov |  |  | For 6 HV values measure 6 Current values
MTF013483 |  | 8 | Source test at tunnel | CERN,IHEP | CERN | Ferioli,Sleptsov |  |  | archive:measure the parameters of IC time to time

## 3.Properties
Properties
Description | Code | Unit | Type | Category | Default | Comment
Leakage Current (Assemb) | P11483 | pA | NUM
Head Producer | P11446 |  | CHAR |  |  | DG Technology(Italy),Aerovac(UK),Alca(Italy),IHEP(Russia)
High Voltage (Assembly) | P11448 | V | NUM
Leakage Current (IHEP) | P11449 | pA | NUM
Batch Number | P11450 |  | NUM |  |  | Number of batch on the vacuum stand
Gas Bottle Number | P11468 |  | NUM
Spectrum (before filling) | P11453 |  | CHAR
Temperature (filling) | P11451 | °C | NUM
Pressure (filling) | P11452 | mbar | NUM
Temperature (heating) | P11454 | °C | NUM
Duration (heating) | P11456 | h | NUM
Spectrum (after filling) | P11455 |  | CHAR
Leakage Current_v | P11457 | pA | NUM |  |  | Measurement after vacuum stand
Leakage Current_f | P11458 | pA | CHAR |  |  | Measurement after el. box soldering
Lorry Number | P11461 |  | NUM |  |  | Number of lorry from Protvino
IHEP Departure Date | P11462 |  | CHAR
CERN Arrival Date | P11463 |  | CHAR
High Voltage1 | P11471 | V | NUM
Current1 | P11477 | pA | NUM
High Voltage2 | P11472 | V | NUM
Current2 | P11478 | pA | NUM
High Voltage3 | P11473 | V | NUM
Current3 | P11479 | pA | NUM
High Voltage4 | P11474 | V | NUM
Current4 | P11480 | pA | NUM
High Voltage5 | P11475 | V | NUM
Current5 | P11481 | pA | NUM
High Voltage6 | P11476 | V | NUM
Current6 | P11482 | pA | NUM
Current | P11470 | pA | NUM |  |  | Rad source current measurements
Dose Rate | P11467 | millisievert par heure | NUM
Source Activity (Surface) | P11469 | GBq | NUM
High Voltage (Leak-I-I) | P11459 | V | NUM
High Voltage (Leak-C-C) | P11464 | V | NUM
High Voltage (Leak-I-C) | P11460 | V | NUM
High Voltage (Tunnel) | P11466 | V | NUM
Leakage Current (CERN) | P11465 | pA | NUM
Source Activity (Tunnel) | P11447 | GBq | NUM

## Personal Information
Personal Information
EDMSlogin | JJGRAS
EDMSGroup | AB-BI-BLM-MTF
Created On | 2005-01-01