## HCBLM_I001
Beam Loss Monitor: Ionization Chamber
Identification |  |  |  |  | Step 1 (): Assembly | [6 empty] | Step 1: Non-Conformities |  | Properties |  |  | Step 2 (): Vacuum stand | [6 empty] | Step 2: Non-Conformities |  | Properties | [7 empty] | Step 3 (): Leakage current measurement at IHEP by IHEP | [6 empty] | Step 3: Non-Conformities |  | Properties |  |  | Step 4 (): Leakage current measurement at IHEP by CERN | [6 empty] | Step 4: Non-Conformities |  | Properties |  | Step 5 (): Shipping | [6 empty] | Step 5: Non-Conformities |  | Properties |  |  | Step 6 (): Leakage current measurement at CERN | [6 empty] | Step 6: Non-Conformities |  | Properties |  | Step 7 (): Source test | [6 empty] | Step 7: Non-Conformities |  | Properties | [13 empty] | Step 8 (): Source test at tunnel | [6 empty] | Step 8: Non-Conformities |  | Properties |  |  | Equipment Non-Conformities |  | Additional document 1 | Additional document 2 | Additional document 3
Equipment ID | Other ID | Manufacturer | Status | Comment | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Leakage Current (Assemb) | Head Producer | High Voltage (Assembly) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Batch Number | Gas Bottle Number | Spectrum (before filling) | Temperature (filling) | Pressure (filling) | Temperature (heating) | Duration (heating) | Spectrum (after filling) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Leakage Current_f | Leakage Current_v | High Voltage (Leak-I-I) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Leakage Current (IHEP) | High Voltage (Leak-I-C) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Lorry Number | IHEP Departure Date | CERN Arrival Date | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | High Voltage (Leak-C-C) | Leakage Current (CERN) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | High Voltage1 | Current1 | High Voltage2 | Current2 | High Voltage3 | Current3 | High Voltage4 | Current4 | High Voltage5 | Current5 | High Voltage6 | Current6 | Dose Rate | Source Activity (Surface) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Current | High Voltage (Tunnel) | Source Activity (Tunnel) | Doc. Number | Type | Doc. Number | Doc. Number | Doc. Number
[14 empty] | [pA] | [] | [V] | [9 empty] | [] | [] | [] | [°C] | [mbar] | [°C] | [h] | [] | [9 empty] | [pA] | [pA] | [V] | [9 empty] | [pA] | [V] | [9 empty] | [] | [] | [] | [9 empty] | [V] | [pA] | [9 empty] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [millisievert par heure] | [GBq] | [9 empty] | [pA] | [V] | [GBq]
(CERN ID) | [42 empty] | <1 pA
BL01 | HCBLM_I001 | Beam Loss Monitor: Ionization Chamber |  |  | MTF013476 | [7 empty] | 3 | P11483 | P11446 | P11448 | MTF013477 | [7 empty] | 8 | P11450 | P11468 | P11453 | P11451 | P11452 | P11454 | P11456 | P11455 | MTF013478 | [7 empty] | 3 | P11458 | P11457 | P11459 | MTF013479 | [7 empty] | 2 | P11449 | P11460 | MTF013480 | [7 empty] | 3 | P11461 | P11462 | P11463 | MTF013481 | [7 empty] | 2 | P11464 | P11465 | MTF013482 | [7 empty] | 14 | P11471 | P11477 | P11472 | P11478 | P11473 | P11479 | P11474 | P11480 | P11475 | P11481 | P11476 | P11482 | P11467 | P11469 | MTF013483 | [7 empty] | 3 | P11470 | P11466 | P11447
| 0 | 8 |  | 3
HCBLM_I001-05001101 | IHEP1101 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 221 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | -0.2 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.5601 | [12 empty] | Pending |  | 8
HCBLM_I001-05001102 | IHEP1102 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 234 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.1035 | [12 empty] | Pending |  | 8
HCBLM_I001-05001103 | IHEP1103 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 240 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.0571 | [12 empty] | Pending |  | 8
HCBLM_I001-05001104 | IHEP1104 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 5 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 224 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.802 | [12 empty] | Pending |  | 8
HCBLM_I001-05001105 | IHEP1105 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.7 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 73.0957 | [12 empty] | Pending |  | 8
HCBLM_I001-05001106 | IHEP1106 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 77.5269 | [12 empty] | Pending |  | 8
HCBLM_I001-05001107 | IHEP1107 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 233 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 72.6416 | [12 empty] | Pending |  | 8
HCBLM_I001-05001108 | IHEP1108 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 232 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 73.8281 | [12 empty] | Pending |  | 8
HCBLM_I001-05001109 | IHEP1109 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 21 | 1102 | 222 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 70.9399 | [12 empty] | Pending |  | 8
HCBLM_I001-05001110 | IHEP1110 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 247 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.176 | [12 empty] | Pending |  | 8
HCBLM_I001-05001111 | IHEP1111 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.824 | [12 empty] | Pending |  | 8
HCBLM_I001-05001112 | IHEP1112 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 240 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.416 | [12 empty] | Pending |  | 8
HCBLM_I001-05001113 | IHEP1113 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1100 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.996 | [12 empty] | Pending |  | 8
HCBLM_I001-05001114 | IHEP1114 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-09 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1094 | 221 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.3574 | [12 empty] | Pending |  | 8
HCBLM_I001-05001115 | IHEP1115 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 230 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.9 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.44 | [12 empty] | Pending |  | 8
HCBLM_I001-05001116 | IHEP1116 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.902 | [12 empty] | Pending |  | 8
HCBLM_I001-05001117 | IHEP1117 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 224 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.2256 | [12 empty] | Pending |  | 8
HCBLM_I001-05001118 | IHEP1118 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 1225 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1100 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.428 | [12 empty] | Pending |  | 8
HCBLM_I001-05001119 | IHEP1119 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 228 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.928 | [12 empty] | Pending |  | 8
HCBLM_I001-05001120 | IHEP1120 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 224 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.642 | [12 empty] | Pending |  | 8
HCBLM_I001-05001121 | IHEP1121 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 227 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 70.4126 | [12 empty] | Pending |  | 8
HCBLM_I001-05001122 | IHEP1122 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.3 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 73.3984 | [12 empty] | Pending |  | 8
HCBLM_I001-05001123 | IHEP1123 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.9116 | [12 empty] | Pending |  | 8
HCBLM_I001-05001124 | IHEP1124 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 230 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 73.0127 | [12 empty] | Pending |  | 8
HCBLM_I001-05001125 | IHEP1125 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.806 | [12 empty] | Pending |  | 8
HCBLM_I001-05001126 | IHEP1126 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 70.3198 | [12 empty] | Pending |  | 8 | [14 empty] | 15L8
HCBLM_I001-05001127 | IHEP1127 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1094 | 240 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.5 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.026 | [12 empty] | Pending |  | 8
HCBLM_I001-05001128 | IHEP1128 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.9946 | [12 empty] | Pending |  | 8
HCBLM_I001-05001129 | IHEP1129 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Seleznev |  |  |  |  | 2.2 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 224 | 6 | no | Done | Ok | 3 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-09-28 | Mishagin |  |  |  |  | 1500 | 72.2119 | [12 empty] | Pending |  | 8
HCBLM_I001-05001130 | IHEP1130 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 256 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.4014 | [12 empty] | Pending |  | 8
HCBLM_I001-05001131 | IHEP1131 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.124 | [12 empty] | Pending |  | 8
HCBLM_I001-05001132 | IHEP1132 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.41 | [12 empty] | Pending |  | 8
HCBLM_I001-05001133 | IHEP1133 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-06 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 220 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 70.0684 | [12 empty] | Pending |  | 8
HCBLM_I001-05001134 | IHEP1134 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 232 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [18 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.946 | [12 empty] | Pending |  | 8
HCBLM_I001-05001135 | IHEP1135 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.1792 | [12 empty] | Pending |  | 8
HCBLM_I001-05001136 | IHEP1136 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-11-09 | Larionov |  |  |  |  | 32 | 2 | lar032_1;lar032_2 | 18 | 1093 | 234 | 6 | no | Done | Ok | 3 | 2006-09-12 | Seleznev |  |  |  |  | 0.4 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 70.9082 | [12 empty] | Pending |  | 8
HCBLM_I001-05001137 | IHEP1137 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 20 | 1100 | 227 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [18 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.856 | [12 empty] | Pending |  | 8
HCBLM_I001-05001138 | IHEP1138 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 223 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 70.4785 | [12 empty] | Pending |  | 8
HCBLM_I001-05001139 | IHEP1139 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 229 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.836 | [12 empty] | Pending |  | 8
HCBLM_I001-05001140 | IHEP1140 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 229 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 70.8008 | [12 empty] | Pending |  | 8
HCBLM_I001-05001141 | IHEP1141 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 234 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.636 | [12 empty] | Pending |  | 8
HCBLM_I001-05001142 | IHEP1142 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 21 | 1101 | 234 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.058 | [12 empty] | Pending |  | 8
HCBLM_I001-05001143 | IHEP1143 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-07 | Tikhonov |  |  |  |  | 38 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 252 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.9 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 69.7827 | [12 empty] | Pending |  | 8
HCBLM_I001-05001144 | IHEP1144 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 225 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.7 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.222 | [12 empty] | Pending |  | 8
HCBLM_I001-05001145 | IHEP1145 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 21 | 1102 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 72.0239 | [12 empty] | Pending |  | 8
HCBLM_I001-05001146 | IHEP1146 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 240 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.2 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.6 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 70.647 | [12 empty] | Pending |  | 8
HCBLM_I001-05001147 | IHEP1147 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-08 | Tikhonov |  |  |  |  | 10 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 | 2006-09-29 | Mishagin |  |  |  |  | 2000 | 0.5 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 67.4072 | [12 empty] | Pending |  | 8
HCBLM_I001-05001148 | IHEP1148 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Pending |  | 6 | [8 empty] | Done | Not Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 2.744 | [12 empty] | Pending |  | 8
HCBLM_I001-05001149 | IHEP1149 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 227 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Pending |  | 6 | [8 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.062 | [12 empty] | Pending |  | 8
HCBLM_I001-05001150 | IHEP1150 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 21 | 1101 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.966 | [12 empty] | Pending |  | 8
HCBLM_I001-05001151 | IHEP1151 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 241 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.678 | [12 empty] | Pending |  | 8
HCBLM_I001-05001152 | IHEP1152 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 18 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Pending |  | 6 | [8 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.338 | [12 empty] | Pending |  | 8
HCBLM_I001-05001153 | IHEP1153 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 223 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.844 | [12 empty] | Pending |  | 8
HCBLM_I001-05001154 | IHEP1154 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.384 | [12 empty] | Pending |  | 8
HCBLM_I001-05001155 | IHEP1155 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 227 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.3 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Pending |  | 7 | [20 empty] | Pending |  | 8
HCBLM_I001-05001156 | IHEP1156 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-20 | Seleznev |  |  |  |  | 37 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.83 | [12 empty] | Pending |  | 8
HCBLM_I001-05001157 | IHEP1157 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-08-20 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 221 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.532 | [12 empty] | Pending |  | 8
HCBLM_I001-05001158 | IHEP1158 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.3 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.7 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [18 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 73.926 | [12 empty] | Pending |  | 8
HCBLM_I001-05001159 | IHEP1159 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-11 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 19 | 1094 | 220 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 |  | Mishagin |  |  |  |  | 2000 | 0.2 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 73.4302 | [12 empty] | Pending |  | 8
HCBLM_I001-05001160 | IHEP1160 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-11 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 224 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 |  | Mishagin |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.04 | [12 empty] | Pending |  | 8
HCBLM_I001-05001161 | IHEP1161 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-20 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 239 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [18 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 73.858 | [12 empty] | Pending |  | 8
HCBLM_I001-05001162 | IHEP1162 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 226 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.184 | [12 empty] | Pending |  | 8
HCBLM_I001-05001163 | IHEP1163 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-20 | Seleznev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 225 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 1.3 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.726 | [12 empty] | Pending |  | 8
HCBLM_I001-05001164 | IHEP1164 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 221 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 |  | Mishagin |  |  |  |  | 2000 | 0.4 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 71.0449 | [12 empty] | Pending |  | 8
HCBLM_I001-05001165 | IHEP1165 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 18 | 1085 | 233 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.302 | [12 empty] | Pending |  | 8
HCBLM_I001-05001166 | IHEP1166 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.74 | [12 empty] | Pending |  | 8
HCBLM_I001-05001167 | IHEP1167 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 224 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [18 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.59 | [12 empty] | Pending |  | 8
HCBLM_I001-05001168 | IHEP1168 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 18 | 1085 | 226 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [18 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.602 | [12 empty] | Pending |  | 8
HCBLM_I001-05001169 | IHEP1169 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-20 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.7 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.748 | [12 empty] | Pending |  | 8
HCBLM_I001-05001170 | IHEP1170 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-11 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 238 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 | [12 empty] | Done | Cancelled | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 144.16 | [12 empty] | Pending |  | 8
HCBLM_I001-05001171 | IHEP1171 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.014 | [12 empty] | Pending |  | 8
HCBLM_I001-05001172 | IHEP1172 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-20 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 234 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.146 | [12 empty] | Pending |  | 8
HCBLM_I001-05001173 | IHEP1173 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 21 | 1101 | 222 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 |  | Grishin | [7 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.83 | [12 empty] | Pending |  | 8
HCBLM_I001-05001174 | IHEP1174 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-13 | Tereshin |  |  |  |  | 33 | 2 | ter033_1;ter033_2 | 20 | 1100 | 257 | 6 | ter033_3;ter033_4 | Done | Ok | 3 | 2006-09-14 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [8 empty] | Done | Ok | 5 | 2006-09-19 | Grishin |  |  |  |  | 4 | 2006-09-19 |  | Done | Ok | 6 |  | Mishagin |  |  |  |  | 2000 | 0.3 | Done | Ok | 7 | 2006-10-13 | Mishagin |  |  |  |  | 1500 | 69.8706 | [12 empty] | Pending |  | 8
HCBLM_I001-05001175 | IHEP1175 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 237 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.54 | [12 empty] | Pending |  | 8
HCBLM_I001-05001176 | IHEP1176 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 19 | 1085 | 221 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.766 | [12 empty] | Pending |  | 8
HCBLM_I001-05001177 | IHEP1177 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-20 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.84 | [12 empty] | Pending |  | 8
HCBLM_I001-05001178 | IHEP1178 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-20 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 255 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 66.936 | [12 empty] | Pending |  | 8
HCBLM_I001-05001179 | IHEP1179 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-11 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.368 | [12 empty] | Pending |  | 8
HCBLM_I001-05001180 | IHEP1180 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-11 | Seleznev |  |  |  |  | 1250 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 234 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [31 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.292 | [12 empty] | Pending |  | 8
HCBLM_I001-05001181 | IHEP1181 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-11 | Seleznev |  |  |  |  | 1.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.2 | 0.2 | 2000 | Cancelled | Cancelled | 4 | [31 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 78.272 | [12 empty] | Pending |  | 8
HCBLM_I001-05001182 | IHEP1182 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 223 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-18 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.068 | [12 empty] | Pending |  | 8
HCBLM_I001-05001183 | IHEP1183 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-14 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [31 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 72.258 | [12 empty] | Pending |  | 8
HCBLM_I001-05001184 | IHEP1184 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 222 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.388 | [12 empty] | Pending |  | 8
HCBLM_I001-05001185 | IHEP1185 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 241 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.996 | [12 empty] | Pending |  | 8
HCBLM_I001-05001186 | IHEP1186 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.4 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [31 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.24 | [12 empty] | Pending |  | 8
HCBLM_I001-05001187 | IHEP1187 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-13 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 220 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.18 | [12 empty] | Pending |  | 8
HCBLM_I001-05001188 | IHEP1188 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 233 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.766 | [12 empty] | Pending |  | 8
HCBLM_I001-05001189 | IHEP1189 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-20 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 233 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.926 | [12 empty] | Pending |  | 8
HCBLM_I001-05001190 | IHEP1190 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 237 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.748 | [12 empty] | Pending |  | 8
HCBLM_I001-05001191 | IHEP1191 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 226 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.688 | [12 empty] | Pending |  | 8
HCBLM_I001-05001192 | IHEP1192 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 229 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.9 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.158 | [12 empty] | Pending |  | 8
HCBLM_I001-05001193 | IHEP1193 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-14 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 228 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.55 | [12 empty] | Pending |  | 8
HCBLM_I001-05001194 | IHEP1194 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-19 | Seleznev |  |  |  |  | 20.5 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 220 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-02-25 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.332 | [12 empty] | Pending |  | 8
HCBLM_I001-05001195 | IHEP1195 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-11 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 230 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-25 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.362 | [12 empty] | Pending |  | 8
HCBLM_I001-05001196 | IHEP1196 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-12 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-15 | Larionov |  |  |  |  | 34 | 2 | lar034_1;lar034_2 | 19 | 1094 | 256 | 6 | lar034_3;lar034_4 | Done | Ok | 3 | 2006-09-20 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.842 | [12 empty] | Pending |  | 8
HCBLM_I001-05001197 | IHEP1197 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-11 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 16 | 1085 | 230 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.6 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.754 | [12 empty] | Pending |  | 8
HCBLM_I001-05001198 | IHEP1198 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-14 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-22 | Larionov |  |  |  |  | 35 | 2 | lar035_1;lar035_2 | 17 | 1085 | 251 | 6 | lar035_3;lar035_4 | Done | Ok | 3 | 2006-09-22 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.242 | [12 empty] | Pending |  | 8
HCBLM_I001-05001199 | IHEP1199 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-20 | Seleznev |  |  |  |  | 2.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.3 | 1089 | 247 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.5 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 73.574 | [12 empty] | Pending |  | 8
HCBLM_I001-05001200 | IHEP1200 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.2 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.5 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.258 | [12 empty] | Pending |  | 8
HCBLM_I001-05001201 | IHEP1201 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 25 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.2 | 1089 | 228 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.114 | [12 empty] | Pending |  | 8
HCBLM_I001-05001202 | IHEP1202 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.1 | 1089 | 220 | 6 | lar037_1;lar037_2 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.9 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.572 | [12 empty] | Pending |  | 8
HCBLM_I001-05001203 | IHEP1203 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.2 | 1089 | 240 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 1.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.452 | [12 empty] | Pending |  | 8
HCBLM_I001-05001204 | IHEP1204 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 10.5 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.3 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.09 | [12 empty] | Pending |  | 8
HCBLM_I001-05001205 | IHEP1205 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.5 | 1089 | 229 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.51 | [12 empty] | Pending |  | 8
HCBLM_I001-05001206 | IHEP1206 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.9 | 1089 | 227 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.324 | [12 empty] | Pending |  | 8
HCBLM_I001-05001207 | IHEP1207 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.9 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.9 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.91 | [12 empty] | Pending |  | 8
HCBLM_I001-05001208 | IHEP1208 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.7 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.102 | [12 empty] | Pending |  | 8
HCBLM_I001-05001209 | IHEP1209 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 2.5 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 19.1 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.25 | [12 empty] | Pending |  | 8
HCBLM_I001-05001210 | IHEP1210 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 19.2 | 1089 | 222 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.352 | [12 empty] | Pending |  | 8
HCBLM_I001-05001211 | IHEP1211 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-20 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.8 | 1089 | 234 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.4 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.31 | [12 empty] | Pending |  | 8
HCBLM_I001-05001212 | IHEP1212 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.4 | 1089 | 224 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.472 | [12 empty] | Pending |  | 8
HCBLM_I001-05001213 | IHEP1213 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 2 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 18.1 | 1089 | 223 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.18 | [12 empty] | Pending |  | 8
HCBLM_I001-05001214 | IHEP1214 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.8 | 1089 | 230 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.7 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.9 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.328 | [12 empty] | Pending |  | 8
HCBLM_I001-05001215 | IHEP1215 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-22 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.8 | 1089 | 234 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.5 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.9 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.342 | [12 empty] | Pending |  | 8
HCBLM_I001-05001216 | IHEP1216 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.9 | 1089 | 240 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 67.654 | [12 empty] | Pending |  | 8
HCBLM_I001-05001217 | IHEP1217 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.7 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.7 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.932 | [12 empty] | Pending |  | 8
HCBLM_I001-05001218 | IHEP1218 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 1.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.9 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.974 | [12 empty] | Pending |  | 8
HCBLM_I001-05001219 | IHEP1219 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-22 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.6 | 1089 | 227 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.9 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.37 | [12 empty] | Pending |  | 8
HCBLM_I001-05001220 | IHEP1220 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.8 | 1089 | 232 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.914 | [12 empty] | Pending |  | 8
HCBLM_I001-05001221 | IHEP1221 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.9 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.922 | [12 empty] | Pending |  | 8
HCBLM_I001-05001222 | IHEP1222 | IHEP | Manufacturing |  | Done | Ok | 1 | 1902-07-09 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.6 | 1089 | 225 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.296 | [12 empty] | Pending |  | 8
HCBLM_I001-05001223 | IHEP1223 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-22 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.6 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.864 | [12 empty] | Pending |  | 8
HCBLM_I001-05001224 | IHEP1224 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.5 | 1089 | 229 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.572 | [12 empty] | Pending |  | 8
HCBLM_I001-05001225 | IHEP1225 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.9 | 1089 | 240 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.9 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.716 | [12 empty] | Pending |  | 8
HCBLM_I001-05001226 | IHEP1226 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.9 | 1089 | 229 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.942 | [12 empty] | Pending |  | 8
HCBLM_I001-05001227 | IHEP1227 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.5 | 1089 | 235 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.8 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.508 | [12 empty] | Pending |  | 8
HCBLM_I001-05001228 | IHEP1228 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 18.1 | 1089 | 220 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.45 | [12 empty] | Pending |  | 8
HCBLM_I001-05001229 | IHEP1229 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 18.7 | 1089 | 224 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.4 | 0.2 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.814 | [12 empty] | Pending |  | 8
HCBLM_I001-05001230 | IHEP1230 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-22 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.5 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.632 | [12 empty] | Pending |  | 8
HCBLM_I001-05001231 | IHEP1231 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 2.2 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.9 | 1089 | 234 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.984 | [12 empty] | Pending |  | 8
HCBLM_I001-05001232 | IHEP1232 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-22 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.1 | 1089 | 220 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.8 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 72.052 | [12 empty] | Pending |  | 8
HCBLM_I001-05001233 | IHEP1233 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-22 | Seleznev |  |  |  |  | 7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.2 | 1089 | 234 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.9 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.8 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.496 | [12 empty] | Pending |  | 8
HCBLM_I001-05001234 | IHEP1234 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.9 | 1089 | 237 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.84 | [12 empty] | Pending |  | 8
HCBLM_I001-05001235 | IHEP1235 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.4 | 1089 | 225 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.5 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.958 | [12 empty] | Pending |  | 8
HCBLM_I001-05001236 | IHEP1236 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-22 | Seleznev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.3 | 1089 | 240 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.4 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.332 | [12 empty] | Pending |  | 8
HCBLM_I001-05001237 | IHEP1237 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.7 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.07 | [12 empty] | Pending |  | 8
HCBLM_I001-05001238 | IHEP1238 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.8 | 1089 | 223 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [54 empty] | Pending |  | 8
HCBLM_I001-05001239 | IHEP1239 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Seleznev |  |  |  |  | 120 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 18 | 1084 | 223 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.792 | [12 empty] | Pending |  | 8
HCBLM_I001-05001240 | IHEP1240 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-22 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.3 | 1089 | 220 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 66.388 | [12 empty] | Pending |  | 8
HCBLM_I001-05001241 | IHEP1241 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.7 | 1089 | 256 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.704 | [12 empty] | Pending |  | 8
HCBLM_I001-05001242 | IHEP1242 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 1257 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.1 | 1089 | 224 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 71.808 | [12 empty] | Pending |  | 8
HCBLM_I001-05001243 | IHEP1243 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-22 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.2 | 1089 | 221 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.434 | [12 empty] | Pending |  | 8
HCBLM_I001-05001244 | IHEP1244 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 1.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.8 | 1089 | 221 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-27 | Seleznev |  |  |  |  | 0.5 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.62 | [12 empty] | Pending |  | 8
HCBLM_I001-05001245 | IHEP1245 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.1 | 1089 | 255 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.8 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.79 | [12 empty] | Pending |  | 8
HCBLM_I001-05001246 | IHEP1246 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.7 | 1089 | 224 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.956 | [12 empty] | Pending |  | 8
HCBLM_I001-05001247 | IHEP1247 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Galyaev |  |  |  |  | 270 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.432 | [12 empty] | Pending |  | 8
HCBLM_I001-05001248 | IHEP1248 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Galyaev |  |  |  |  | 9 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 221 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.464 | [12 empty] | Pending |  | 8
HCBLM_I001-05001249 | IHEP1249 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Seleznev |  |  |  |  | 840 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.376 | [12 empty] | Pending |  | 8
HCBLM_I001-05001250 | IHEP1250 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-21 | Seleznev |  |  |  |  | 1.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-26 | Larionov |  |  |  |  | 36 | 2 | lar036_1;lar036_2 | 17.6 | 1089 | 220 | 6 | lar036_3;lar036_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.26 | [12 empty] | Pending |  | 8
HCBLM_I001-05001251 | IHEP1251 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.3 | 1089 | 241 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.36 | [12 empty] | Pending |  | 8
HCBLM_I001-05001252 | IHEP1252 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 225 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.3 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.808 | [12 empty] | Pending |  | 8
HCBLM_I001-05001253 | IHEP1253 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 6.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 16.7 | 1089 | 225 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.8 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.604 | [12 empty] | Pending |  | 8
HCBLM_I001-05001254 | IHEP1254 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.2 | 1089 | 220 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.7 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.742 | [12 empty] | Pending |  | 8
HCBLM_I001-05001255 | IHEP1255 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 6.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17 | 1089 | 224 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.8 | [12 empty] | Pending |  | 8
HCBLM_I001-05001256 | IHEP1256 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.3 | 1089 | 220 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.8 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.152 | [12 empty] | Pending |  | 8
HCBLM_I001-05001257 | IHEP1257 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 228 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.3 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.926 | [12 empty] | Pending |  | 8
HCBLM_I001-05001258 | IHEP1258 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17 | 1089 | 230 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.5 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.022 | [12 empty] | Pending |  | 8
HCBLM_I001-05001259 | IHEP1259 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 235 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.794 | [12 empty] | Pending |  | 8
HCBLM_I001-05001260 | IHEP1260 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 18.4 | 1089 | 232 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.6 | 0.2 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.978 | [12 empty] | Pending |  | 8
HCBLM_I001-05001261 | IHEP1261 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 231 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.718 | [12 empty] | Pending |  | 8
HCBLM_I001-05001262 | IHEP1262 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 238 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.1 | [12 empty] | Pending |  | 8
HCBLM_I001-05001263 | IHEP1263 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 224 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.6 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.134 | [12 empty] | Pending |  | 8
HCBLM_I001-05001264 | IHEP1264 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 62 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 258 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.6 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.28 | [12 empty] | Pending |  | 8
HCBLM_I001-05001265 | IHEP1265 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 18 | 1089 | 220 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.8 | 0.9 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.9 | [12 empty] | Pending |  | 8
HCBLM_I001-05001266 | IHEP1266 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.4 | 1089 | 221 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 67.634 | [12 empty] | Pending |  | 8
HCBLM_I001-05001267 | IHEP1267 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 16.9 | 1089 | 229 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.4 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.984 | [12 empty] | Pending |  | 8
HCBLM_I001-05001268 | IHEP1268 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.3 | 1089 | 235 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.212 | [12 empty] | Pending |  | 8
HCBLM_I001-05001269 | IHEP1269 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 16.8 | 1089 | 220 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.5 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.362 | [12 empty] | Pending |  | 8
HCBLM_I001-05001270 | IHEP1270 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Seleznev |  |  |  |  | 8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 233 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.492 | [12 empty] | Pending |  | 8
HCBLM_I001-05001271 | IHEP1271 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 228 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70 | [12 empty] | Pending |  | 8
HCBLM_I001-05001272 | IHEP1272 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-10-02 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Pending |  | 2 |  | Larionov | [12 empty] | Done | Ok | 3 | 2006-10-04 | Seleznev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.742 | [12 empty] | Pending |  | 8
HCBLM_I001-05001273 | IHEP1273 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Seleznev |  |  |  |  | 7 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 17 | 1084 | 230 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.85 | [12 empty] | Pending |  | 8
HCBLM_I001-05001274 | IHEP1274 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Seleznev |  |  |  |  | 13 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 17 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.5 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.52 | [12 empty] | Pending |  | 8
HCBLM_I001-05001275 | IHEP1275 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Seleznev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 17 | 1084 | 235 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | [23 empty] | Pending |  | 8
HCBLM_I001-05001276 | IHEP1276 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.772 | [12 empty] | Pending |  | 8
HCBLM_I001-05001277 | IHEP1277 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 67.79 | [12 empty] | Pending |  | 8
HCBLM_I001-05001278 | IHEP1278 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.6 | 0.6 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.7 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.326 | [12 empty] | Pending |  | 8
HCBLM_I001-05001279 | IHEP1279 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 17 | 1084 | 224 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.6 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.742 | [12 empty] | Pending |  | 8
HCBLM_I001-05001280 | IHEP1280 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 3.8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 227 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.6 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.588 | [12 empty] | Pending |  | 8
HCBLM_I001-05001281 | IHEP1281 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 8.2 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 240 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-12 | Seleznev |  |  |  |  | 0.6 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.728 | [12 empty] | Pending |  | 8
HCBLM_I001-05001282 | IHEP1282 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.3 | 1089 | 239 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-24 | Seleznev |  |  |  |  | 0.4 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.9 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.122 | [12 empty] | Pending |  | 8
HCBLM_I001-05001283 | IHEP1283 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 18 | 1084 | 234 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.718 | [12 empty] | Pending |  | 8
HCBLM_I001-05001284 | IHEP1284 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.49 | [12 empty] | Pending |  | 8
HCBLM_I001-05001285 | IHEP1285 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.374 | [12 empty] | Pending |  | 8
HCBLM_I001-05001286 | IHEP1286 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 223 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.352 | [12 empty] | Pending |  | 8
HCBLM_I001-05001287 | IHEP1287 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.6 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.702 | [12 empty] | Pending |  | 8
HCBLM_I001-05001288 | IHEP1288 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 228 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.8 | [12 empty] | Pending |  | 8
HCBLM_I001-05001289 | IHEP1289 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 5.2 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.5 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.402 | [12 empty] | Pending |  | 8
HCBLM_I001-05001290 | IHEP1290 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.6 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.522 | [12 empty] | Pending |  | 8
HCBLM_I001-05001291 | IHEP1291 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.2 | 1089 | 231 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.7 | 0.7 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.932 | [12 empty] | Pending |  | 8
HCBLM_I001-05001292 | IHEP1292 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 16.9 | 1089 | 220 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-29 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.418 | [12 empty] | Pending |  | 8
HCBLM_I001-05001293 | IHEP1293 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.8 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.2 | 1089 | 230 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.8 | 0.8 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.242 | [12 empty] | Pending |  | 8
HCBLM_I001-05001294 | IHEP1294 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 225 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-05 | Galyaev |  |  |  |  | 0.6 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.16 | [12 empty] | Pending |  | 8
HCBLM_I001-05001295 | IHEP1295 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Seleznev |  |  |  |  | 4 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 251 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.6 | 0.5 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.5 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.172 | [12 empty] | Pending |  | 8
HCBLM_I001-05001296 | IHEP1296 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 1200 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 17 | 1084 | 240 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.268 | [12 empty] | Pending |  | 8
HCBLM_I001-05001297 | IHEP1297 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-25 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-09-28 | Larionov |  |  |  |  | 37 | 2 | lar037_1;lar037_2 | 17.2 | 1089 | 220 | 6 | lar037_3;lar037_4 | Done | Ok | 3 | 2006-09-28 | Seleznev |  |  |  |  | 0.7 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-09 | Mishagin |  |  |  |  | 1500 | 0.9 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.082 | [12 empty] | Pending |  | 8
HCBLM_I001-05001298 | IHEP1298 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 17 | 1084 | 220 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-03 | Galyaev |  |  |  |  | 0.5 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.806 | [12 empty] | Pending |  | 8
HCBLM_I001-05001299 | IHEP1299 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-09-28 | Galyaev |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-10-02 | Larionov |  |  |  |  | 38 | 2 | lar038_1;lar038_2 | 16 | 1084 | 226 | 6 | lar038_3;lar038_4 | Done | Ok | 3 | 2006-10-04 | Galyaev |  |  |  |  | 0.5 | 0.4 | 2000 | Cancelled | Cancelled | 4 | [20 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.778 | [12 empty] | Pending |  | 8
[80 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.832
[69 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1
[69 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.784
[69 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.2
[69 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 71.444
[80 empty] | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.408
[69 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.4 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.98
[69 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 70.138
[69 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 69.818
[69 empty] | Done | Ok | 6 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 0.1 | Done | Ok | 7 | 2006-11-10 | Mishagin |  |  |  |  | 1500 | 68.198
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.21
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 74.626
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.154
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.28
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.252
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.33
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.258
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.106
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.194
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.664
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.558
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.19
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.208
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.198
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.572
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.278
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.414
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.286
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.23
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.138
Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | [59 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | -0.07
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.312
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.32
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.192
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.632
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.286
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.28
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.482
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.726
Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | [59 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.312
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.436
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.388
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.244
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.292
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.316
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.636
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.19
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | -0.036
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.194
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.294
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.286
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.194
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.306
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.33
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.21
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.246
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.118
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.38
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.266
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.258
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.366
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | -0.116
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.674
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.256
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.218
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.248
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.224
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.164
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.196
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.61
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.136
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.1
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.42 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.47
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.836
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.054 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 68.272
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 69.842
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.672
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.404 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.816
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.316
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.31 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.548
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 69.988
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 69.236
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.29 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.008
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.44 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 69.94
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.556 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.118
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.226
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.184
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.524
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.22 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.036
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.698
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.804
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.326
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.478
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.586 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.712
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.84 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.344
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 67.402
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.6329
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.688
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.51
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.338
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.406
[80 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 70.06
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.47 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 71.058
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.402 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 72.268
[69 empty] | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 0.312 | Done | Ok | 7 | 21/12/2006 | Tikhonov |  |  |  |  | 1500 | 69.4

## 2.Standard Steps
Standard Steps
Beam Loss Monitor: Ionization Chamber
Standard Step ID | Other identifier | Sequence number | Description | Provided by | Expected by | Responsible | Aplicable standard | Chapter | Comment
MTF013476 |  | 1 | Assembly | IHEP | CERN | Asanov,Seleznev |  |  | Head's welding procedure; Cleaning;Assembly;Current measurement
MTF013477 |  | 2 | Vacuum stand | IHEP | CERN | Larionov |  |  | Vacuum procedure; 36 IC per batch
MTF013478 |  | 3 | Leakage current measurement at IHEP by IHEP | IHEP | CERN | Seleznev |  |  | CERN Picometer noice +/- 0.3 pA
MTF013479 |  | 4 | Leakage current measurement at IHEP by CERN | CERN | CERN | Ferioli |  |  | CERN Picometer noice +/- 0.3 pA
MTF013480 |  | 5 | Shipping | IHEP,CERN | CERN | Seleznev,Grishin |  |  | 360 IC per month
MTF013481 |  | 6 | Leakage current measurement at CERN | CERN,IHEP | CERN | Ferioli,Sleptsov |  |  | CERN Picometer noice +/- 0.3 pA
MTF013482 |  | 7 | Source test | CERN,IHEP | CERN | Ferioli,Sleptsov |  |  | For 6 HV values measure 6 Current values
MTF013483 |  | 8 | Source test at tunnel | CERN,IHEP | CERN | Ferioli,Sleptsov |  |  | archive:measure the parameters of IC time to time

## 3.Properties
Properties
Description | Code | Unit | Type | Category | Default | Comment
Leakage Current (Assemb) | P11483 | pA | NUM
Head Producer | P11446 |  | CHAR |  |  | DG Technology(Italy),Aerovac(UK),Alca(Italy),IHEP(Russia)
High Voltage (Assembly) | P11448 | V | NUM
Leakage Current (IHEP) | P11449 | pA | NUM
Batch Number | P11450 |  | NUM |  |  | Number of batch on the vacuum stand
Gas Bottle Number | P11468 |  | NUM
Spectrum (before filling) | P11453 |  | CHAR
Temperature (filling) | P11451 | °C | NUM
Pressure (filling) | P11452 | mbar | NUM
Temperature (heating) | P11454 | °C | NUM
Duration (heating) | P11456 | h | NUM
Spectrum (after filling) | P11455 |  | CHAR
Leakage Current_v | P11457 | pA | NUM |  |  | Measurement after vacuum stand
Leakage Current_f | P11458 | pA | CHAR |  |  | Measurement after el. box soldering
Lorry Number | P11461 |  | NUM |  |  | Number of lorry from Protvino
IHEP Departure Date | P11462 |  | CHAR
CERN Arrival Date | P11463 |  | CHAR
High Voltage1 | P11471 | V | NUM
Current1 | P11477 | pA | NUM
High Voltage2 | P11472 | V | NUM
Current2 | P11478 | pA | NUM
High Voltage3 | P11473 | V | NUM
Current3 | P11479 | pA | NUM
High Voltage4 | P11474 | V | NUM
Current4 | P11480 | pA | NUM
High Voltage5 | P11475 | V | NUM
Current5 | P11481 | pA | NUM
High Voltage6 | P11476 | V | NUM
Current6 | P11482 | pA | NUM
Current | P11470 | pA | NUM |  |  | Rad source current measurements
Dose Rate | P11467 | millisievert par heure | NUM
Source Activity (Surface) | P11469 | GBq | NUM
High Voltage (Leak-I-I) | P11459 | V | NUM
High Voltage (Leak-C-C) | P11464 | V | NUM
High Voltage (Leak-I-C) | P11460 | V | NUM
High Voltage (Tunnel) | P11466 | V | NUM
Leakage Current (CERN) | P11465 | pA | NUM
Source Activity (Tunnel) | P11447 | GBq | NUM

## Personal Information
Personal Information
EDMSlogin | JJGRAS
EDMSGroup | AB-BI-BLM-MTF
Created On | 2005-01-01