# EDMS No.:  578716

# DR-02061/AB/LHC

# DO-21533/AB/LHC

# Technical Specification

# Alumine Insulator

# Cf. Technical drawing CDD No.: LHCBLM__0005

# The General Technical Specification, EDMS 581010, in its actual version, constitutes an integral part of this document.

# Deliverable

Number of pieces:

Baseline:	7110
Option:	1680

# Material Al2O3 97%

All dimensions and tolerances are given on the technical design drawing. Special care has to be taken that the diameter of the central hole stays between 12.4 and 12.7 mm.

The ceramics acts as electrical insulator in an ultra high vacuum and must not outgas. Special care has to be taken to avoid any contamination of the surface which could impair the application for UHV (see the General Technical Specification section 2, and especially 2.1.1). The pieces will undergo a baking procedure after they have been delivered to CERN to reduce outgasing. They have to withstand 1000°C during 2 hours and the mounting and falling of the temperature at a rate of 100°C per hour.

The final cleaning for UHV and the baking procedure is not part of the deliverable.

# The raw material will be supplied by the contractor.

# Specification and test of electric resistivity of the insulator

Two adjacent holes of the workpiece with the diameter of 4.5 mm have to be equipped with two screws of 4 mm diameter, a plane washer at each side and held by a nut each. The electric resistivity between these two screws at 3000 V in vacuum has to stay below 3 1017 Ohm (10-14 A).

# Specification of the material

## General

### Scope

The insulators are used for high voltages (3000 V) and in ultra-high vacuum. They must, therefore, not be porous and be completely free from degassing.

## Applicable Documents

This specification makes reference to the following publications:

- CERN/FC/1814-II	(1975)		General conditions of CERN contracts

- ASTM 	C20

- ASTM	C351

- ASTM	C372

- ASTM	C408

- ASTM	C528

- ASTM	C565

- ASTM 	D116

- ASTM	D150

- ASTM	D256

- ASTM	D257

- ASTM	E18

- ASTM	F77

## General Characteristics

### Material

Alumina ceramic Al2O3

Minimum alumina content 96 to 98% - SiO2 remainder.

### Others Characteristics

See annex 1.

### Surface Finish

Surface faults, e.g. pitting, iron marks, blistering, scoring, cracking, excess material, scaling, etc. will not be accepted if they impair any of the specification requirements.

### Fabrication

The insulators shall be produced by isostatic pressure.

### Packing

The packing used must be strong enough to protect the surface finish of the various shapes of ceramic material supplied, both during transport and storage.

## Quality Control

### General

The general conditions of CERN/FC/1814-II contracts are applicable.

### Inspection Stages

#### Qualification inspection

##### Phase de qualification

The offer shall also contain full details of the materials used and the documents that have been applied. The offer shall indicate the points where a divergence exists between the material proposed and that specified by CERN.

The bidder shall submit proposals concerning the schedule of tests to be carried out on the material, and provide details of the guarantees.

##### Failures

One or several parameter failures may provoke a qualification refusal without appeal by the tender.

For each batch delivered the following tests shall be carried out by the manufacturer

- mechanical characteristics,

- electrical characteristics,

- thermal characteristics,

- dimensions

- outgasing

- fissures and cracks controls.

CERN reserves the right to delegate a representative to be present at these tests. The results of the tests shall be communicated to CERN at the time of delivery at the latest.

#### Quality conformance inspection

The constant quality control is applied on each partial or complete delivery of the product.

##### Inspection method

At CERN an incoming inspection will be performed to verify the requirements of the technical specifications.

##### Failures

Document CERN/FC/1814-II, clause 13.

# Annex 1

## Others Characteristics

| PROPERTIES | TEST METHOD | UNITS | VALUES | PROPOSED VALUES |
| --- | --- | --- | --- | --- |
| Alumina (Al2O3) content | Spectrographic analysis | % | > 96 |  |
| Apparent density | ASTM C20, F77 | g/cm3 | 3.8 – 4 |  |
| Flexural strength | ASTM D116, ACMA test 2 | kg/cm2 | 3200 |  |
| Compressive strength | ASTM C528 - D116 | kg/cm2 | 21100 – 29800 |  |
| Tensile strength | ASTM C565 | kg/cm2 | 1410 - 2460 |  |
| Impact strength | ASTM D116 D256 Charpy | m kg | 0.08 - 0.09 |  |
| Modulus of elasticity (Young's modulus) | Resonance frequency 1011 | dynes/cm2 | 33 – 39 |  |
| Hardness | ASTM E18 Rockwell 45N |  | 78 – 90 |  |
| Impenetrability (liquids) | ASTM D116 |  | Leak tight |  |
| Impenetrability (gases) | Helium mass spectrometer |  | Leak tight |  |
| Open porosity |  | % | 0 |  |
| Working temperature |  | °C | max. 1600 |  |
| Thermal expansion coef. | - 200° to 25°C | ASTM D116 - C372 | 10-6/°C | 3 - 3.5 |  |
|  | 25° to 200°C |  |  | 5.3 - 6.8 |  |
|  | 25° to 500°C |  |  | 6.1 - 7.6 |  |
|  | 25° to 800°C |  |  | 6.9 - 8.1 |  |
|  | 25° to 1000°C |  |  | 7.4 - 8.6 |  |
| Thermal shock resistance |  |  | Good |  |
| Break down voltage | ASTM D116 | kV/mm | 7.1 - 10.2 |  |
| Electric resistivity |  | ( cm | > 1015 |  |