# List of tasks and occupation time of the vacuum stand

# (SEM)

# Description of Equipment

| Equipment | Description |
| --- | --- |
| Vacuum stand for 36 SEM | Designed and constructed by IHEP |
| Pump station | Station on base TMP 56 l/s supplied by CERN |
| Ion pump | Ion pump 100 l/s supplied by IHEP |
| Leak detector | PhoeniXL300 supplied by CERN |
| Gas analyzer | Prizma QMS200F2 supplied by CERN |
| Pressure gauge controller | MaxiGauge TPG256A supplied by CERN |
| Temperature regulater | TR on 24 channels supplied by CERN |
| Date Acquisition unit | Agilent 34970A supplied by CERN |
| Heating container | Container use for heating IC and creating of He overpressure around IC. |
| N2 | N2 of 99,999 % purity |

# № (1 (3)   -   4 hours

- Installation of the SEM on the stand.

- Remove the rest of the copper tubes from the “Sagana” connectors ; cut off the  cooper tubes for release of  the nuts.

- Write  the numbers of the SEM into the data base (Dbase-vac).

- Wash the Sagana inserts rings with alcohol; put in nuts with inserts on the SEM cooper tubes.

- Introduce the SEM copper tubes in the stand “Sagana” connectors and fix ones tightly.

- Fix the 36 thermocouple with scotch tape on the SEM and connect 12 of the couples to the temperature regulation unit.

- Switch on the vacuum pump station and pump stand with all SEM.

- Leakage test of welds and removable connections of SEM using the procedure of helium overpressure.

- Put the heating containers on SEM’s and tighten with the top covers.

-Connect the He supply tubes to the heating containers (for 9 containers).

Pump out the air from 9 containers and fill with He up to a pressure of somewhat higher   than atmospheric.

  - Control the tightness by observing the He leak rate during of 3 min: in case of the leak detection, close successively the valves of the SEM for determine the leaking SEM.

  - For the each next 9 containers apply the same procedure.

# (Leak test executes with help leak detector)

# Leaks of more than 1E-10 l mbar/s are not allowed.

- Preparation of the heating.

  - Disconnect the He supply tubes from the containers;

  - Dismount the covers of the containers and remove the rubber o-rings from the covers,

  - Dismount the electronic units of the gas analyser and gauges (the valve of gas analyser must be closed)

  - Dismount permanent magnet from ion pump and put on heating jacket,

  - Heat-insulate all the SEM valves and top covers of containers.

- Heating (37 hours)

- Start the stand and SEM heating from 200 C  to 2200 C  and 50 C/hour with help programmator of TR (4 hours).

-  Temperature is kept at  2200 C  (6 hours) with help TR;

- Write into the data base the temperature of the heating for each SEM.

# -  Cool-down stand to 1500 C (1 hours) and this temperature is kept during SEM activation. The SEM temperature is kept 2200 C during cooling stand.

# -  Start the NEG activation, the SEM temperature is increase from 2200 C to 3500 C, (500 C/h) (2.6 hours), time activation is 20 hours.

At the end of SEM activation open gas analyser head valve  for its pumping and switch on heater of the head (hold 150÷200C during 1hour);

# -  Cool-down SEM and stand to 1000 C (3 hours).

# -  Dismount ion pump heater and mount magnets.

# -  Switch on the ion pump and switch off the pump station.

№(5,6)  -  4 hours

5. Leakage test of the welds and the removable connections of the SEM

after heating and activation, elimination of possible leaks.

# (Leak test executes with help gas analyzer)

  - Place the rubber o-rings into the top covers of the heating containers.

  - Close the heating containers and tighten the sealing at the top covers.

  - Connect the He supply tubes to the heating containers (on 9 containers).

  - Switch on the gas analyzer in “leak mode”.

  - Pump out the air of the containers and fill with He up to a pressure of somewhat higher than atmospheric.

  - Control the tightness by observing possible He leaks during 3 min: in case of the appearance of a leak, close successively the valves of the SEM to determine the leaking SEM.

  - Repeat control tightness procedure for all SEM.

# Leaks of more than 1E-10 l mbar/s are not allowed. Data entered into the data base.

-    Measure and record residual gas spectrum (percentage composition and partial   pressure) to folder ‘Dbase-vac’.

-    Close all SEM valves.

- Dismounting of SEM.

-    Pinch off the cooper tubes and dismount the SEMs from the stand .

-    Open all SEM valves on the stand.

  - Verify the tightness of the cut-off copper tubes which remained on the vacuum stand. (The tightness on the stand side gives a good indication of the tightness of the SEM side of the pinch-off).

-    Switch off the gas analyser and close the valve at the head of the gas analyser.

-   Close ion pump valve and switch of ion pump.

  - Equip the sharp rest of the SEM copper tube with a shrink tube.