30 August 2007

# List of tasks and occupation time of the vacuum stand

# (SEM)

# Description of Equipment

| Equipment | Description |
| --- | --- |
| Vacuum stand for 18-36 SEM | Designed and constructed by IHEP; only 18 SEMs produced at the same time |
| Pump station | Station on base TMP 56 l/s supplied by CERN |
| Pump station | Station on base TMP 170 l/s supplied by CERN |
| Ion pump | Ion pump 100 l/s supplied by IHEP |
| Leak detector | PhoeniXL300 supplied by CERN |
| Gas analyzer | Prizma QMS200F2 supplied by CERN |
| Pressure gauge controller | MaxiGauge TPG256A supplied by CERN |
| Temperature regulater | TR on 24 channels supplied by CERN |
| Date Acquisition unit | Agilent 34970A supplied by CERN |
| Heating container | Container use for heating SEM |

- Preparation procedure made:

  - – ultimate pressure is 3E-10 mbar (measured by gas analyser),

  - –desorbtion rate is 2E-13 l*mbar/s*cm2,

  - –the main rest gas is H2.

(15-00)-              Installation of the SEM on the stand.

- Remove the rest of the copper tubes from the “Sagana” connectors ; cut off the  cooper tubes for release of  the nuts.

- Write  the numbers of the SEM into the data base (Dbase-vac).

- Wash the Sagana insert rings with alcohol; put in nuts with inserts on the SEM cooper tubes.

- Introduce the SEM copper tubes in the stand “Sagana” connectors and fix ones tightly.

- Fix the 18 thermocouples with a wire on the SEM and connect 6 of the couples to the temperature regulation unit.

- Switch on the vacuum pump station and pump stand with all SEM.

(16-00)-    Leakage test of welds and removable connections of SEM using

the procedure  of helium overpressure.

- Put the heating containers on SEM’s and tighten with the top covers.

- Connect the He supply tubes to the heating containers (for 9 containers).

- Pump out the air from 9 containers and fill with He up to a pressure of somewhat higher   than atmospheric.

- Control the tightness by observing the He leak rate during of 3 min: in case of the leak detection, close successively the valves of the SEM for determine the leaking SEM.

- For each of the next 9 containers apply the same procedure (Leak test is execute with help leak detector).

# Leaks of more than 1E-10 l mbar/s are not allowed.

(18-00)-                Preparation of the heating.

- Disconnect the He supply tubes from the containers;

- Dismount the covers of the containers and remove the rubber o-rings from the covers,

- Dismount control unit of the analyser. (The valve of gas analyser must be closed),

- Heat-insulate all the SEM valves. Continue the pumping of the stand.

Heating. Leakage test of the welds and the removable connections of the SEM

after heating and activation, elimination of possible leaks.

(20-00)-  1. Start the stand and SEM heating from 200 C to 2200 C and 50 C/hour

with help  programmator of TR (4 hours).

(The pressure not must be more than a few 1E-4mbar)

# 2. Switch on the heating of the ion pump (to1000 C) and the gas analyser (1900).

# The valves of the pump (V5) and gas analyser (V4) must be opened. Ion pump                      must be switched on. The valve V2 must be closed.

Temperature is kept at 2200 C  (12 hours) with help TR;

# (12-00)-  3. Cool-down the stand to 1500 C and this temperature is kept during

# SEM activation. The SEM temperature is kept 2200 C during cooling stand.

(14-00)-  5. Input He successively to the each heating container for control tightness

of the SEM with help leak detector. SEM temperature is 2200 C.

(The valve of back  pump must be closed).

(15-00)- 7. Start the NEG activation (The start pressure not must be more  than 1E-6mbar)                    the SEM temperature is increase from 2200 C to 3500 C, (500 C/h)

(2.6 hours), time activation is   14 hours.

# (07-30)-   8. Cool-down SEM to 3000 C and it is kept about 5 hours, cool-down stand to TR0 C.

# (08-00)-  9. Switch off heating of analyser and ion pump.

# (08-30)-10. Switch on gas analyser. Record spectrum (partial pressure, analog current)

# versus  time. Open V2 valve. Record spectrum up to closing all SEM valves.

# (10-00)- 11. Check tightness of SEM with help analyser (observe for N2/CO).

# The valve of pump station must be closed. Checking execute successively

# SEM valve closing.

# Open pump station valve

# Check tightness of SEM with leakdetector with valve of ion pump (V5) closed

# (13-00)- 13. Cool-down of SEM from 3000 C at a rate of 1000 C/hour.

# (14-00)-                           Dismounting of SEM.

# At the temperature of SEM about 2000 C and pressure <1E-9 mbar close all SEM valves.

# Dismount heating containers.

# Pinch of all SEMs.

# Switch off analyser, close V2 valve.

- Open all SEM valves and verify the tightness of the cut-off copper tubes which remained on the vacuum stand. (The tightness on the stand side gives a good indication of the tightness of the SEM side of the pinch-off). Check the tightness with help of leak detector.

- Close pump station valve.

# Cover the sharp rest of the SEM copper tube with shrinking tube.