## Installation overview
BLECF installation (tunnel electronic)
| ARC |  |  |  |  | straight section
| BLECF |  |  |  |  | BLECF |  |  |  |  | Source test planning
| Dismounted | Calibrated | Installed | Checked and working | Comments | Dismounted | Calibrated | Installed | Checked and working | Comments
1 LEFT | 22 | 22 | 22 | 22 |  | 19 | 19 | 19 | 19 |  | 4
1 RIGHT | 23 | 23 | 23 | 23 |  | 19 | 19 | 19 | 19 |  | 3
2 LEFT | 22 | 22 | 22 | 22 |  | 22 | 22 | 22 | 22 |  | 3
2 RIGHT | 23 | 23 | 23 | 23 |  | 21 | 21 | 21 | 21 |  | 8
3 LEFT | 22 | 22 | 22 | 22 |  | 18 | 18 | 18 | 18 |  | 8
3 RIGHT | 23 | 23 | 23 | 23 |  | 18 | 18 | 18 | 18 |  | 2
4 LEFT | 22 | 22 | 22 | 22 |  | 8 | 8 | 8 | 8 |  | 2
4 RIGHT | 23 | 23 | 23 | 23 |  | 8 | 8 | 8 | 8 |  | 6
5 LEFT | 22 | 22 | 22 | 22 |  | 19 | 19 | 19 | 19 |  | 6
5 RIGHT | 23 | 23 | 23 | 23 |  | 19 | 19 | 19 | 19 |  | 7
6 LEFT | 22 | 22 | 22 | 22 |  | 21 | 21 | 21 | 21 |  | 7
6 RIGHT | 23 | 23 | 23 | 23 |  | 21 | 21 | 21 | 21 |  | 1
7 LEFT | 22 | 22 | 22 | 22 |  | 28 | 28 | 28 | 28 |  | 1
7 RIGHT | 23 | 23 | 23 | 23 |  | 28 | 28 | 28 | 28 |  | 5
8 LEFT | 22 | 22 | 22 | 22 |  | 19 | 19 | 19 | 19 |  | 5
8 RIGHT | 23 | 23 | 23 | 23 |  | 21 | 21 | 20 | 21 | C11.R8 (BLECF_0300) not installed | 4
Total | 360 | 360 | 360 | 360 | 0 | 309 | 309 | 308 | 309
in % |  | 100 | 100 | 100 | 0 |  | 100 | 99.6764 | 100
Missing |  | 0 | 0 | 0 | 360 |  | 0 | 1 | 0
Sum BLECF complete installation |  |  |  |  | 669
Installed BLECF |  |  |  |  | 669 | 100
Missing BLECF complete installation |  |  |  |  | 0 | 0

## Point 1
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L1 | BLECF_ | 285 | 0x | 011D | 6 | Arc | BYPLM.A33L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
2 | BJBHT.A32L1 | BLECF_ | 280 | 0x | 0118 | 6 | Arc | BYPLM.A32L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 00 | [9 empty] | No
3 | BJBHT.A31L1 | BLECF_ | 276 | 0x | 0114 | 6 | Arc | BYPLM.A31L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | yes
4 | BJBHT.A30L1 | BLECF_ | 201 | 0x | 00C9 | 6 | Arc | BYPLM.A30L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
5 | BJBHT.A29L1 | BLECF_ | 141 | 0x | 008D | 6 | Arc | BYPLM.A29L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
6 | BJBHT.A28L1 | BLECF_ | 249 | 0x | 00F9 | 6 | Arc | BYPLM.A28L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
7 | BJBHT.A27L1 | BLECF_ | 241 | 0x | 00F1 | 6 | Arc | BYPLM.A27L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
8 | BJBHT.A26L1 | BLECF_ | 282 | 0x | 011A | 6 | Arc | BYPLM.A26L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
9 | BJBHT.A25L1 | BLECF_ | 218 | 0x | 00DA | 6 | Arc | BYPLM.A25L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
10 | BJBHT.A24L1 | BLECF_ | 204 | 0x | 00CC | 6 | Arc | BYPLM.A24L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | yes
11 | BJBHT.A23L1 | BLECF_ | 212 | 0x | 00D4 | 6 | Arc | BYPLM.A23L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
12 | BJBHT.A22L1 | BLECF_ | 284 | 0x | 011C | 6 | Arc | BYPLM.A22L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
13 | BJBHT.A21L1 | BLECF_ | 22 | 0x | 0016 | 6 | Arc | BYPLM.A21L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
14 | BJBHT.A20L1 | BLECF_ | 754 | 0x | 02F2 | 6 | Arc | BYPLM.A20L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Wrong number -> cheking of card number in the tunnel
15 | BJBHT.A19L1 | BLECF_ | 35 | 0x | 0023 | 6 | Arc | BYPLM.A19L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | DAC reset needed (didn't reset with HV request)
16 | BJBHT.A18L1 | BLECF_ | 400 | 0x | 0190 | 6 | Arc | BYPLM.A18L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
17 | BJBHT.A17L1 | BLECF_ | 36 | 0x | 0024 | 6 | Arc | BYPLM.A17L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
18 | BJBHT.A16L1 | BLECF_ | 64 | 0x | 0040 | 6 | Arc | BYPLM.A16L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | O.K | O.K | O.K | ↑↑ | O.K | O.K | O.K | [10 empty] | No |  | Wrong number -> cheking of card number in the tunnel
19 | BJBHT.A15L1 | BLECF_ | 59 | 0x | 003B | 6 | Arc | BYPLM.A15L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
20 | BJBHT.A14L1 | BLECF_ | 58 | 0x | 003A | 6 | Arc | BYPLM.A14L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
21 | BJBHT.A13L1 | BLECF_ | 77 | 0x | 004D | 8 | Arc | BYPLM.A13L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
22 | BJBHT.A12L1 | BLECF_ | 79 | 0x | 004F | 6 | Arc | BYPLM.A12L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
23 | BJBHT.B11L1 | BLECF_ | 667 | 0x | 029B | 8 | RR13 | BY02 | 1 | CL | 2 | 9 | B | 13-14 | SR1 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Level blinking -> HV o.k?, ground o.k? !!! Isolation of a HV cable was toucinh the machine ground!
24 | BJBHT.A11L1 | BLECF_ | 594 | 0x | 0252 | 8 | RR13 | BY02 | 1 | CL | 2 | 8 | B | 11-12 | SR1 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
25 | BJBHT.B10L1 | BLECF_ | 409 | 0x | 0199 | 6 | RR13 | BY02 | 1 | CL | 2 | 7 | B | 9-10 | SR1 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
26 | BJBHT.A10L1 | BLECF_ | 550 | 0x | 0226 | 2 | RR13 | BY02 | 1 | CL | 2 | 6 | B | 7-8 | SR1 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
27 | BJBHT.A9L1 | BLECF_ | 518 | 0x | 0206 | 6 | RR13 | BY02 | 1 | CL | 2 | 5 | B | 5-6 | SR1 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Level blinking -> HV o.k?, ground o.k? !!! Isolation of a Signal cable was toucinh the machine ground!
28 | BJBHT.B8L1 | BLECF_ | 457 | 0x | 01C9 | 6 | RR13 | BY02 | 1 | CL | 2 | 4 | B | 3-4 | SR1 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
29 | BJBHT.A8L1 | BLECF_ | 588 | 0x | 024C | 6 | RR13 | BY02 | 1 | CL | 2 | 3 | B | 1-2 | SR1 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
30 | BJBHT.B7L1 | BLECF_ | 322 | 0x | 0142 | 7 | RR13 | BY02 | 1 | CL | 2 | 2 | A | 23-24 | SR1 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
31 | BJBHT.A7L1 | BLECF_ | 194 | 0x | 00C2 | 1 | RR13 | BY02 | 1 | CL | 2 | 1 | A | 21-22 | SR1 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
32 | BJBHT.A6L1 | BLECF_ | 88 | 0x | 0058 | 6 | RR13 | BY02 | 1 | CL | 1 | 10 | A | 19-20 | SR1 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | Wrong number -> cheking of card number in the tunnel
33 | BJBHT.B5L1 | BLECF_ | 530 | 0x | 0212 | 7 | RR13 | BY02 | 1 | CL | 1 | 9 | A | 17-18 | SR1 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
34 | BJBHT.A5L1 | BLECF_ | 597 | 0x | 0255 | 1 | RR13 | BY02 | 1 | CL | 1 | 8 | A | 15-16 | SR1 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
35 | BJBHT.C4L1 | BLECF_ | 398 | 0x | 018E | 6 | RR13 | BY02 | 1 | CL | 1 | 7 | A | 13-14 | SR1 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
36 | BJBHT.B4L1 | BLECF_ | 374 | 0x | 0176 | 4 | RR13 | BY02 | 1 | CL | 1 | 6 | A | 11-12 | SR1 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
37 | BJBHT.A4L1 | BLECF_ | 401 | 0x | 0191 | 4 | RR13 | BY02 | 1 | CL | 1 | 5 | A | 9-10 | SR1 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Wrong number -> cheking of card number in the tunnel
38 | BJBHT.A3L1 | BLECF_ | 593 | 0x | 0251 | 5 | RR13 | BY02 | 1 | CL | 1 | 4 | A | 7-8 | SR1 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
39 | BJBHT.A2L1 | BLECF_ | 592 | 0x | 0250 | 8 | RR13 | BY02 | 1 | CL | 1 | 3 | A | 5-6 | SR1 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | Level blinking -> HV o.k?, ground o.k? !!! New cable NES-18 wasn't connected! Now it is o.k !
40 | BJBHT.B1L1 | BLECF_ | 635 | 0x | 027B | 5 | RR13 | BY02 | 1 | CL | 1 | 2 | A | 3-4 | SR1 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | O.K | O.K | ↑↑↑ | ↑↑↑ | O.K | O.K | O.K | [9 empty] | NES18 | No
41 | BJBHT.A1L1 | BLECF_ | 639 | 0x | 027F | 1 | RR13 | BY02 | 1 | CL | 1 | 1 | A | 1-2 | SR1 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Ch4 and Ch5 (BLMQI.01L1.B1I30 & BLMQI.01L1.B2E10) are not modulating!!!!
42 | BJBHT.A1R1 | BLECF_ | 377 | 0x | 0179 | 1 | RR17 | BY02 | 1 | CR | 1 | 1 | A | 1-2 | SR1 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | All Levels are blinking, a lot of noise on all channels
43 | BJBHT.B1R1 | BLECF_ | 331 | 0x | 014B | 5 | RR17 | BY02 | 1 | CR | 1 | 2 | A | 3-4 | SR1 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
44 | BJBHT.A2R1 | BLECF_ | 501 | 0x | 01F5 | 8 | RR17 | BY02 | 1 | CR | 1 | 3 | A | 5-6 | SR1 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
45 | BJBHT.A3R1 | BLECF_ | 636 | 0x | 027C | 5 | RR17 | BY02 | 1 | CR | 1 | 4 | A | 7-8 | SR1 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
46 | BJBHT.A4R1 | BLECF_ | 676 | 0x | 02A4 | 4 | RR17 | BY02 | 1 | CR | 1 | 5 | A | 9-10 | SR1 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
47 | BJBHT.B4R1 | BLECF_ | 653 | 0x | 028D | 4 | RR17 | BY02 | 1 | CR | 1 | 6 | A | 11-12 | SR1 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
48 | BJBHT.C4R1 | BLECF_ | 376 | 0x | 0178 | 6 | RR17 | BY02 | 1 | CR | 1 | 7 | A | 13-14 | SR1 | BY05 | C | 11 | 1-2 | BY02 | R | 3 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
49 | BJBHT.A5R1 | BLECF_ | 437 | 0x | 01B5 | 1 | RR17 | BY02 | 1 | CR | 1 | 8 | A | 15-16 | SR1 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
50 | BJBHT.B5R1 | BLECF_ | 442 | 0x | 01BA | 7 | RR17 | BY02 | 1 | CR | 1 | 9 | A | 17-18 | SR1 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
51 | BJBHT.A6R1 | BLECF_ | 355 | 0x | 0163 | 6 | RR17 | BY02 | 1 | CR | 1 | 10 | A | 19-20 | SR1 | BY05 | C | 11 | 7-8 | BY02 | R | 3 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
52 | BJBHT.A7R1 | BLECF_ | 323 | 0x | 0143 | 1 | RR17 | BY02 | 1 | CR | 2 | 1 | A | 21-22 | SR1 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
53 | BJBHT.B7R1 | BLECF_ | 512 | 0x | 0200 | 7 | RR17 | BY02 | 1 | CR | 2 | 2 | A | 23-24 | SR1 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
54 | BJBHT.A8R1 | BLECF_ | 325 | 0x | 0145 | 6 | RR17 | BY02 | 1 | CR | 2 | 3 | B | 1-2 | SR1 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
55 | BJBHT.B8R1 | BLECF_ | 647 | 0x | 0287 | 6 | RR17 | BY02 | 1 | CR | 2 | 4 | B | 3-4 | SR1 | BY05 | C | 10 | 3-4 | BY02 | R | 3 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Channel 2 is blocked to 0 -> CFC ERROR is not on!!!, (DAC above 155) All the other channels are noisy!
56 | BJBHT.A9R1 | BLECF_ | 85 | 0x | 0055 | 6 | RR17 | BY02 | 1 | CR | 2 | 5 | B | 5-6 | SR1 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | Channel 3,4,5 is blocked to 0 -> CFC ERROR is not on!!! All the other channels are noisy!
57 | BJBHT.A10R1 | BLECF_ | 338 | 0x | 0152 | 2 | RR17 | BY02 | 1 | CR | 2 | 6 | B | 7-8 | SR1 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
58 | BJBHT.B10R1 | BLECF_ | 567 | 0x | 0237 | 6 | RR17 | BY02 | 1 | CR | 2 | 7 | B | 9-10 | SR1 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Levels are slowly blinking , HV missing and a lot of noise on all channels; Channel 2,4,5,8 is blocked to 0
59 | BJBHT.A11R1 | BLECF_ | 349 | 0x | 015D | 8 | RR17 | BY02 | 1 | CR | 2 | 8 | B | 11-12 | SR1 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | HV missing, all Levels are blinking
60 | BJBHT.B11R1 | BLECF_ | 327 | 0x | 0147 | 8 | RR17 | BY03 | 1 | CR | 2 | 9 | B | 13-14 | SR2 | BY06 | C | 9 | 1-2 | BY02 | R | 3 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
61 | BJBHT.A12R1 | BLECF_ | 539 | 0x | 021B | 6 | Arc | BYPLM.A12R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | O.K | O.K | OK | OK | O.K | O.K | O.K | [10 empty] | No
62 | BJBHT.A13R1 | BLECF_ | 361 | 0x | 0169 | 8 | Arc | BYPLM.A13R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | O.K | O.K | OK | OK | O.K | O.K | O.K | [10 empty] | No
63 | BJBHT.A14R1 | BLECF_ | 570 | 0x | 023A | 6 | Arc | BYPLM.A14R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | O.K | O.K | OK | OK | O.K | O.K | O.K | [10 empty] | No
64 | BJBHT.A15R1 | BLECF_ | 306 | 0x | 0132 | 6 | Arc | BYPLM.A15R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
65 | BJBHT.A16R1 | BLECF_ | 311 | 0x | 0137 | 6 | Arc | BYPLM.A16R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | No Power on the Main
66 | BJBHT.A17R1 | BLECF_ | 303 | 0x | 012F | 6 | Arc | BYPLM.A17R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | O.K | O.K | ↑↑↑ | ↑↑↑ | O.K | O.K | O.K | [10 empty] | No
67 | BJBHT.A18R1 | BLECF_ | 515 | 0x | 0203 | 6 | Arc | BYPLM.A18R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
68 | BJBHT.A19R1 | BLECF_ | 481 | 0x | 01E1 | 6 | Arc | BYPLM.A19R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
69 | BJBHT.A20R1 | BLECF_ | 81 | 0x | 0051 | 6 | Arc | BYPLM.A20R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
70 | BJBHT.A21R1 | BLECF_ | 522 | 0x | 020A | 6 | Arc | BYPLM.A21R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
71 | BJBHT.A22R1 | BLECF_ | 373 | 0x | 0175 | 6 | Arc | BYPLM.A22R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
72 | BJBHT.A23R1 | BLECF_ | 423 | 0x | 01A7 | 6 | Arc | BYPLM.A23R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
73 | BJBHT.A24R1 | BLECF_ | 352 | 0x | 0160 | 6 | Arc | BYPLM.A24R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
74 | BJBHT.A25R1 | BLECF_ | 500 | 0x | 01F4 | 6 | Arc | BYPLM.A25R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
75 | BJBHT.A26R1 | BLECF_ | 479 | 0x | 01DF | 6 | Arc | BYPLM.A26R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
76 | BJBHT.A27R1 | BLECF_ | 478 | 0x | 01DE | 6 | Arc | BYPLM.A27R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | O.K | O.K | OK | OK | O.K | O.K | O.K | [10 empty] | No |  | all connected channels go to 0! HV needed to be checked
77 | BJBHT.A28R1 | BLECF_ | 317 | 0x | 013D | 6 | Arc | BYPLM.A28R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
78 | BJBHT.A29R1 | BLECF_ | 310 | 0x | 0136 | 6 | Arc | BYPLM.A29R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Cable disconnected from crate to workaround Problem with the WorldFip reset
79 | BJBHT.A30R1 | BLECF_ | 329 | 0x | 0149 | 6 | Arc | BYPLM.A30R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | O.K | O.K | ↑↑ | O.K | O.K | O.K | O.K | [10 empty] | No
80 | BJBHT.A31R1 | BLECF_ | 495 | 0x | 01EF | 6 | Arc | BYPLM.A31R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
81 | BJBHT.A32R1 | BLECF_ | 384 | 0x | 0180 | 6 | Arc | BYPLM.A32R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
82 | BJBHT.A33R1 | BLECF_ | 417 | 0x | 01A1 | 6 | Arc | BYPLM.A33R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
83 | BJBHT.A34R1 | BLECF_ | 321 | 0x | 0141 | 6 | Arc | BYPLM.A34R1 | 1 | R |  |  |  | 7-9 | SR1 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
|  |  | 81 | 0x | 0051
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Card number wrong while checking on the 26.11
Note on T79: effinger: 13.08.08 need to check if the fibre are changed to 5-6 --> checked 18.8.08 connected to 7-8
Note on T14: effinger: 13.08.08 need to check if the fibre are changed to 5-6 --> 18.8.08 was connected to 5&6 reconnected to 7&8
Note on D15: effinger: Card replaced because channels 3 -6 blocked to negative: BLECF_0283 replaced by BLECF_0212
Note on F84: Ewald Effinger: 0x013D = 317 BLECF_0317 is installed
Note on F76: Ewald Effinger: 0x0051 = 51 BLECF_0051 is installed
Note on F57: Ewald Effinger: 0x0163 to be checked with list of Ion
Note on AC83: Ewald Effinger: LF at 17:08 at 27/08/14 LF at 9:10 28.8 still the same
Note on AC68: Ewald Effinger: LF at 17:08 at 27/08/14 LF at 9:10 28.8 still the same
Note on AC69: Ewald Effinger: LF at 17:08 at 27/08/14 LF at 9:10 28.8 still the same
Note on AC70: Ewald Effinger: LF at 17:08 at 27/08/14 LF at 9:10 28.8 still the same
Note on F42: Ewald Effinger: 0x0191 to be check in Ions list
Note on AS83: Ewald Effinger: After disconneting the HV on the first BLM noise dissapeared. Disconncted HV number 4 noise dispeared
Note on AT83: Ewald Effinger: Exchange of two HV connectors from the russian team, noise disappeared!
Note on AT64: Ewald Effinger: Test with Andrey: All signal cables discconected on the signal box and there was still noise! HV disconnected on the return and all channels have been negative and didn't recuperate. Next test disconnetion of NG 18 cable on the electronic side.
Note on AQ50: Ewald Effinger: Installed XmasStop 2011/12
Note on AQ51: Ewald Effinger: Installed XmasStop 2011/12
Note on AC86: Ewald Effinger: After reset (15/12/14) it looked o.k!

## Point 2
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L2 | BLECF_ | 399 | 0x | 018F | 6 | Arc | BYPLM.A33L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
2 | BJBHT.A32L2 | BLECF_ | 348 | 0x | 015C | 6 | Arc | BYPLM.A32L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
3 | BJBHT.A31L2 | BLECF_ | 497 | 0x | 01F1 | 6 | Arc | BYPLM.A31L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | O.K | O.K | ↑↑↑ | ↑↑↑ | O.K | O.K | O.K | [10 empty] | No
4 | BJBHT.A30L2 | BLECF_ | 432 | 0x | 01B0 | 6 | Arc | BYPLM.A30L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
5 | BJBHT.A29L2 | BLECF_ | 520 | 0x | 0208 | 6 | Arc | BYPLM.A29L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | O.K | O.K | ↑↑↑ | ↑↑↑ | O.K | O.K | O.K | [10 empty] | No
6 | BJBHT.A28L2 | BLECF_ | 439 | 0x | 01B7 | 6 | Arc | BYPLM.A28L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | A lot of noise an the 6 connected channels!
7 | BJBHT.A27L2 | BLECF_ | 342 | 0x | 0156 | 6 | Arc | BYPLM.A27L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
8 | BJBHT.A26L2 | BLECF_ | 460 | 0x | 01CC | 6 | Arc | BYPLM.A26L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
9 | BJBHT.A25L2 | BLECF_ | 388 | 0x | 0184 | 6 | Arc | BYPLM.A25L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
10 | BJBHT.A24L2 | BLECF_ | 440 | 0x | 01B8 | 6 | Arc | BYPLM.A24L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | RSM
11 | BJBHT.A23L2 | BLECF_ | 468 | 0x | 01D4 | 6 | Arc | BYPLM.A23L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | RSM
12 | BJBHT.A22L2 | BLECF_ | 456 | 0x | 01C8 | 6 | Arc | BYPLM.A22L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
13 | BJBHT.A21L2 | BLECF_ | 111 | 0x | 006F | 6 | Arc | BYPLM.A21L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
14 | BJBHT.A20L2 | BLECF_ | 11 | 0x | 000B | 6 | Arc | BYPLM.A20L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
15 | BJBHT.A19L2 | BLECF_ | 715 | 0x | 02CB | 6 | Arc | BYPLM.A19L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | RSM
16 | BJBHT.A18L2 | BLECF_ | 366 | 0x | 016E | 6 | Arc | BYPLM.A18L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
17 | BJBHT.A17L2 | BLECF_ | 427 | 0x | 01AB | 6 | Arc | BYPLM.A17L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
18 | BJBHT.A16L2 | BLECF_ | 385 | 0x | 0181 | 6 | Arc | BYPLM.A16L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
19 | BJBHT.A15L2 | BLECF_ | 444 | 0x | 01BC | 6 | Arc | BYPLM.A15L2 | 2 | L |  |  |  | 5-6 | SR2 | BY05 | F | 4 | 5-6 | BY02 | L | 1 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | RSM
20 | BJBHT.A14L2 | BLECF_ | 380 | 0x | 017C | 6 | Arc | BYPLM.A14L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
21 | BJBHT.A13L2 | BLECF_ | 347 | 0x | 015B | 6 | Arc | BYPLM.A13L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
22 | BJBHT.A12L2 | BLECF_ | 200 | 0x | 00C8 | 8 | Arc | BYPLM.A12L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
23 | BJBHT.C11L2 | BLECF_ | 583 | 0x | 0247 | 6 | UA23 | BY02 | 2 | CL | 3 | 2 | C | 19-20 | SR2 | BY05 | C | 4 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
24 | BJBHT.B11L2 | BLECF_ | 343 | 0x | 0157 | 6 | UA23 | BY02 | 2 | CL | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Channels are more noise than all the other arc cards
25 | BJBHT.A11L2 | BLECF_ | 378 | 0x | 017A | 6 | UA23 | BY02 | 2 | CL | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
26 | BJBHT.B10L2 | BLECF_ | 98 | 0x | 0062 | 6 | UA23 | BY02 | 2 | CL | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
27 | BJBHT.A10L2 | BLECF_ | 363 | 0x | 016B | 7 | UA23 | BY02 | 2 | CL | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
28 | BJBHT.A9L2 | BLECF_ | 517 | 0x | 0205 | 6 | UA23 | BY02 | 2 | CL | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
29 | BJBHT.C8L2 | BLECF_ | 357 | 0x | 0165 | 6 | UA23 | BY02 | 2 | CL | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
30 | BJBHT.B8L2 | BLECF_ | 333 | 0x | 014D | 6 | UA23 | BY02 | 2 | CL | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
31 | BJBHT.A8L2 | BLECF_ | 491 | 0x | 01EB | 6 | UA23 | BY02 | 2 | CL | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
32 | BJBHT.A7L2 | BLECF_ | 362 | 0x | 016A | 6 | UA23 | BY02 | 2 | CL | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
33 | BJBHT.C6L2 | BLECF_ | 189 | 0x | 00BD | 6 | UA23 | BY02 | 2 | CL | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
34 | BJBHT.B6L2 | BLECF_ | 344 | 0x | 0158 | 6 | UA23 | BY02 | 2 | CL | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
35 | BJBHT.A6L2 | BLECF_ | 474 | 0x | 01DA | 6 | UA23 | BY02 | 2 | CL | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
36 | BJBHT.A5L2 | BLECF_ | 328 | 0x | 0148 | 6 | UA23 | BY02 | 2 | CL | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Voltage missing, because of missing BLM (MKI support)
37 | BJBHT.D4L2 | BLECF_ | 371 | 0x | 0173 | 8 | UA23 | BY02 | 2 | CL | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Reset on the CFC needed
38 | BJBHT.C4L2 | BLECF_ | 181 | 0x | 00B5 | 1 | UA23 | BY02 | 2 | CL | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | 1 LIC missing ! Check if 2nd is a LIC
39 | BJBHT.B4L2 | BLECF_ | 320 | 0x | 0140 | 4 | UA23 | BY02 | 2 | CL | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | DAC reset needed! Modulation not seen on CH4
40 | BJBHT.A4L2 | BLECF_ | 516 | 0x | 0204 | 7 | UA23 | BY02 | 2 | CL | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | Two BLMs are missing!
41 | BJBHT.A3L2 | BLECF_ | 324 | 0x | 0144 | 4 | UA23 | BY02 | 2 | CL | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
42 | BJBHT.A2L2 | BLECF_ | 188 | 0x | 00BC | 8 | UA23 | BY02 | 2 | CL | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
43 | BJBHT.B1L2 | BLECF_ | 510 | 0x | 01FE | 5 | UA23 | BY02 | 2 | CL | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | cable exchanged
44 | BJBHT.A1L2 | BLECF_ | 476 | 0x | 01DC | 1 | UA23 | BY02 | 2 | CL | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
45 | BJBHT.A1R2 | BLECF_ | 706 | 0x | 02C2 | 1 | UA27 | BY03 | 2 | CR | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
46 | BJBHT.B1R2 | BLECF_ | 606 | 0x | 025E | 5 | UA27 | BY03 | 2 | CR | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
47 | BJBHT.A2R2 | BLECF_ | 675 | 0x | 02A3 | 8 | UA27 | BY03 | 2 | CR | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | (Fibers wrongly plugged !)
48 | BJBHT.A3R2 | BLECF_ | 641 | 0x | 0281 | 4 | UA27 | BY03 | 2 | CR | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
49 | BJBHT.A4R2 | BLECF_ | 719 | 0x | 02CF | 5 | UA27 | BY03 | 2 | CR | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | BLMs where missing the russian team installed them 14.10
50 | BJBHT.B4R2 | BLECF_ | 738 | 0x | 02E2 | 2 | UA27 | BY03 | 2 | CR | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | BLMs where missing the russian team installed them 14.11
51 | BJBHT.C4R2 | BLECF_ | 652 | 0x | 028C | 1 | UA27 | BY03 | 2 | CR | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | BLMs where missing the russian team installed them 14.12
52 | BJBHT.D4R2 | BLECF_ | 665 | 0x | 0299 | 8 | UA27 | BY03 | 2 | CR | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
53 | BJBHT.A5R2 | BLECF_ | 740 | 0x | 02E4 | 6 | UA27 | BY03 | 2 | CR | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | NES18
54 | BJBHT.A6R2 | BLECF_ | 666 | 0x | 029A | 1 | UA27 | BY03 | 2 | CR | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
55 | BJBHT.B6R2 | BLECF_ | 722 | 0x | 02D2 | 7 | UA27 | BY03 | 2 | CR | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
56 | BJBHT.A7R2 | BLECF_ | 658 | 0x | 0292 | 6 | UA27 | BY03 | 2 | CR | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
57 | BJBHT.A8R2 | BLECF_ | 626 | 0x | 0272 | 6 | UA27 | BY03 | 2 | CR | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Order of fiber in the TC needed to be adjusted
58 | BJBHT.B8R2 | BLECF_ | 406 | 0x | 0196 | 6 | UA27 | BY03 | 2 | CR | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Fiber where inverted insight the TC. Ok now!
59 | BJBHT.C8R2 | BLECF_ | 4 | 0x | 0004 | 6 | UA27 | BY03 | 2 | CR | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Fibre between BY02 and BY05 brocken
60 | BJBHT.A9R2 | BLECF_ | 279 | 0x | 0117 | 6 | UA27 | BY03 | 2 | CR | 2 | 6 | C | 7 | SR2 | BY05 | C | 10 | 7 | BY02 | R | 3 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Fibre between tunnel and BY05 brocken
61 | BJBHT.A10R2 | BLECF_ | 97 | 0x | 0061 | 6 | UA27 | BY03 | 2 | CR | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
62 | BJBHT.B10R2 | BLECF_ | 721 | 0x | 02D1 | 6 | UA27 | BY03 | 2 | CR | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
63 | BJBHT.A11R2 | BLECF_ | 128 | 0x | 0080 | 6 | UA27 | BY03 | 2 | CR | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | Noise on all channels! Cable changed by NES18
64 | BJBHT.B11R2 | BLECF_ | 627 | 0x | 0273 | 6 | UA27 | BY03 | 2 | CR | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
65 | BJBHT.C11R2 | BLECF_ | 689 | 0x | 02B1 | 6 | UA27 | BY03 | 2 | CR | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
66 | BJBHT.A12R2 | BLECF_ | 681 | 0x | 02A9 | 8 | Arc | BYPLM.A12R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
67 | BJBHT.A13R2 | BLECF_ | 687 | 0x | 02AF | 6 | Arc | BYPLM.A13R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
68 | BJBHT.A14R2 | BLECF_ | 684 | 0x | 02AC | 6 | Arc | BYPLM.A14R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Some channels are not displayed in the Fix display
69 | BJBHT.A15R2 | BLECF_ | 484 | 0x | 01E4 | 6 | Arc | BYPLM.A15R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
70 | BJBHT.A16R2 | BLECF_ | 686 | 0x | 02AE | 6 | Arc | BYPLM.A16R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
71 | BJBHT.A17R2 | BLECF_ | 143 | 0x | 008F | 6 | Arc | BYPLM.A17R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
72 | BJBHT.A18R2 | BLECF_ | 463 | 0x | 01CF | 6 | Arc | BYPLM.A18R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Noise on the channel 6 , o.k. now and no problem while RSM | BMLEI.B2E10_Drift (modulation signal is very small)
73 | BJBHT.A19R2 | BLECF_ | 379 | 0x | 017B | 6 | Arc | BYPLM.A19R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
74 | BJBHT.A20R2 | BLECF_ | 382 | 0x | 017E | 6 | Arc | BYPLM.A20R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
75 | BJBHT.A21R2 | BLECF_ | 451 | 0x | 01C3 | 6 | Arc | BYPLM.A21R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
76 | BJBHT.A22R2 | BLECF_ | 404 | 0x | 0194 | 6 | Arc | BYPLM.A22R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
77 | BJBHT.A23R2 | BLECF_ | 55 | 0x | 0037 | 6 | Arc | BYPLM.A23R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
78 | BJBHT.A24R2 | BLECF_ | 418 | 0x | 01A2 | 6 | Arc | BYPLM.A24R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
79 | BJBHT.A25R2 | BLECF_ | 390 | 0x | 0186 | 6 | Arc | BYPLM.A25R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | DAC RESET NEEDED
80 | BJBHT.A26R2 | BLECF_ | 214 | 0x | 00D6 | 6 | Arc | BYPLM.A26R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
81 | BJBHT.A27R2 | BLECF_ | 694 | 0x | 02B6 | 6 | Arc | BYPLM.A27R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | HV missing. O.k. now!
82 | BJBHT.A28R2 | BLECF_ | 685 | 0x | 02AD | 6 | Arc | BYPLM.A28R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
83 | BJBHT.A29R2 | BLECF_ | 725 | 0x | 02D5 | 6 | Arc | BYPLM.A29R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
84 | BJBHT.A30R2 | BLECF_ | 195 | 0x | 00C3 | 6 | Arc | BYPLM.A30R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
85 | BJBHT.A31R2 | BLECF_ | 690 | 0x | 02B2 | 6 | Arc | BYPLM.A31R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
86 | BJBHT.A32R2 | BLECF_ | 82 | 0x | 0052 | 6 | Arc | BYPLM.A32R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
87 | BJBHT.A33R2 | BLECF_ | 751 | 0x | 02EF | 6 | Arc | BYPLM.A33R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | BLM exchanged (P1027) due to problem with HV connector
88 | BJBHT.A34R2 | BLECF_ | 724 | 0x | 02D4 | 6 | Arc | BYPLM.A34R2 | 2 | R |  |  |  | 7-9 | SR2 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | HV connector ?
89 |  | BLECF_ | 288 | 0x | 0120 | 7 | Arc | BYPLM.A34R2 | 3 | R |  |  |  | 7-10 | SR3 | BY06 | I | 13 | 7-9 | BY03 | R | 3 | 20 | 16 | 1-9 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | 1024 | o
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on F81: effinger: Checked-----> BLECF_0382 installed! Wrong number 0x017E (382) instead of 0x0188 (392) ; to be check in the tunnel; Card number checked: --> BLECF_0382 installed; sent number 382
Note on T23: effinger: Fibre 7 & 8 broken, spare fibre used! Still the case after LS1!!!!
Note on AC48: Ewald Effinger: CFC working checked with the fiber connection of card 1 Problem must me in the surface! Fibers where not plugged in properly
Note on AC46: Ewald Effinger: CFC working problem at the surface Fibers where not plugged in properly
Note on AC43: Ewald Effinger: CFC working problem at the surface Fibers where not plugged in properly
Note on AC39: Ewald Effinger: Fiber mostprobably internaly not changed to the new configuration!! ! Verify with Steve what he has changed! LF due to bad connection (10) in the patch panel (insight) Fiber 9-10 swaped on the patch panel Fiber mixed up in the TC.
Note on AC36: Ewald Effinger: Fiber mostprobably internaly not changed to the new configuration!! ! Verify with Steve what he has changed!
Note on E32: Ewald Effinger: 0x0168 installed! In the list from Ion the old number was 0050
Note on AS43: Ewald Effinger: HV ok. By using a HV jumper!
Note on AI29: Ewald Effinger: cable was missing, hidden in the cable tray.
Note on AT42: Ewald Effinger: Missing BLM LIC ,
Note on AT10: Ewald Effinger: Andrey made a reset and checked the HV. All loocked fine.
Note on AR42: Ewald Effinger: Reset done by Andrey
Note on AS41: Ewald Effinger: Andrey could not check since the cable is on the cealing! WE wait till all is installed.
Note on AC78: Ewald Effinger: 0 LF @16:30 1/10/2014 after a CPU reset still 0 LF after 23 hours!
Note on D84: Ewald Effinger: was BLECF_0085
Note on D93: Ewald Effinger: The old number was 360 ??
Note on AA53: Ewald Effinger: ↑↑↑
Note on AE53: Ewald Effinger: ↑↑↑
Note on AE63: Ewald Effinger: ↑↑↑
Note on AF63: Ewald Effinger: ↑↑↑
Note on AE64: Ewald Effinger: ↑↑↑
Note on AF64: Ewald Effinger: ↑↑↑
Note on AG63: Ewald Effinger: ↑↑↑
Note on AG64: Ewald Effinger: ↑↑↑
Note on AI69: Ewald Effinger: Noise on all channels! Cable changed by NES18
Note on AS55: Ewald Effinger: HV was missing
Note on AS56: Ewald Effinger: HV was missing
Note on AS57: Ewald Effinger: HV was missing
Note on AC65: Ewald Effinger: ↑↑↑
Note on AD66: Ewald Effinger: ↑↑↑
Note on T66: Ewald Effinger: Second fiber : C 9 - 7
Note on O66: Ewald Effinger: Second fibre: C 19
Note on AS69: Ewald Effinger: HV was missing, repaired!
Note on AQ47: Ewald Effinger: Installed XmasStop 2011/12
Note on AQ45: Ewald Effinger: Installed XmasStop 2011/12
Note on AQ53: Ewald Effinger: Installed XmasStop 2011/12
Note on AQ52: Ewald Effinger: Installed XmasStop 2011/12
Note on AT59: Ewald Effinger: Installed XmasStop 2011/12

## Point 3
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L3 | BLECF_ | 619 | 0x | 026B | 6 | Arc | BYPLM.A33L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
2 | BJBHT.A32L3 | BLECF_ | 601 | 0x | 0259 | 6 | Arc | BYPLM.A32L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
3 | BJBHT.A31L3 | BLECF_ | 727 | 0x | 02D7 | 6 | Arc | BYPLM.A31L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
4 | BJBHT.A30L3 | BLECF_ | 705 | 0x | 02C1 | 6 | Arc | BYPLM.A30L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Problem with RSM (Huge peek after conecting the source) All concetors verified!
5 | BJBHT.A29L3 | BLECF_ | 646 | 0x | 0286 | 6 | Arc | BYPLM.A29L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | HV was disconnected
6 | BJBHT.A28L3 | BLECF_ | 702 | 0x | 02BE | 6 | Arc | BYPLM.A28L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Signal and HV was disconnected
7 | BJBHT.A27L3 | BLECF_ | 607 | 0x | 025F | 6 | Arc | BYPLM.A27L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
8 | BJBHT.A26L3 | BLECF_ | 623 | 0x | 026F | 6 | Arc | BYPLM.A26L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
9 | BJBHT.A25L3 | BLECF_ | 717 | 0x | 02CD | 6 | Arc | BYPLM.A25L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
10 | BJBHT.A24L3 | BLECF_ | 319 | 0x | 013F | 6 | Arc | BYPLM.A24L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
11 | BJBHT.A23L3 | BLECF_ | 745 | 0x | 02E9 | 6 | Arc | BYPLM.A23L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Fiber exchanged by Steph 15/10 o.k. ! RSM
12 | BJBHT.A22L3 | BLECF_ | 645 | 0x | 0285 | 6 | Arc | BYPLM.A22L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Fiber exchanged by Steph 15/10 o.k. ! RSM
13 | BJBHT.A21L3 | BLECF_ | 741 | 0x | 02E5 | 6 | Arc | BYPLM.A21L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
14 | BJBHT.A20L3 | BLECF_ | 614 | 0x | 0266 | 6 | Arc | BYPLM.A20L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
15 | BJBHT.A19L3 | BLECF_ | 620 | 0x | 026C | 6 | Arc | BYPLM.A19L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
16 | BJBHT.A18L3 | BLECF_ | 720 | 0x | 02D0 | 6 | Arc | BYPLM.A18L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
17 | BJBHT.A17L3 | BLECF_ | 668 | 0x | 029C | 6 | Arc | BYPLM.A17L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
18 | BJBHT.A16L3 | BLECF_ | 604 | 0x | 025C | 6 | Arc | BYPLM.A16L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
19 | BJBHT.A15L3 | BLECF_ | 622 | 0x | 026E | 6 | Arc | BYPLM.A15L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
20 | BJBHT.A14L3 | BLECF_ | 305 | 0x | 0131 | 6 | Arc | BYPLM.A14L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
21 | BJBHT.A13L3 | BLECF_ | 600 | 0x | 0258 | 10 | Arc | BYPLM.A13L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | RSM strange signal seen! All connection have been checked ! To be remeasured
22 | BJBHT.A12L3 | BLECF_ | 703 | 0x | 02BF | 10 | Arc | BYPLM.A12L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | O.k. now, only the feedback was not connected correctly! RSM showed o.k!
23 | BJBHT.B11L3 | BLECF_ | 657 | 0x | 0291 | 7 | UJ33 | BY02 | 3 | L | 2 | 8 | C | 11-12 | SR3 | BY05 | C | 3 | 11-12 | BY02 | L | 1 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
24 | BJBHT.A11L3 | BLECF_ | 654 | 0x | 028E | 12 | UJ33 | BY02 | 3 | L | 2 | 7 | C | 9-10 | SR3 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Fiber exchanged by Steph 15/10 o.k. ! RSM ! Divider missing
25 | BJBHT.A10L3 | BLECF_ | 574 | 0x | 023E | 11 | UJ33 | BY02 | 3 | L | 2 | 6 | C | 7-8 | SR3 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
26 | BJBHT.B9L3 | BLECF_ | 674 | 0x | 02A2 | 11 | UJ33 | BY02 | 3 | L | 2 | 5 | C | 5-6 | SR3 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Divider missing (ion installed!)
27 | BJBHT.A9L3 | BLECF_ | 609 | 0x | 0261 | 10 | UJ33 | BY02 | 3 | L | 2 | 4 | C | 3-4 | SR3 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Fiber exchanged by Steph 15/10 o.k. ! RSM
28 | BJBHT.B8L3 | BLECF_ | 672 | 0x | 02A0 | 6 | UJ33 | BY02 | 3 | L | 2 | 3 | C | 1-2 | SR3 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
29 | BJBHT.A8L3 | BLECF_ | 103 | 0x | 0067 | 6 | UJ33 | BY02 | 3 | L | 2 | 2 | B | 23-24 | SR3 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
30 | BJBHT.B7L3 | BLECF_ | 576 | 0x | 0240 | 7 | UJ33 | BY02 | 3 | L | 2 | 1 | B | 21-22 | SR3 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | HV was o.k after the first check
31 | BJBHT.A7L3 | BLECF_ | 637 | 0x | 027D | 1 | UJ33 | BY02 | 3 | L | 1 | 10 | B | 19-20 | SR3 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
32 | BJBHT.C6L3 | BLECF_ | 211 | 0x | 00D3 | 6 | UJ33 | BY02 | 3 | L | 1 | 9 | B | 17-18 | SR3 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
33 | BJBHT.B6L3 | BLECF_ | 629 | 0x | 0275 | 4 | UJ33 | BY02 | 3 | L | 1 | 8 | B | 15-16 | SR3 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
34 | BJBHT.A6L3 | BLECF_ | 536 | 0x | 0218 | 4 | UJ33 | BY02 | 3 | L | 1 | 7 | B | 13-14 | SR3 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
35 | BJBHT.D5L3 | BLECF_ | 175 | 0x | 00AF | 2 | UJ33 | BY02 | 3 | L | 1 | 6 | B | 11-12 | SR3 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
36 | BJBHT.C5L3 | BLECF_ | 577 | 0x | 0241 | 8 | UJ33 | BY02 | 3 | L | 1 | 5 | B | 9-10 | SR3 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
37 | BJBHT.B5L3 | BLECF_ | 575 | 0x | 023F | 6 | UJ33 | BY02 | 3 | L | 1 | 4 | B | 7-8 | SR3 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
38 | BJBHT.A5L3 | BLECF_ | 602 | 0x | 025A | 6 | UJ33 | BY02 | 3 | L | 1 | 3 | B | 5-6 | SR3 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
39 | BJBHT.B4L3 | BLECF_ | 711 | 0x | 02C7 | 2 | UJ33 | BY02 | 3 | L | 1 | 2 | B | 3-4 | SR3 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
40 | BJBHT.A4L3 | BLECF_ | 729 | 0x | 02D9 | 8 | UJ33 | BY02 | 3 | L | 1 | 1 | B | 1-2 | SR3 | BY05 | C | 1 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
41 | BJBHT.A4R3 | BLECF_ | 543 | 0x | 021F | 8 | UJ33 | BY03 | 3 | R | 1 | 1 | C | 13-14 | SR3 | BY05 | C | 4 | 1-2 | BY02 | R | 3 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
42 | BJBHT.B4R3 | BLECF_ | 677 | 0x | 02A5 | 2 | UJ33 | BY03 | 3 | R | 1 | 2 | C | 15-16 | SR3 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
43 | BJBHT.A5R3 | BLECF_ | 182 | 0x | 00B6 | 6 | UJ33 | BY03 | 3 | R | 1 | 3 | C | 17-18 | SR3 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
44 | BJBHT.B5R3 | BLECF_ | 326 | 0x | 0146 | 6 | UJ33 | BY03 | 3 | R | 1 | 4 | C | 19-20 | SR3 | BY05 | C | 4 | 7-8 | BY02 | C | 2 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
45 | BJBHT.C5R3 | BLECF_ | 487 | 0x | 01E7 | 8 | UJ33 | BY03 | 3 | R | 1 | 5 | C | 21-22 | SR3 | BY05 | C | 4 | 9-10 | BY02 | R | 3 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
46 | BJBHT.D5R3 | BLECF_ | 393 | 0x | 0189 | 2 | UJ33 | BY03 | 3 | R | 1 | 6 | C | 23-24 | SR3 | BY05 | C | 4 | 11-12 | BY02 | C | 2 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
47 | BJBHT.A6R3 | BLECF_ | 541 | 0x | 021D | 4 | UJ33 | BY03 | 3 | R | 1 | 7 | D | 1-2 | SR3 | BY05 | C | 5 | 1-2 | BY02 | C | 2 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
48 | BJBHT.B6R3 | BLECF_ | 341 | 0x | 0155 | 4 | UJ33 | BY03 | 3 | R | 1 | 8 | D | 3-4 | SR3 | BY05 | C | 5 | 3-4 | BY02 | C | 2 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
49 | BJBHT.C6R3 | BLECF_ | 339 | 0x | 0153 | 6 | UJ33 | BY03 | 3 | R | 1 | 9 | D | 5-6 | SR3 | BY05 | C | 5 | 5-6 | BY02 | R | 3 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
50 | BJBHT.A7R3 | BLECF_ | 367 | 0x | 016F | 1 | UJ33 | BY03 | 3 | R | 1 | 10 | D | 7-8 | SR3 | BY05 | C | 5 | 7-8 | BY02 | C | 2 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Divider missing
51 | BJBHT.B7R3 | BLECF_ | 709 | 0x | 02C5 | 7 | UJ33 | BY03 | 3 | R | 2 | 1 | D | 9-10 | SR3 | BY05 | C | 5 | 9-10 | BY02 | R | 3 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | no modulation signal
52 | BJBHT.A8R3 | BLECF_ | 502 | 0x | 01F6 | 6 | UJ33 | BY03 | 3 | R | 2 | 2 | D | 11-12 | SR3 | BY05 | C | 5 | 11-12 | BY02 | C | 2 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
53 | BJBHT.B8R3 | BLECF_ | 183 | 0x | 00B7 | 6 | UJ33 | BY03 | 3 | R | 2 | 3 | D | 13-14 | SR3 | BY05 | C | 6 | 1-2 | BY02 | R | 3 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
54 | BJBHT.A9R3 | BLECF_ | 375 | 0x | 0177 | 10 | UJ33 | BY03 | 3 | R | 2 | 4 | D | 15-16 | SR3 | BY05 | C | 6 | 3-4 | BY02 | C | 2 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
55 | BJBHT.B9R3 | BLECF_ | 410 | 0x | 019A | 11 | UJ33 | BY03 | 3 | R | 2 | 5 | D | 17-18 | SR3 | BY05 | C | 6 | 5-6 | BY02 | R | 3 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | no
56 | BJBHT.A10R3 | BLECF_ | 186 | 0x | 00BA | 11 | UJ33 | BY03 | 3 | R | 2 | 6 | D | 19-20 | SR3 | BY05 | C | 6 | 7-8 | BY02 | R | 3 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
57 | BJBHT.A11R3 | BLECF_ | 562 | 0x | 0232 | 12 | UJ33 | BY03 | 3 | R | 2 | 7 | D | 21-22 | SR3 | BY05 | C | 6 | 9-10 | BY02 | C | 2 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
58 | BJBHT.B11R3 | BLECF_ | 459 | 0x | 01CB | 7 | UJ33 | BY03 | 3 | R | 2 | 8 | D | 23-24 | SR3 | BY05 | C | 6 | 11-12 | BY02 | R | 3 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | no |  | no modulation signal
59 | BJBHT.A12R3 | BLECF_ | 68 | 0x | 0044 | 10 | Arc | BYPLM.A12R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | card is regulary reset!!!
60 | BJBHT.A13R3 | BLECF_ | 527 | 0x | 020F | 10 | Arc | BYPLM.A13R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | was 611, removed to install at 12r6 !!! no modulation signal !!!
61 | BJBHT.A14R3 | BLECF_ | 504 | 0x | 01F8 | 6 | Arc | BYPLM.A14R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
62 | BJBHT.A15R3 | BLECF_ | 499 | 0x | 01F3 | 6 | Arc | BYPLM.A15R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
63 | BJBHT.A16R3 | BLECF_ | 514 | 0x | 0202 | 6 | Arc | BYPLM.A16R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
64 | BJBHT.A17R3 | BLECF_ | 493 | 0x | 01ED | 6 | Arc | BYPLM.A17R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
65 | BJBHT.A18R3 | BLECF_ | 496 | 0x | 01F0 | 6 | Arc | BYPLM.A18R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
66 | BJBHT.A19R3 | BLECF_ | 314 | 0x | 013A | 6 | Arc | BYPLM.A19R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | no modulation signal
67 | BJBHT.A20R3 | BLECF_ | 523 | 0x | 020B | 6 | Arc | BYPLM.A20R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
68 | BJBHT.A21R3 | BLECF_ | 185 | 0x | 00B9 | 6 | Arc | BYPLM.A21R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
69 | BJBHT.A22R3 | BLECF_ | 387 | 0x | 0183 | 6 | Arc | BYPLM.A22R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
70 | BJBHT.A23R3 | BLECF_ | 485 | 0x | 01E5 | 6 | Arc | BYPLM.A23R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
71 | BJBHT.A24R3 | BLECF_ | 309 | 0x | 0135 | 6 | Arc | BYPLM.A24R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
72 | BJBHT.A25R3 | BLECF_ | 184 | 0x | 00B8 | 6 | Arc | BYPLM.A25R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
73 | BJBHT.A26R3 | BLECF_ | 337 | 0x | 0151 | 6 | Arc | BYPLM.A26R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
74 | BJBHT.A27R3 | BLECF_ | 519 | 0x | 0207 | 6 | Arc | BYPLM.A27R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
75 | BJBHT.A28R3 | BLECF_ | 448 | 0x | 01C0 | 6 | Arc | BYPLM.A28R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | no modulation signal
76 | BJBHT.A29R3 | BLECF_ | 489 | 0x | 01E9 | 6 | Arc | BYPLM.A29R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
77 | BJBHT.A30R3 | BLECF_ | 492 | 0x | 01EC | 6 | Arc | BYPLM.A30R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
78 | BJBHT.A31R3 | BLECF_ | 488 | 0x | 01E8 | 6 | Arc | BYPLM.A31R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
79 | BJBHT.A32R3 | BLECF_ | 407 | 0x | 0197 | 6 | Arc | BYPLM.A32R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
80 | BJBHT.A33R3 | BLECF_ | 453 | 0x | 01C5 | 6 | Arc | BYPLM.A33R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
81 | BJBHT.A34R3 | BLECF_ | 109 | 0x | 006D | 6 | Arc | BYPLM.A34R3 | 3 | R |  |  |  | 6-7 | SR3 | BY05 | I | 12 | 6-7 | BY02 | R | 3 | 20 | 16 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on AD88: Ewald Effinger: was on 7-8
Note on AD85: Ewald Effinger: LF cleaned and reconnected than it was o.k.
Note on AE66: Ewald Effinger: Fibers mixed in the TC
Note on AL10: Ewald Effinger: Signal was disconnected
Note on D15: Ewald Effinger: was BLECF_0730
Note on D30: Ewald Effinger: was 0754
Note on D34: Ewald Effinger: was 574
Note on AT32: Ewald Effinger: Comment from Matti after RSM: • Signals from monitors connected to BJBHT.A11L3 and BJBHT.A9L3 are swapped Fiber must be exchanged!!!!
Note on AT29: Ewald Effinger: Comment from Matti after RSM: • Signals from monitors connected to BJBHT.A11L3 and BJBHT.A9L3 are swapped Fiber must be exchanged!!!!
Note on AT15: Ewald Effinger: Comment from Matti after RSM: • Signals from monitors connected to BJBHT.A23L3 and BJBHT.A22L3 are swapped Fiber should be inversed in the surface! (RSM confirms the swap)
Note on AT16: Ewald Effinger: Comment from Matti after RSM: • Signals from monitors connected to BJBHT.A23L3 and BJBHT.A22L3 are swapped Fiber should be inversed in the surface! (RSM confirms the swap)
Note on AS26: Ewald Effinger: HV distribution need to be checked!
Note on F66: Ewald Effinger: was 472
Note on AE60: Ewald Effinger: ↑↑↑
Note on AF60: Ewald Effinger: ↑↑↑
Note on AG60: Ewald Effinger: ↑↑↑
Note on AE63: Ewald Effinger: ↑↑↑
Note on AF63: Ewald Effinger: ↑↑↑
Note on AG63: Ewald Effinger: ↑↑↑
Note on AG47: Ewald Effinger: 1
Note on AQ64: Ewald Effinger: Installed in september 2011

## Point 4
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L4 | BLECF_ | 452 | 0x | 01C4 | 6 | Arc | BYPLM.A33L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8
2 | BJBHT.A32L4 | BLECF_ | 748 | 0x | 02EC | 6 | Arc | BYPLM.A32L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 |  |  |  | ↑↑↑
3 | BJBHT.A31L4 | BLECF_ | 616 | 0x | 0268 | 6 | Arc | BYPLM.A31L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8
4 | BJBHT.A30L4 | BLECF_ | 662 | 0x | 0296 | 6 | Arc | BYPLM.A30L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 |  |  | ↑↑↑
5 | BJBHT.A29L4 | BLECF_ | 38 | 0x | 0026 | 6 | Arc | BYPLM.A29L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 |  |  |  | ↑↑↑ | [15 empty] | GOH2 ready signal off!
6 | BJBHT.A28L4 | BLECF_ | 76 | 0x | 004C | 6 | Arc | BYPLM.A28L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16
7 | BJBHT.A27L4 | BLECF_ | 572 | 0x | 023C | 6 | Arc | BYPLM.A27L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8
8 | BJBHT.A26L4 | BLECF_ | 21 | 0x | 0015 | 6 | Arc | BYPLM.A26L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16
9 | BJBHT.A25L4 | BLECF_ | 613 | 0x | 0265 | 6 | Arc | BYPLM.A25L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8
10 | BJBHT.A24L4 | BLECF_ | 655 | 0x | 028F | 6 | Arc | BYPLM.A24L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16
11 | BJBHT.A23L4 | BLECF_ | 631 | 0x | 0277 | 6 | Arc | BYPLM.A23L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8
12 | BJBHT.A22L4 | BLECF_ | 743 | 0x | 02E7 | 6 | Arc | BYPLM.A22L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16
13 | BJBHT.A21L4 | BLECF_ | 664 | 0x | 0298 | 6 | Arc | BYPLM.A21L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8
14 | BJBHT.A20L4 | BLECF_ | 471 | 0x | 01D7 | 6 | Arc | BYPLM.A20L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16
15 | BJBHT.A19L4 | BLECF_ | 552 | 0x | 0228 | 6 | Arc | BYPLM.A19L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 |  |  | ↑↑↑ | ↑↑↑
16 | BJBHT.A18L4 | BLECF_ | 395 | 0x | 018B | 6 | Arc | BYPLM.A18L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16
17 | BJBHT.A17L4 | BLECF_ | 615 | 0x | 0267 | 6 | Arc | BYPLM.A17L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8
18 | BJBHT.A16L4 | BLECF_ | 591 | 0x | 024F | 6 | Arc | BYPLM.A16L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16
19 | BJBHT.A15L4 | BLECF_ | 642 | 0x | 0282 | 6 | Arc | BYPLM.A15L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8
20 | BJBHT.A14L4 | BLECF_ | 670 | 0x | 029E | 6 | Arc | BYPLM.A14L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16
21 | BJBHT.A13L4 | BLECF_ | 599 | 0x | 0257 | 6 | Arc | BYPLM.A13L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8
22 | BJBHT.A12L4 | BLECF_ | 659 | 0x | 0293 | 6 | Arc | BYPLM.A12L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16
23 | BJBHT.B5L4 | BLECF_ | 251 | 0x | 00FB | 0 | UA43 | BY12 | 4 | CL | 1 | 9 | A | 17-18 | SX4 | BY04 | C | 2 | 5-6 | BY02 | C | 2 | 4 | 1 | 9-16
24 | BJBHT.A11L4 | BLECF_ | 643 | 0x | 0283 | 6 | UA43 | BY12 | 4 | CL | 1 | 8 | A | 15-16 | SX4 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 7 | 4 | 1-8
25 | BJBHT.A10L4 | BLECF_ | 578 | 0x | 0242 | 6 | UA43 | BY12 | 4 | CL | 1 | 7 | A | 13-14 | SX4 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 6 | 3 | 9-16 | [16 empty] | NES18 |  |  | NES 18 connected ? Needs to be checked!
26 | BJBHT.A9L4 | BLECF_ | 408 | 0x | 0198 | 6 | UA43 | BY12 | 4 | CL | 1 | 6 | A | 11-12 | SX4 | BY05 | C | 1 | 11-12 | BY02 | L | 1 | 6 | 3 | 1-8
27 | BJBHT.A8L4 | BLECF_ | 651 | 0x | 028B | 6 | UA43 | BY12 | 4 | CL | 1 | 5 | A | 9-10 | SX4 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | [16 empty] | NES18 |  |  | NES 18 connected ? Needs to be checked!
28 | BJBHT.A7L4 | BLECF_ | 386 | 0x | 0182 | 6 | UA43 | BY12 | 4 | CL | 1 | 4 | A | 7-8 | SX4 | BY05 | C | 1 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8
29 | BJBHT.A6L4 | BLECF_ | 157 | 0x | 009D | 6 | UA43 | BY12 | 4 | CL | 1 | 3 | A | 5-6 | SX4 | BY05 | C | 1 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 |  |  |  | ↑↑↑
30 | BJBHT.C5L4 | BLECF_ | 744 | 0x | 02E8 | 8 | UA43 | BY12 | 4 | CL | 1 | 2 | A | 3-4 | SX4 | BY05 | C | 1 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8
31 | BJBHT.B5L4 | BLECF_ | 624 | 0x | 0270 | 2 | UA43 | BY12 | 4 | CL | 1 | 1 | A | 1-2 | SX4 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8
32 | BJBHT.B5R4 | BLECF_ | 8 | 0x | 0008 | 2 | UA47 | BY09 | 4 | CR | 1 | 1 | A | 1-2 | SX4 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8
33 | BJBHT.C5R4 | BLECF_ | 5 | 0x | 0005 | 8 | UA47 | BY09 | 4 | CR | 1 | 2 | A | 3-4 | SX4 | BY05 | C | 12 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8 | [16 empty] | NES18 |  |  | NES 18 connected ? Needs to be checked!
34 | BJBHT.A6R4 | BLECF_ | 316 | 0x | 013C | 6 | UA47 | BY09 | 4 | CR | 1 | 3 | A | 5-6 | SX4 | BY05 | C | 12 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16
35 | BJBHT.A7R4 | BLECF_ | 235 | 0x | 00EB | 6 | UA47 | BY09 | 4 | CR | 1 | 4 | A | 7-8 | SX4 | BY05 | C | 12 | 7-8 | BY02 | R | 3 | 5 | 2 | 1-8 | [16 empty] | NES18 |  |  | NES 18 connected ? Needs to be checked!
36 | BJBHT.A8R4 | BLECF_ | 584 | 0x | 0248 | 6 | UA47 | BY09 | 4 | CR | 1 | 5 | A | 9-10 | SX4 | BY05 | C | 12 | 9-10 | BY02 | R | 3 | 5 | 2 | 9-16
37 | BJBHT.A9R4 | BLECF_ | 509 | 0x | 01FD | 6 | UA47 | BY09 | 4 | CR | 1 | 6 | A | 11-12 | SX4 | BY05 | C | 12 | 11-12 | BY02 | R | 3 | 6 | 3 | 1-8
38 | BJBHT.A10R4 | BLECF_ | 7 | 0x | 0007 | 6 | UA47 | BY09 | 4 | CR | 1 | 7 | A | 13-14 | SX4 | BY05 | C | 11 | 1-2 | BY02 | R | 3 | 6 | 3 | 9-16 | [19 empty] | Card reset needed!
39 | BJBHT.A11R4 | BLECF_ | 10 | 0x | 000A | 6 | UA47 | BY09 | 4 | CR | 1 | 8 | A | 15-16 | SX4 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 7 | 4 | 1-8
40 | BJBHT.A5R4 | BLECF_ | 119 | 0x | 0077 | 0 | UA47 | BY09 | 4 | CR | 1 | 9 | A | 17-18 | SX4 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 13 | 9 | 9-16
41 | BJBHT.A12R4 | BLECF_ | 217 | 0x | 00D9 | 6 | Arc | BYPLM.A12R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 7 | 4 | 9-16 | OK | OK | OK | OK | OK | OK | OK
42 | BJBHT.A13R4 | BLECF_ | 531 | 0x | 0213 | 6 | Arc | BYPLM.A13R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | OK | OK | OK | OK | OK | OK | OK
43 | BJBHT.A14R4 | BLECF_ | 568 | 0x | 0238 | 6 | Arc | BYPLM.A14R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | OK | OK | OK | OK | OK | OK | OK
44 | BJBHT.A15R4 | BLECF_ | 57 | 0x | 0039 | 6 | Arc | BYPLM.A15R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | OK | OK | OK | OK | OK | OK | OK
45 | BJBHT.A16R4 | BLECF_ | 57 | 0x | 0039 | 6 | Arc | BYPLM.A16R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | OK | OK | OK | OK | OK | OK | OK
46 | BJBHT.A17R4 | BLECF_ | 403 | 0x | 0193 | 6 | Arc | BYPLM.A17R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | OK | ↑↑↑ | OK | OK | ↑↑↑ | OK | OK
47 | BJBHT.A18R4 | BLECF_ | 558 | 0x | 022E | 6 | Arc | BYPLM.A18R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | OK | OK | OK | OK | OK | OK | OK
48 | BJBHT.A19R4 | BLECF_ | 304 | 0x | 0130 | 6 | Arc | BYPLM.A19R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | OK | OK | OK | OK | OK | OK | OK
49 | BJBHT.A20R4 | BLECF_ | 585 | 0x | 0249 | 6 | Arc | BYPLM.A20R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | OK | OK | OK | OK | OK | OK | OK
50 | BJBHT.A21R4 | BLECF_ | 315 | 0x | 013B | 6 | Arc | BYPLM.A21R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | OK | OK | OK | OK | OK | OK | OK
51 | BJBHT.A22R4 | BLECF_ | 340 | 0x | 0154 | 6 | Arc | BYPLM.A22R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | OK | OK | OK | OK | OK | OK | OK
52 | BJBHT.A23R4 | BLECF_ | 537 | 0x | 0219 | 6 | Arc | BYPLM.A23R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | OK | OK | OK | OK | OK | OK | OK
53 | BJBHT.A24R4 | BLECF_ | 405 | 0x | 0195 | 6 | Arc | BYPLM.A24R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | OK | OK | OK | OK | OK | OK | OK
54 | BJBHT.A25R4 | BLECF_ | 483 | 0x | 01E3 | 6 | Arc | BYPLM.A25R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | OK | OK | OK | OK | OK | OK | OK
55 | BJBHT.A26R4 | BLECF_ | 438 | 0x | 01B6 | 6 | Arc | BYPLM.A26R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | OK | OK | OK | OK | OK | OK | OK
56 | BJBHT.A27R4 | BLECF_ | 547 | 0x | 0223 | 6 | Arc | BYPLM.A27R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | OK | OK | OK | OK | OK | OK | OK
57 | BJBHT.A28R4 | BLECF_ | 551 | 0x | 0227 | 6 | Arc | BYPLM.A28R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | OK | OK | OK | OK | OK | OK | OK
58 | BJBHT.A29R4 | BLECF_ | 83 | 0x | 0053 | 6 | Arc | BYPLM.A29R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | OK | OK | OK | OK | OK | OK | OK
59 | BJBHT.A30R4 | BLECF_ | 121 | 0x | 0079 | 6 | Arc | BYPLM.A30R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | OK | OK | OK | OK | OK | OK | OK
60 | BJBHT.A31R4 | BLECF_ | 40 | 0x | 0028 | 6 | Arc | BYPLM.A31R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | OK | OK | OK | OK | OK | OK | OK
61 | BJBHT.A32R4 | BLECF_ | 230 | 0x | 00E6 | 6 | Arc | BYPLM.A32R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | OK | OK | OK | OK | OK | OK | OK
62 | BJBHT.A33R4 | BLECF_ | 582 | 0x | 0246 | 6 | Arc | BYPLM.A33R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | OK | OK | OK | OK | OK | OK | OK
63 | BJBHT.A34R4 | BLECF_ | 389 | 0x | 0185 | 6 | Arc | BYPLM.A34R4 | 4 | R |  |  |  | 7-9 | SX4 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | OK | OK | OK | OK | OK | OK | OK
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on D12: Ewald Effinger: Number from the calibration list of Ion
Note on D41: Ewald Effinger: Number from the calibration list of Ion
Note on D68: Ewald Effinger: Number from the calibration list of Ion
Note on D66: Ewald Effinger: Number from the calibration list of Ion
Note on D55: Ewald Effinger: Number from the calibration list of Ion
Note on D52: Ewald Effinger: Number (BLECF_0134) from the calibration list of Ion. Installed BLECF_0057
Note on F52: Ewald Effinger: BLECF_0057 (x0039) installed
Note on AD46: Ewald Effinger: ↑↑↑ o.k 9.12.2014
Note on AS56: Ewald Effinger: HV problem

## Point 5
| Tunnel - position | [13 empty] | Surface position | [19 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | Comment | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | HV | 1024 | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L5 | BLECF_ | 335 | 0x | 014F | 6 | Arc | BYPLM.A33L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16
2 | BJBHT.A32L5 | BLECF_ | 330 | 0x | 014A | 6 | Arc | BYPLM.A32L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | Test in tunnel needed! | [19 empty] | CFC reset needed! Done!
3 | BJBHT.A31L5 | BLECF_ | 650 | 0x | 028A | 6 | Arc | BYPLM.A31L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | Test in tunnel needed!
4 | BJBHT.A30L5 | BLECF_ | 414 | 0x | 019E | 6 | Arc | BYPLM.A30L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8
5 | BJBHT.A29L5 | BLECF_ | 307 | 0x | 0133 | 6 | Arc | BYPLM.A29L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | [20 empty] | 220 V ? Was working !
6 | BJBHT.A28L5 | BLECF_ | 447 | 0x | 01BF | 6 | Arc | BYPLM.A28L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8
7 | BJBHT.A27L5 | BLECF_ | 226 | 0x | 00E2 | 6 | Arc | BYPLM.A27L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16
8 | BJBHT.A26L5 | BLECF_ | 680 | 0x | 02A8 | 6 | Arc | BYPLM.A26L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | Test in tunnel needed!
9 | BJBHT.A25L5 | BLECF_ | 424 | 0x | 01A8 | 6 | Arc | BYPLM.A25L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16
10 | BJBHT.A24L5 | BLECF_ | 419 | 0x | 01A3 | 6 | Arc | BYPLM.A24L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8
11 | BJBHT.A23L5 | BLECF_ | 34 | 0x | 0022 | 6 | Arc | BYPLM.A23L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | Test in tunnel needed!
12 | BJBHT.A22L5 | BLECF_ | 625 | 0x | 0271 | 6 | Arc | BYPLM.A22L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | Test in tunnel needed!
13 | BJBHT.A21L5 | BLECF_ | 506 | 0x | 01FA | 6 | Arc | BYPLM.A21L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | Test in tunnel needed! | [19 empty] | BLM behind the door!
14 | BJBHT.A20L5 | BLECF_ | 589 | 0x | 024D | 6 | Arc | BYPLM.A20L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8
15 | BJBHT.A19L5 | BLECF_ | 546 | 0x | 0222 | 6 | Arc | BYPLM.A19L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16
16 | BJBHT.A18L5 | BLECF_ | 473 | 0x | 01D9 | 6 | Arc | BYPLM.A18L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | Test in tunnel needed!
17 | BJBHT.A17L5 | BLECF_ | 475 | 0x | 01DB | 6 | Arc | BYPLM.A17L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16
18 | BJBHT.A16L5 | BLECF_ | 196 | 0x | 00C4 | 6 | Arc | BYPLM.A16L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8
19 | BJBHT.A15L5 | BLECF_ | 346 | 0x | 015A | 6 | Arc | BYPLM.A15L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16
20 | BJBHT.A14L5 | BLECF_ | 48 | 0x | 0030 | 6 | Arc | BYPLM.A14L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8
21 | BJBHT.A13L5 | BLECF_ | 640 | 0x | 0280 | 6 | Arc | BYPLM.A13L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | Test in tunnel needed!
22 | BJBHT.A12L5 | BLECF_ | 477 | 0x | 01DD | 6 | Arc | BYPLM.A12L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8
23 | BJBHT.B11L5 | BLECF_ | 753 | 0x | 02F1 | 6 | RR53 | BY02 | 5 | CL | 2 | 9 | B | 13-14 | SR5 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16
24 | BJBHT.A11L5 | BLECF_ | 581 | 0x | 0245 | 7 | RR53 | BY02 | 5 | CL | 2 | 8 | B | 11-12 | SR5 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8
25 | BJBHT.B10L5 | BLECF_ | 14 | 0x | 000E | 6 | RR53 | BY02 | 5 | CL | 2 | 7 | B | 9-10 | SR5 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 7 | 4 | 1-8
26 | BJBHT.A10L5 | BLECF_ | 124 | 0x | 007C | 2 | RR53 | BY02 | 5 | CL | 2 | 6 | B | 7-8 | SR5 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16
27 | BJBHT.A9L5 | BLECF_ | 757 | 0x | 02F5 | 6 | RR53 | BY02 | 5 | CL | 2 | 5 | B | 5-6 | SR5 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16
28 | BJBHT.B8L5 | BLECF_ | 190 | 0x | 00BE | 6 | RR53 | BY02 | 5 | CL | 2 | 4 | B | 3-4 | SR5 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8
29 | BJBHT.A8L5 | BLECF_ | 9 | 0x | 0009 | 6 | RR53 | BY02 | 5 | CL | 2 | 3 | B | 1-2 | SR5 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | [17 empty] | NES18
30 | BJBHT.A7L5 | BLECF_ | 525 | 0x | 020D | 6 | RR53 | BY02 | 5 | CL | 2 | 2 | A | 23-24 | SR5 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 9-16
31 | BJBHT.B6L5 | BLECF_ | 632 | 0x | 0278 | 7 | RR53 | BY02 | 5 | CL | 2 | 1 | A | 21-22 | SR5 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 1-8 | [17 empty] | NES18
32 | BJBHT.A6L5 | BLECF_ | 535 | 0x | 0217 | 1 | RR53 | BY02 | 5 | CL | 1 | 10 | A | 19-20 | SR5 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16 | Test in tunnel needed!
33 | BJBHT.B5L5 | BLECF_ | 313 | 0x | 0139 | 6 | RR53 | BY02 | 5 | CL | 1 | 9 | A | 17-18 | SR5 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | [17 empty] | NES18
34 | BJBHT.A5L5 | BLECF_ | 521 | 0x | 0209 | 2 | RR53 | BY02 | 5 | CL | 1 | 8 | A | 15-16 | SR5 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8
35 | BJBHT.C4L5 | BLECF_ | 18 | 0x | 0012 | 6 | RR53 | BY02 | 5 | CL | 1 | 7 | A | 13-14 | SR5 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | [17 empty] | NES18
36 | BJBHT.B4L5 | BLECF_ | 370 | 0x | 0172 | 7 | RR53 | BY02 | 5 | CL | 1 | 6 | A | 11-12 | SR5 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16
37 | BJBHT.A4L5 | BLECF_ | 3 | 0x | 0003 | 2 | RR53 | BY02 | 5 | CL | 1 | 5 | A | 9-10 | SR5 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | [17 empty] | NES18
38 | BJBHT.A3L5 | BLECF_ | 72 | 0x | 0048 | 5 | RR53 | BY02 | 5 | CL | 1 | 4 | A | 7-8 | SR5 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | [17 empty] | NES18 |  |  | NES18
39 | BJBHT.A2L5 | BLECF_ | 42 | 0x | 002A | 8 | RR53 | BY02 | 5 | CL | 1 | 3 | A | 5-6 | SR5 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | [17 empty] | NES18
40 | BJBHT.B1L5 | BLECF_ | 586 | 0x | 024A | 4 | RR53 | BY02 | 5 | CL | 1 | 2 | A | 3-4 | SR5 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16
41 | BJBHT.A1L5 | BLECF_ | 579 | 0x | 0243 | 2 | RR53 | BY02 | 5 | CL | 1 | 1 | A | 1-2 | SR5 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | [20 empty] | BLMES.01L5.B2E10_BPMSW.1L5 is not modulating
42 | BJBHT.A1R5 | BLECF_ | 193 | 0x | 00C1 | 2 | RR57 | BY02 | 5 | CR | 1 | 1 | A | 1-2 | SR5 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8
43 | BJBHT.B1R5 | BLECF_ | 749 | 0x | 02ED | 6 | RR57 | BY02 | 5 | CR | 1 | 2 | A | 3-4 | SR5 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | [17 empty] | NES18 |  |  | NES18
44 | BJBHT.A2R5 | BLECF_ | 671 | 0x | 029F | 6 | RR57 | BY02 | 5 | CR | 1 | 3 | A | 5-6 | SR5 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8
45 | BJBHT.A3R5 | BLECF_ | 750 | 0x | 02EE | 6 | RR57 | BY02 | 5 | CR | 1 | 4 | A | 7-8 | SR5 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16
46 | BJBHT.A4R5 | BLECF_ | 726 | 0x | 02D6 | 2 | RR57 | BY02 | 5 | CR | 1 | 5 | A | 9-10 | SR5 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8
47 | BJBHT.B4R5 | BLECF_ | 663 | 0x | 0297 | 8 | RR57 | BY02 | 5 | CR | 1 | 6 | A | 11-12 | SR5 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16
48 | BJBHT.C4R5 | BLECF_ | 710 | 0x | 02C6 | 6 | RR57 | BY02 | 5 | CR | 1 | 7 | A | 13-14 | SR5 | BY05 | C | 11 | 1-2 | BY02 | R | 3 | 4 | 1 | 1-8
49 | BJBHT.A5R5 | BLECF_ | 598 | 0x | 0256 | 2 | RR57 | BY02 | 5 | CR | 1 | 8 | A | 15-16 | SR5 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8
50 | BJBHT.B5R5 | BLECF_ | 697 | 0x | 02B9 | 6 | RR57 | BY02 | 5 | CR | 1 | 9 | A | 17-18 | SR5 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | [20 empty] | Channel 7 (constant 10 blmbits @40us)
51 | BJBHT.A6R5 | BLECF_ | 648 | 0x | 0288 | 2 | RR57 | BY02 | 5 | CR | 1 | 10 | A | 19-20 | SR5 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16
52 | BJBHT.B6R5 | BLECF_ | 656 | 0x | 0290 | 6 | RR57 | BY02 | 5 | CR | 2 | 1 | A | 21-22 | SR5 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 5 | 2 | 1-8 | [17 empty] | NES18 |  |  | NES18
53 | BJBHT.A7R5 | BLECF_ | 383 | 0x | 017F | 6 | RR57 | BY02 | 5 | CR | 2 | 2 | A | 23-24 | SR5 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16
54 | BJBHT.A8R5 | BLECF_ | 596 | 0x | 0254 | 6 | RR57 | BY02 | 5 | CR | 2 | 3 | B | 1-2 | SR5 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8
55 | BJBHT.B8R5 | BLECF_ | 660 | 0x | 0294 | 6 | RR57 | BY02 | 5 | CR | 2 | 4 | B | 3-4 | SR5 | BY05 | C | 10 | 3-4 | BY02 | R | 3 | 6 | 3 | 1-8
56 | BJBHT.A9R5 | BLECF_ | 149 | 0x | 0095 | 6 | RR57 | BY02 | 5 | CR | 2 | 5 | B | 5-6 | SR5 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 9-16
57 | BJBHT.A10R5 | BLECF_ | 605 | 0x | 025D | 2 | RR57 | BY02 | 5 | CR | 2 | 6 | B | 7-8 | SR5 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16
58 | BJBHT.B10R5 | BLECF_ | 673 | 0x | 02A1 | 6 | RR57 | BY02 | 5 | CR | 2 | 7 | B | 9-10 | SR5 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 7 | 4 | 1-8 | [17 empty] | NES18
59 | BJBHT.A11R5 | BLECF_ | 621 | 0x | 026D | 8 | RR57 | BY02 | 5 | CR | 2 | 8 | B | 11-12 | SR5 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8
60 | BJBHT.B11R5 | BLECF_ | 603 | 0x | 025B | 6 | RR57 | BY02 | 5 | CR | 2 | 9 | B | 13-14 | SR5 | BY05 | C | 9 | 1-2 | BY02 | R | 3 | 7 | 4 | 9-16 | [20 empty] | All channles up to 100 blmbits @40us
61 | BJBHT.A12R5 | BLECF_ | 699 | 0x | 02BB | 6 | Arc | BYPLM.A12R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | Test in tunnel needed!
62 | BJBHT.A13R5 | BLECF_ | 728 | 0x | 02D8 | 6 | Arc | BYPLM.A13R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16
63 | BJBHT.A14R5 | BLECF_ | 698 | 0x | 02BA | 6 | Arc | BYPLM.A14R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8
64 | BJBHT.A15R5 | BLECF_ | 737 | 0x | 02E1 | 6 | Arc | BYPLM.A15R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16
65 | BJBHT.A16R5 | BLECF_ | 733 | 0x | 02DD | 6 | Arc | BYPLM.A16R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8
66 | BJBHT.A17R5 | BLECF_ | 634 | 0x | 027A | 6 | Arc | BYPLM.A17R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16
67 | BJBHT.A18R5 | BLECF_ | 732 | 0x | 02DC | 6 | Arc | BYPLM.A18R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8
68 | BJBHT.A19R5 | BLECF_ | 742 | 0x | 02E6 | 6 | Arc | BYPLM.A19R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | [20 empty] | Channel 4 (peaks up to 50 blmbits @40us)
69 | BJBHT.A20R5 | BLECF_ | 695 | 0x | 02B7 | 6 | Arc | BYPLM.A20R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | Test in tunnel needed!
70 | BJBHT.A21R5 | BLECF_ | 688 | 0x | 02B0 | 6 | Arc | BYPLM.A21R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16
71 | BJBHT.A22R5 | BLECF_ | 739 | 0x | 02E3 | 6 | Arc | BYPLM.A22R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8
72 | BJBHT.A23R5 | BLECF_ | 691 | 0x | 02B3 | 6 | Arc | BYPLM.A23R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | Test in tunnel needed!
73 | BJBHT.A24R5 | BLECF_ | 617 | 0x | 0269 | 6 | Arc | BYPLM.A24R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8
74 | BJBHT.A25R5 | BLECF_ | 736 | 0x | 02E0 | 6 | Arc | BYPLM.A25R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16
75 | BJBHT.A26R5 | BLECF_ | 683 | 0x | 02AB | 6 | Arc | BYPLM.A26R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8
76 | BJBHT.A27R5 | BLECF_ | 490 | 0x | 01EA | 6 | Arc | BYPLM.A27R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16
77 | BJBHT.A28R5 | BLECF_ | 50 | 0x | 0032 | 6 | Arc | BYPLM.A28R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8
78 | BJBHT.A29R5 | BLECF_ | 701 | 0x | 02BD | 6 | Arc | BYPLM.A29R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | [20 empty] | Blinking (all channels peaks up to 50 blmbits @40us)
79 | BJBHT.A30R5 | BLECF_ | 704 | 0x | 02C0 | 6 | Arc | BYPLM.A30R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8
80 | BJBHT.A31R5 | BLECF_ | 693 | 0x | 02B5 | 6 | Arc | BYPLM.A31R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16
81 | BJBHT.A32R5 | BLECF_ | 735 | 0x | 02DF | 6 | Arc | BYPLM.A32R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8
82 | BJBHT.A33R5 | BLECF_ | 731 | 0x | 02DB | 6 | Arc | BYPLM.A33R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 |  |  |  | ↑↑↑ | ↑↑↑ | [15 empty] | Was removed to check the input channels
83 | BJBHT.A34R5 | BLECF_ | 661 | 0x | 0295 | 6 | Arc | BYPLM.A34R5 | 5 | R |  |  |  | 7-9 | SR5 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8
Door
[25 empty] | ↑↑↑ | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on AA68: effinger: both jumper not correctly pluged, Card o.k.,
Note on AA76: effinger: jumper not correctly pluged, Card o.k.,
Note on AA79: effinger: jumper not correctly pluged, Card o.k.,
Note on AS80: Ewald Effinger: HV not correctly connected.
Note on AO51: Ewald Effinger: Checked for 5 min, I didn't see anymore the level on this channel blinkimg.
Note on AS85: Ewald Effinger: HV connector with bad
Note on AO85: Ewald Effinger: Level blinking for ch1 -6
Note on D84: Ewald Effinger: Number from the calibration list of Ion
Note on AR44: Ewald Effinger: Installed XmasStop 2011/12
Note on AU43: Ewald Effinger: Installed XmasStop 2011/12
Note on AU49: Ewald Effinger: Installed XmasStop 2011/12
Note on AU58: Ewald Effinger: Installed XmasStop 2011/12
Note on D37: Ewald Effinger: Number from the calibration list of Ion
Note on D40: Ewald Effinger: Number from the calibration list of Ion
Note on D43: Ewald Effinger: Number from the calibration list of Ion
Note on D44: Ewald Effinger: Number from the calibration list of Ion
Note on D58: Ewald Effinger: Number from the calibration list of Ion
Note on AS76: Ewald Effinger: Problem with HV
Note on AS15: Ewald Effinger: Problem HV: BLM dismanteled!
Note on AD37: Ewald Effinger: ↑↑↑ due to a broken fiber from the patch down to upstairs. Spare fiber used!
Note on T37: Ewald Effinger: C - 4 - 3 to B - 15
Note on O37: Ewald Effinger: C - 4 - 3 to B - 15
Note on AS14: Ewald Effinger: Adrey checked all connections, no o.k.
Note on AS13: Ewald Effinger: Adrey checked all connections, no o.k.
Note on AS10: Ewald Effinger: HV problem now o.k.
Note on AS12: Ewald Effinger: HV problem now o.k.
Note on AS17: Ewald Effinger: Ion and Adrien
Note on AS22: Ewald Effinger: Problem HV: BLM dismanteled!
Note on AS25: Ewald Effinger: Ion and Adrien : BLM reinstalled o.k.
Note on AS8: Ewald Effinger: HV problem
Note on AD9: Ewald Effinger: ↑↑↑ After removing and repluging the CFC it worked!
Note on AE9: Ewald Effinger: ↑↑↑
Note on AS26: Ewald Effinger: HV problems
Note on AS45: Ewald Effinger: HV problem

## Point 6
| Tunnel - position | [13 empty] | Surface position | [19 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | Comment | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L6 | BLECF_ | 533 | 0x | 0215 | 6 | Arc | BYPLM.A33L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
2 | BJBHT.A32L6 | BLECF_ | 117 | 0x | 0075 | 6 | Arc | BYPLM.A32L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
3 | BJBHT.A31L6 | BLECF_ | 232 | 0x | 00E8 | 6 | Arc | BYPLM.A31L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
4 | BJBHT.A30L6 | BLECF_ | 93 | 0x | 005D | 6 | Arc | BYPLM.A30L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
5 | BJBHT.A29L6 | BLECF_ | 163 | 0x | 00A3 | 6 | Arc | BYPLM.A29L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
6 | BJBHT.A28L6 | BLECF_ | 549 | 0x | 0225 | 6 | Arc | BYPLM.A28L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
7 | BJBHT.A27L6 | BLECF_ | 166 | 0x | 00A6 | 6 | Arc | BYPLM.A27L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | dac ovrflow
8 | BJBHT.A26L6 | BLECF_ | 266 | 0x | 010A | 6 | Arc | BYPLM.A26L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
9 | BJBHT.A25L6 | BLECF_ | 162 | 0x | 00A2 | 6 | Arc | BYPLM.A25L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
10 | BJBHT.A24L6 | BLECF_ | 168 | 0x | 00A8 | 6 | Arc | BYPLM.A24L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
11 | BJBHT.A23L6 | BLECF_ | 177 | 0x | 00B1 | 6 | Arc | BYPLM.A23L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
12 | BJBHT.A22L6 | BLECF_ | 197 | 0x | 00C5 | 6 | Arc | BYPLM.A22L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
13 | BJBHT.A21L6 | BLECF_ | 170 | 0x | 00AA | 6 | Arc | BYPLM.A21L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
14 | BJBHT.A20L6 | BLECF_ | 179 | 0x | 00B3 | 6 | Arc | BYPLM.A20L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
15 | BJBHT.A19L6 | BLECF_ | 178 | 0x | 00B2 | 6 | Arc | BYPLM.A19L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
16 | BJBHT.A18L6 | BLECF_ | 207 | 0x | 00CF | 6 | Arc | BYPLM.A18L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
17 | BJBHT.A17L6 | BLECF_ | 125 | 0x | 007D | 6 | Arc | BYPLM.A17L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
18 | BJBHT.A16L6 | BLECF_ | 708 | 0x | 02C4 | 6 | Arc | BYPLM.A16L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
19 | BJBHT.A15L6 | BLECF_ | 299 | 0x | 012B | 6 | Arc | BYPLM.A15L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
20 | BJBHT.A14L6 | BLECF_ | 281 | 0x | 0119 | 6 | Arc | BYPLM.A14L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
21 | BJBHT.A13L6 | BLECF_ | 268 | 0x | 010C | 6 | Arc | BYPLM.A13L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
22 | BJBHT.A12L6 | BLECF_ | 172 | 0x | 00AC | 6 | Arc | BYPLM.A12L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
23 | UD62 | BLECF_ | 187 | 0x | 00BB | 8 | UJ63 | BY03 | 6 | CL | 3 | 1 | B | 17-18 | SR6 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
24 | TD62.2 | BLECF_ | 255 | 0x | 00FF | 8 | UJ63 | BY03 | 6 | CL | 2 | 10 | B | 15-16 | SR6 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
25 | TD62.1 | BLECF_ | 511 | 0x | 01FF | 8 | UJ63 | BY03 | 6 | CL | 2 | 9 | B | 13-14 | SR6 | BY05 | C | 4 | 1-2 | BY02 | C | 2 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
26 | BJBHT.A11L6 | BLECF_ | 108 | 0x | 006C | 6 | UJ63 | BY03 | 6 | CL | 2 | 8 | B | 11-12 | SR6 | BY05 | C | 3 | 11-12 | BY02 | L | 1 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
27 | BJBHT.A10L6 | BLECF_ | 91 | 0x | 005B | 6 | UJ63 | BY03 | 6 | CL | 2 | 7 | B | 9-10 | SR6 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 |  |  | Problem with HV (20mA)!!!!
28 | BJBHT.A9L6 | BLECF_ | 161 | 0x | 00A1 | 6 | UJ63 | BY03 | 6 | CL | 2 | 6 | B | 7-8 | SR6 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
29 | BJBHT.A8L6 | BLECF_ | 115 | 0x | 0073 | 6 | UJ63 | BY03 | 6 | CL | 2 | 5 | B | 5-6 | SR6 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [8 empty] | NES18 |  |  | Problem with HV (20mA)!!!!
30 | BJBHT.C5L6 | BLECF_ | 49 | 0x | 0031 | 1 | UJ63 | BY03 | 6 | CL | 2 | 4 | B | 3-4 | SR6 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [11 empty] | Problem with HV (20mA)!!!!
31 | BJBHT.B5L6 | BLECF_ | 429 | 0x | 01AD | 8 | UJ63 | BY03 | 6 | CL | 2 | 3 | B | 1-2 | SR6 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
32 | BJBHT.A5L6 | BLECF_ | 107 | 0x | 006B | 3 | UJ63 | BY03 | 6 | CL | 2 | 2 | A | 23-24 | SR6 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
33 | BJBHT.K4L6 | BLECF_ | 223 | 0x | 00DF | 8 | UJ63 | BY03 | 6 | CL | 2 | 1 | A | 21-22 | SR6 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
34 | BJBHT.J4L6 | BLECF_ | 548 | 0x | 0224 | 5 | UJ63 | BY03 | 6 | CL | 1 | 10 | A | 19-20 | SR6 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
35 | BJBHT.I4L6 | BLECF_ | 508 | 0x | 01FC | 5 | UJ63 | BY03 | 6 | CL | 1 | 9 | A | 17-18 | SR6 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
36 | BJBHT.H4L6 | BLECF_ | 336 | 0x | 0150 | 3 | UJ63 | BY03 | 6 | CL | 1 | 8 | A | 15-16 | SR6 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [11 empty] | Problem with HV (20mA)!!!!
37 | BJBHT.G4L6 | BLECF_ | 66 | 0x | 0042 | 3 | UJ63 | BY03 | 6 | CL | 1 | 7 | A | 13-14 | SR6 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [11 empty] | Problem with HV (20mA)!!!!
38 | BJBHT.F4L6 | BLECF_ | 610 | 0x | 0262 | 2 | UJ63 | BY03 | 6 | CL | 1 | 6 | A | 11-12 | SR6 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
39 | BJBHT.E4L6 | BLECF_ | 392 | 0x | 0188 | 2 | UJ63 | BY03 | 6 | CL | 1 | 5 | A | 9-10 | SR6 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 |  |  | Problem with HV (20mA)!!!!
40 | BJBHT.D4L6 | BLECF_ | 612 | 0x | 0264 | 8 | UJ63 | BY03 | 6 | CL | 1 | 4 | A | 7-8 | SR6 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
41 | BJBHT.C4L6 | BLECF_ | 12 | 0x | 000C | 8 | UJ63 | BY03 | 6 | CL | 1 | 3 | A | 5-6 | SR6 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 |  |  | Problem with HV (20mA)!!!!
42 | BJBHT.B4L6 | BLECF_ | 436 | 0x | 01B4 | 7 | UJ63 | BY03 | 6 | CL | 1 | 2 | A | 3-4 | SR6 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
43 | BJBHT.A4L6 | BLECF_ | 144 | 0x | 0090 | 7 | UJ63 | BY03 | 6 | CL | 1 | 1 | A | 1-2 | SR6 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | Problem with HV (20mA)!!!!
44 | BJBHT.A4R6 | BLECF_ | 102 | 0x | 0066 | 7 | UJ67 | BY03 | 6 | CR | 1 | 1 | A | 1-2 | SR6 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
45 | BJBHT.B4R6 | BLECF_ | 116 | 0x | 0074 | 7 | UJ67 | BY03 | 6 | CR | 1 | 2 | A | 3-4 | SR6 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
46 | BJBHT.C4R6 | BLECF_ | 156 | 0x | 009C | 8 | UJ67 | BY03 | 6 | CR | 1 | 3 | A | 5-6 | SR6 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
47 | BJBHT.D4R6 | BLECF_ | 202 | 0x | 00CA | 8 | UJ67 | BY03 | 6 | CR | 1 | 4 | A | 7-8 | SR6 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
48 | BJBHT.E4R6 | BLECF_ | 167 | 0x | 00A7 | 2 | UJ67 | BY03 | 6 | CR | 1 | 5 | A | 9-10 | SR6 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
49 | BJBHT.F4R6 | BLECF_ | 545 | 0x | 0221 | 2 | UJ67 | BY03 | 6 | CR | 1 | 6 | A | 11-12 | SR6 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
50 | BJBHT.G4R6 | BLECF_ | 219 | 0x | 00DB | 3 | UJ67 | BY03 | 6 | CR | 1 | 7 | A | 13-14 | SR6 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
51 | BJBHT.H4R6 | BLECF_ | 45 | 0x | 002D | 3 | UJ67 | BY03 | 6 | CR | 1 | 8 | A | 15-16 | SR6 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
52 | BJBHT.I4R6 | BLECF_ | 173 | 0x | 00AD | 5 | UJ67 | BY03 | 6 | CR | 1 | 9 | A | 17-18 | SR6 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
53 | BJBHT.J4R6 | BLECF_ | 262 | 0x | 0106 | 4 | UJ67 | BY03 | 6 | CR | 1 | 10 | A | 19-20 | SR6 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
54 | BJBHT.K4R6 | BLECF_ | 198 | 0x | 00C6 | 8 | UJ67 | BY03 | 6 | CR | 2 | 1 | A | 21-22 | SR6 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
55 | BJBHT.A5R6 | BLECF_ | 289 | 0x | 0121 | 3 | UJ67 | BY03 | 6 | CR | 2 | 2 | A | 23-24 | SR6 | BY05 | C | 11 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
56 | BJBHT.B5R6 | BLECF_ | 165 | 0x | 00A5 | 8 | UJ67 | BY03 | 6 | CR | 2 | 3 | B | 1-2 | SR6 | BY05 | C | 10 | 1-2 | BY02 | R | 3 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
57 | BJBHT.C5R6 | BLECF_ | 88 | 0x | 0058 | 1 | UJ67 | BY03 | 6 | CR | 2 | 4 | B | 3-4 | SR6 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
58 | BJBHT.A8R6 | BLECF_ | 159 | 0x | 009F | 6 | UJ67 | BY03 | 6 | CR | 2 | 5 | B | 5-6 | SR6 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
59 | BJBHT.A9R6 | BLECF_ | 542 | 0x | 021E | 6 | UJ67 | BY03 | 6 | CR | 2 | 6 | B | 7-8 | SR6 | BY05 | C | 10 | 7-8 | BY02 | R | 3 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
60 | BJBHT.A10R6 | BLECF_ | 164 | 0x | 00A4 | 6 | UJ67 | BY03 | 6 | CR | 2 | 7 | B | 9-10 | SR6 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
61 | BJBHT.A11R6 | BLECF_ | 844 | 0x | 034C | 6 | UJ67 | BY03 | 6 | CR | 2 | 8 | B | 11-12 | SR6 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
62 | TD68.1 | BLECF_ | 391 | 0x | 0187 | 8 | UJ67 | BY03 | 6 | CR | 2 | 9 | B | 13-14 | SR6 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
63 | TD68.2 | BLECF_ | 56 | 0x | 0038 | 8 | UJ67 | BY03 | 6 | CR | 2 | 10 | B | 15-16 | SR6 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
64 | UD68 | BLECF_ | 480 | 0x | 01E0 | 8 | UJ67 | BY03 | 6 | CR | 3 | 1 | B | 17-18 | SR6 | BY05 | C | 9 | 5-6 | BY02 | C | 2 | 20 | 16 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
65 | BJBHT.A12R6 | BLECF_ | 611 | 0x | 0263 | 6 | Arc | BYPLM.A12R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [12 empty] | was 527, new comes from (13r3)
66 | BJBHT.A13R6 | BLECF_ | 730 | 0x | 02DA | 6 | Arc | BYPLM.A13R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
67 | BJBHT.A14R6 | BLECF_ | 532 | 0x | 0214 | 6 | Arc | BYPLM.A14R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
68 | BJBHT.A15R6 | BLECF_ | 569 | 0x | 0239 | 6 | Arc | BYPLM.A15R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
69 | BJBHT.A16R6 | BLECF_ | 274 | 0x | 0112 | 6 | Arc | BYPLM.A16R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
70 | BJBHT.A17R6 | BLECF_ | 301 | 0x | 012D | 6 | Arc | BYPLM.A17R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
71 | BJBHT.A18R6 | BLECF_ | 19 | 0x | 0013 | 6 | Arc | BYPLM.A18R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
72 | BJBHT.A19R6 | BLECF_ | 563 | 0x | 0233 | 6 | Arc | BYPLM.A19R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0
73 | BJBHT.A20R6 | BLECF_ | 566 | 0x | 0236 | 6 | Arc | BYPLM.A20R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
74 | BJBHT.A21R6 | BLECF_ | 295 | 0x | 0127 | 6 | Arc | BYPLM.A21R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
75 | BJBHT.A22R6 | BLECF_ | 242 | 0x | 00F2 | 6 | Arc | BYPLM.A22R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
76 | BJBHT.A23R6 | BLECF_ | 540 | 0x | 021C | 6 | Arc | BYPLM.A23R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
77 | BJBHT.A24R6 | BLECF_ | 234 | 0x | 00EA | 6 | Arc | BYPLM.A24R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
78 | BJBHT.A25R6 | BLECF_ | 555 | 0x | 022B | 6 | Arc | BYPLM.A25R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
79 | BJBHT.A26R6 | BLECF_ | 564 | 0x | 0234 | 6 | Arc | BYPLM.A26R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
80 | BJBHT.A27R6 | BLECF_ | 557 | 0x | 022D | 6 | Arc | BYPLM.A27R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
81 | BJBHT.A28R6 | BLECF_ | 63 | 0x | 003F | 6 | Arc | BYPLM.A28R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
82 | BJBHT.A29R6 | BLECF_ | 561 | 0x | 0231 | 6 | Arc | BYPLM.A29R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
83 | BJBHT.A30R6 | BLECF_ | 565 | 0x | 0235 | 6 | Arc | BYPLM.A30R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
84 | BJBHT.A31R6 | BLECF_ | 210 | 0x | 00D2 | 6 | Arc | BYPLM.A31R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
85 | BJBHT.A32R6 | BLECF_ | 154 | 0x | 009A | 6 | Arc | BYPLM.A32R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
86 | BJBHT.A33R6 | BLECF_ | 538 | 0x | 021A | 6 | Arc | BYPLM.A33R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
87 | BJBHT.A34R6 | BLECF_ | 176 | 0x | 00B0 | 6 | Arc | BYPLM.A34R6 | 6 | R |  |  |  | 7-9 | SR6 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on D22: effinger: BLECF_0279 repalced by BLECF_0708 at the 5.6.08 by Ion
Note on AU28: Ewald Effinger: Problem with HV (20mA)!!!! After checking the powersupply the 20mA disapeard! The cable of the first HV suppl was not connected.

## Point 7
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES-18 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L7 | BLECF_ | 69 | 0x | 0045 | 6 | Arc | BYPLM.A33L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | [17 empty] | NO
2 | BJBHT.A32L7 | BLECF_ | 122 | 0x | 007A | 6 | Arc | BYPLM.A32L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | [17 empty] | NO
3 | BJBHT.A31L7 | BLECF_ | 426 | 0x | 01AA | 6 | Arc | BYPLM.A31L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | [17 empty] | NO
4 | BJBHT.A30L7 | BLECF_ | 580 | 0x | 0244 | 6 | Arc | BYPLM.A30L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 |  |  | LF | LF | [13 empty] | NO |  | Power ?
5 | BJBHT.A29L7 | BLECF_ | 123 | 0x | 007B | 6 | Arc | BYPLM.A29L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | [17 empty] | NO
6 | BJBHT.A28L7 | BLECF_ | 20 | 0x | 0014 | 6 | Arc | BYPLM.A28L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | [17 empty] | NO
7 | BJBHT.A27L7 | BLECF_ | 140 | 0x | 008C | 6 | Arc | BYPLM.A27L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | [17 empty] | NO
8 | BJBHT.A26L7 | BLECF_ | 126 | 0x | 007E | 6 | Arc | BYPLM.A26L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | [17 empty] | NO
9 | BJBHT.A25L7 | BLECF_ | 43 | 0x | 002B | 6 | Arc | BYPLM.A25L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | [17 empty] | NO
10 | BJBHT.A24L7 | BLECF_ | 454 | 0x | 01C6 | 6 | Arc | BYPLM.A24L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | [17 empty] | NO |  | DAC Reset with HV not executed ( reset of card needed and need to be checked again)
11 | BJBHT.A23L7 | BLECF_ | 17 | 0x | 0011 | 6 | Arc | BYPLM.A23L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 |  |  | LF | [14 empty] | NO
12 | BJBHT.A22L7 | BLECF_ | 47 | 0x | 002F | 6 | Arc | BYPLM.A22L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | [17 empty] | NO |  | Fiber connection to be checked!
13 | BJBHT.A21L7 | BLECF_ | 87 | 0x | 0057 | 6 | Arc | BYPLM.A21L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | [17 empty] | NO
14 | BJBHT.A20L7 | BLECF_ | 74 | 0x | 004A | 6 | Arc | BYPLM.A20L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | [17 empty] | NO
15 | BJBHT.A19L7 | BLECF_ | 466 | 0x | 01D2 | 6 | Arc | BYPLM.A19L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | [17 empty] | NO
16 | BJBHT.A18L7 | BLECF_ | 571 | 0x | 023B | 6 | Arc | BYPLM.A18L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 |  |  | LF | LF |  |  |  | FF | [9 empty] | NO |  | Power ?
17 | BJBHT.A17L7 | BLECF_ | 334 | 0x | 014E | 6 | Arc | BYPLM.A17L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | [17 empty] | NO
18 | BJBHT.A16L7 | BLECF_ | 23 | 0x | 0017 | 6 | Arc | BYPLM.A16L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | [17 empty] | NO |  | DAC Reset with HV not executed ( reset of card needed and need to be checked again)
19 | BJBHT.A15L7 | BLECF_ | 534 | 0x | 0216 | 6 | Arc | BYPLM.A15L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | [17 empty] | NO |  | Reset needed
20 | BJBHT.A14L7 | BLECF_ | 461 | 0x | 01CD | 6 | Arc | BYPLM.A14L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | [17 empty] | NO |  | Reset needed
21 | BJBHT.A13L7 | BLECF_ | 449 | 0x | 01C1 | 6 | Arc | BYPLM.A13L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | [17 empty] | NO |  | DAC Reset with HV not executed ( reset of card needed and need to be checked again)
22 | BJBHT.A12L7 | BLECF_ | 472 | 0x | 01D8 | 6 | Arc | BYPLM.A12L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | [17 empty] | NO |  | was 0068, new comes from 12r3
23 | BJBHT.B11L7 | BLECF_ | 469 | 0x | 01D5 | 8 | RR73 | BY01 | 7 | CL | 3 | 8 | F | 1-2 | SR7 | BY05 | C | 5 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | [17 empty] | NO
24 | BJBHT.A11L7 | BLECF_ | 679 | 0x | 02A7 | 8 | RR73 | BY01 | 7 | CL | 3 | 7 | E | 5-6 | SR7 | BY05 | C | 5 | 5-6 | BY02 | E | 4 | 5 | 2 | 9-16 | [17 empty] | NO
25 | BJBHT.A10L7 | BLECF_ | 544 | 0x | 0220 | 8 | RR73 | BY01 | 7 | CL | 3 | 6 | E | 3-4 | SR7 | BY05 | C | 5 | 3-4 | BY02 | L | 1 | 7 | 4 | 1-8 | [17 empty] | NO
26 | BJBHT.B9L7 | BLECF_ | 678 | 0x | 02A6 | 8 | RR73 | BY01 | 7 | CL | 3 | 5 | E | 1-2 | SR7 | BY05 | C | 5 | 1-2 | BY02 | L | 1 | 6 | 3 | 9-16 | [7 empty] | 00 | [9 empty] | NO
27 | BJBHT.A9L7 | BLECF_ | 559 | 0x | 022F | 8 | RR73 | BY01 | 7 | CL | 3 | 4 | D | 23-24 | SR7 | BY05 | C | 4 | 11-12 | BY02 | E | 4 | 5 | 2 | 1-8 | [17 empty] | NO
28 | BJBHT.B8L7 | BLECF_ | 649 | 0x | 0289 | 6 | RR73 | BY01 | 7 | CL | 3 | 3 | D | 21-22 | SR7 | BY05 | C | 4 | 9-10 | BY02 | L | 1 | 6 | 3 | 1-8 | [17 empty] | NO
29 | BJBHT.A8L7 | BLECF_ | 462 | 0x | 01CE | 3 | RR73 | BY01 | 7 | CL | 3 | 2 | D | 19-20 | SR7 | BY05 | C | 4 | 7-8 | BY02 | E | 4 | 4 | 1 | 9-16 | [7 empty] | 00 | [9 empty] | NO
30 | BJBHT.B7L7 | BLECF_ | 513 | 0x | 0201 | 7 | RR73 | BY01 | 7 | CL | 3 | 1 | D | 17-18 | SR7 | BY05 | C | 4 | 5-6 | BY02 | L | 1 | 5 | 2 | 9-16 | [17 empty] | NO
31 | BJBHT.A7L7 | BLECF_ | 411 | 0x | 019B | 2 | RR73 | BY01 | 7 | CL | 2 | 10 | D | 15-16 | SR7 | BY05 | C | 4 | 3-4 | BY02 | E | 4 | 4 | 1 | 1-8 | [17 empty] | NO
32 | BJBHT.I6L7 | BLECF_ | 587 | 0x | 024B | 7 | RR73 | BY01 | 7 | CL | 2 | 9 | D | 13-14 | SR7 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 5 | 2 | 1-8 | [17 empty] | NO
33 | BJBHT.H6L7 | BLECF_ | 118 | 0x | 0076 | 4 | RR73 | BY01 | 7 | CL | 2 | 8 | D | 11-12 | SR7 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 11 | 8 | 9-16 | [17 empty] | NO
34 | BJBHT.G6L7 | BLECF_ | 191 | 0x | 00BF | 4 | RR73 | BY01 | 7 | CL | 2 | 7 | D | 9-10 | SR7 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 11 | 8 | 1-8 | [17 empty] | NO
35 | BJBHT.F6L7 | BLECF_ | 434 | 0x | 01B2 | 7 | RR73 | BY01 | 7 | CL | 2 | 6 | D | 7-8 | SR7 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 10 | 7 | 9-16 | [17 empty] | NO
36 | BJBHT.E6L7 | BLECF_ | 638 | 0x | 027E | 7 | RR73 | BY01 | 7 | CL | 2 | 5 | D | 5-6 | SR7 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 10 | 7 | 1-8 | [17 empty] | NO
37 | BJBHT.D6L7 | BLECF_ | 364 | 0x | 016C | 3 | RR73 | BY01 | 7 | CL | 2 | 4 | D | 3-4 | SR7 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 9 | 6 | 9-16 | [17 empty] | NO
38 | BJBHT.C6L7 | BLECF_ | 360 | 0x | 0168 | 3 | RR73 | BY01 | 7 | CL | 2 | 3 | D | 1-2 | SR7 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 9 | 6 | 1-8 | [17 empty] | NO
39 | BJBHT.B6L7 | BLECF_ | 682 | 0x | 02AA | 7 | RR73 | BY01 | 7 | CL | 2 | 2 | C | 23-24 | SR7 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 8 | 5 | 9-16 | [7 empty] | 00 | [9 empty] | NO
40 | BJBHT.A6L7 | BLECF_ | 265 | 0x | 0109 | 7 | RR73 | BY01 | 7 | CL | 2 | 1 | C | 21-22 | SR7 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8 | [17 empty] | NO
41 | BJBHT.E5L7 | BLECF_ | 96 | 0x | 0060 | 6 | RR73 | BY01 | 7 | CL | 1 | 10 | C | 19-20 | SR7 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 4 | 1 | 9-16 | [16 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
42 | BJBHT.D5L7 | BLECF_ | 94 | 0x | 005E | 6 | RR73 | BY01 | 7 | CL | 1 | 9 | C | 17-18 | SR7 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 7 | 4 | 9-16 | [7 empty] | 00 | [9 empty] | NO
43 | BJBHT.C5L7 | BLECF_ | 273 | 0x | 0111 | 6 | RR73 | BY01 | 7 | CL | 1 | 8 | C | 15-16 | SR7 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8 | [16 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
44 | BJBHT.B5L7 | BLECF_ | 590 | 0x | 024E | 8 | RR73 | BY01 | 7 | CL | 1 | 7 | C | 13-14 | SR7 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 6 | 3 | 9-16 | [17 empty] | NO
45 | BJBHT.A5L7 | BLECF_ | 244 | 0x | 00F4 | 8 | RR73 | BY01 | 7 | CL | 1 | 6 | C | 11-12 | SR7 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 1-8 | [17 empty] | NO
46 | BJBHT.E4L7 | BLECF_ | 372 | 0x | 0174 | 6 | RR73 | BY01 | 7 | CL | 1 | 5 | C | 9-10 | SR7 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 4 | 1 | 1-8 | [17 empty] | NO
47 | BJBHT.D4L7 | BLECF_ | 445 | 0x | 01BD | 6 | RR73 | BY01 | 7 | CL | 1 | 4 | C | 7-8 | SR7 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | [17 empty] | NO
48 | BJBHT.C4L7 | BLECF_ | 369 | 0x | 0171 | 6 | RR73 | BY01 | 7 | CL | 1 | 3 | C | 5-6 | SR7 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | [17 empty] | NO
49 | BJBHT.B4L7 | BLECF_ | 633 | 0x | 0279 | 6 | RR73 | BY01 | 7 | CL | 1 | 2 | C | 3-4 | SR7 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | [17 empty] | NO
50 | BJBHT.A4L7 | BLECF_ | 308 | 0x | 0134 | 6 | RR73 | BY01 | 7 | CL | 1 | 1 | C | 1-2 | SR7 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | [17 empty] | NO
51 | BJBHT.A4R7 | BLECF_ | 443 | 0x | 01BB | 6 | RR77 | BY02 | 7 | CR | 1 | 1 | C | 1-2 | SR7 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | [7 empty] | 0 | [9 empty] | NO
52 | BJBHT.B4R7 | BLECF_ | 455 | 0x | 01C7 | 6 | RR77 | BY02 | 7 | CR | 1 | 2 | C | 3-4 | SR7 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | [7 empty] | 0 | [9 empty] | NO
53 | BJBHT.C4R7 | BLECF_ | 351 | 0x | 015F | 6 | RR77 | BY02 | 7 | CR | 1 | 3 | C | 5-6 | SR7 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | [7 empty] | 0 | [9 empty] | NO
54 | BJBHT.D4R7 | BLECF_ | 431 | 0x | 01AF | 6 | RR77 | BY02 | 7 | CR | 1 | 4 | C | 7-8 | SR7 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | [17 empty] | NO
55 | BJBHT.E4R7 | BLECF_ | 420 | 0x | 01A4 | 6 | RR77 | BY02 | 7 | CR | 1 | 5 | C | 9-10 | SR7 | BY05 | C | 12 | 9-10 | BY02 | R | 3 | 4 | 1 | 1-8 | [17 empty] | NO
56 | BJBHT.A5R7 | BLECF_ | 713 | 0x | 02C9 | 8 | RR77 | BY02 | 7 | CR | 1 | 6 | C | 11-12 | SR7 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 1-8 | [17 empty] | NO |  | 2c9
57 | BJBHT.B5R7 | BLECF_ | 446 | 0x | 01BE | 8 | RR77 | BY02 | 7 | CR | 1 | 7 | C | 13-14 | SR7 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 15 | 11 | 9-16 | [17 empty] | NO
58 | BJBHT.C5R7 | BLECF_ | 421 | 0x | 01A5 | 6 | RR77 | BY02 | 7 | CR | 1 | 8 | C | 15-16 | SR7 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | [7 empty] | 0 | [8 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
59 | BJBHT.D5R7 | BLECF_ | 415 | 0x | 019F | 6 | RR77 | BY02 | 7 | CR | 1 | 9 | C | 17-18 | SR7 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 16 | 12 | 9-16 | [17 empty] | NO
60 | BJBHT.E5R7 | BLECF_ | 368 | 0x | 0170 | 6 | RR77 | BY02 | 7 | CR | 1 | 10 | C | 19-20 | SR7 | BY05 | C | 11 | 7-8 | BY02 | R | 3 | 4 | 1 | 9-16 | [17 empty] | NO
61 | BJBHT.A6R7 | BLECF_ | 192 | 0x | 00C0 | 4 | RR77 | BY02 | 7 | CR | 2 | 1 | C | 21-22 | SR7 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 17 | 13 | 1-8 | [7 empty] | 0 | [9 empty] | NO |  | cd
62 | BJBHT.B6R7 | BLECF_ | 221 | 0x | 00DD | 4 | RR77 | BY02 | 7 | CR | 2 | 2 | C | 23-24 | SR7 | BY05 | C | 11 | 11-12 | BY02 | C | 2 | 17 | 13 | 9-16 | [17 empty] | NO |  | dd
63 | BJBHT.C6R7 | BLECF_ | 228 | 0x | 00E4 | 5 | RR77 | BY02 | 7 | CR | 2 | 3 | D | 1-2 | SR7 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 18 | 14 | 1-8 | [17 empty] | NO |  | (e4) HV cable needed jumper on the place where BLM was removed!
64 | BJBHT.D6R7 | BLECF_ | 245 | 0x | 00F5 | 5 | RR77 | BY02 | 7 | CR | 2 | 4 | D | 3-4 | SR7 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 18 | 14 | 9-16 | [17 empty] | NO |  | f5
65 | BJBHT.E6R7 | BLECF_ | 31 | 0x | 001F | 8 | RR77 | BY02 | 7 | CR | 2 | 5 | D | 5-6 | SR7 | BY05 | C | 10 | 5-6 | BY02 | C | 2 | 19 | 15 | 1-8 | [17 empty] | NO |  | 1f
66 | BJBHT.F6R7 | BLECF_ | 13 | 0x | 000D | 8 | RR77 | BY02 | 7 | CR | 2 | 6 | D | 7-8 | SR7 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 19 | 15 | 9-16 | [17 empty] | NO |  | Fiber connection to be checked! (Ewald: Adrien cleaned the fibers and checked if the GOH was proper connected)
67 | BJBHT.G6R7 | BLECF_ | 433 | 0x | 01B1 | 4 | RR77 | BY02 | 7 | CR | 2 | 7 | D | 9-10 | SR7 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 20 | 16 | 1-8 | [17 empty] | NO
68 | BJBHT.H6R7 | BLECF_ | 482 | 0x | 01E2 | 4 | RR77 | BY02 | 7 | CR | 2 | 8 | D | 11-12 | SR7 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 20 | 16 | 9-16 | [7 empty] | 0 | [9 empty] | NO
69 | BJBHT.I6R7 | BLECF_ | 464 | 0x | 01D0 | 7 | RR77 | BY02 | 7 | CR | 2 | 9 | D | 13-14 | SR7 | BY05 | C | 9 | 1-2 | BY02 | R | 3 | 5 | 2 | 1-8 | [17 empty] | NO
70 | BJBHT.A7R7 | BLECF_ | 359 | 0x | 0167 | 2 | RR77 | BY02 | 7 | CR | 2 | 10 | D | 15-16 | SR7 | BY05 | C | 9 | 3-4 | BY02 | E | 4 | 13 | 9 | 1-8 | [17 empty] | NO |  | BNC cable between HV divider and signal box had a shortcut!
71 | BJBHT.B7R7 | BLECF_ | 470 | 0x | 01D6 | 7 | RR77 | BY02 | 7 | CR | 3 | 1 | D | 17-18 | SR7 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 5 | 2 | 9-16 | [17 empty] | NO |  | HV not o.k. (Ewald: HV Cable was wrong connected
72 | BJBHT.A8R7 | BLECF_ | 700 | 0x | 02BC | 3 | RR77 | BY02 | 7 | CR | 3 | 2 | D | 19-20 | SR7 | BY05 | C | 9 | 7-8 | BY02 | E | 4 | 13 | 9 | 9-16 | [17 empty] | NO
73 | BJBHT.B8R7 | BLECF_ | 358 | 0x | 0166 | 6 | RR77 | BY02 | 7 | CR | 3 | 3 | D | 21-22 | SR7 | BY05 | C | 9 | 9-10 | BY02 | R | 3 | 6 | 3 | 1-8 | [17 empty] | NO |  | 17d
74 | BJBHT.A9R7 | BLECF_ | 505 | 0x | 01F9 | 8 | RR77 | BY02 | 7 | CR | 3 | 4 | D | 23-24 | SR7 | BY05 | C | 9 | 11-12 | BY02 | E | 4 | 14 | 10 | 1-8 | [17 empty] | NO
75 | BJBHT.B9R7 | BLECF_ | 345 | 0x | 0159 | 8 | RR77 | BY02 | 7 | CR | 3 | 5 | E | 1-2 | SR7 | BY05 | C | 8 | 1-2 | BY02 | R | 3 | 6 | 3 | 9-16 | [17 empty] | NO |  | 1b9 (Ewald: HV cable problem bad contact afterwards noise was gone)
76 | BJBHT.A10R7 | BLECF_ | 381 | 0x | 017D | 8 | RR77 | BY02 | 7 | CR | 3 | 6 | E | 3-4 | SR7 | BY05 | C | 8 | 3-4 | BY02 | R | 3 | 7 | 4 | 1-8 | [17 empty] | NO |  | 166
77 | BJBHT.A11R7 | BLECF_ | 428 | 0x | 01AC | 8 | RR77 | BY02 | 7 | CR | 3 | 7 | E | 5-6 | SR7 | BY05 | C | 8 | 5-6 | BY02 | E | 4 | 14 | 10 | 9-16 | [17 empty] | NO |  | broken fiber surface
78 | BJBHT.B11R7 | BLECF_ | 441 | 0x | 01B9 | 7 | RR77 | BY02 | 7 | CR | 3 | 8 | F | 1-2 | SR7 | BY05 | C | 7 | 1-2 | BY02 | R | 3 | 7 | 4 | 9-16 | [7 empty] | 159 | [8 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
79 | BJBHT.A12R7 | BLECF_ | 283 | 0x | 011B | 6 | Arc | BYPLM.A12R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | [7 empty] | 11b | [9 empty] | NO |  | DAC Reset with HV not executed ( reset of card needed and need to be checked again)
80 | BJBHT.A13R7 | BLECF_ | 716 | 0x | 02CC | 6 | Arc | BYPLM.A13R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | [17 empty] | NO
81 | BJBHT.A14R7 | BLECF_ | 44 | 0x | 002C | 6 | Arc | BYPLM.A14R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | [17 empty] | NO
82 | BJBHT.A15R7 | BLECF_ | 264 | 0x | 0108 | 6 | Arc | BYPLM.A15R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | [17 empty] | NO
83 | BJBHT.A16R7 | BLECF_ | 413 | 0x | 019D | 6 | Arc | BYPLM.A16R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | [17 empty] | NO
84 | BJBHT.A17R7 | BLECF_ | 467 | 0x | 01D3 | 6 | Arc | BYPLM.A17R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | [17 empty] | NO |  | DAC Reset with HV not executed ( reset of card needed and need to be checked again)
85 | BJBHT.A18R7 | BLECF_ | 263 | 0x | 0107 | 6 | Arc | BYPLM.A18R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | [17 empty] | NO
86 | BJBHT.A19R7 | BLECF_ | 291 | 0x | 0123 | 6 | Arc | BYPLM.A19R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | [17 empty] | NO
87 | BJBHT.A20R7 | BLECF_ | 236 | 0x | 00EC | 6 | Arc | BYPLM.A20R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | [17 empty] | NO
88 | BJBHT.A21R7 | BLECF_ | 158 | 0x | 009E | 6 | Arc | BYPLM.A21R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | [17 empty] | NO
89 | BJBHT.A22R7 | BLECF_ | 256 | 0x | 0100 | 6 | Arc | BYPLM.A22R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | [17 empty] | NO
90 | BJBHT.A23R7 | BLECF_ | 65 | 0x | 0041 | 6 | Arc | BYPLM.A23R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | [17 empty] | NO
91 | BJBHT.A24R7 | BLECF_ | 152 | 0x | 0098 | 6 | Arc | BYPLM.A24R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | [17 empty] | NO
92 | BJBHT.A25R7 | BLECF_ | 220 | 0x | 00DC | 6 | Arc | BYPLM.A25R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | [17 empty] | NO
93 | BJBHT.A26R7 | BLECF_ | 133 | 0x | 0085 | 6 | Arc | BYPLM.A26R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | [17 empty] | NO
94 | BJBHT.A27R7 | BLECF_ | 138 | 0x | 008A | 6 | Arc | BYPLM.A27R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | [17 empty] | NO
95 | BJBHT.A28R7 | BLECF_ | 238 | 0x | 00EE | 6 | Arc | BYPLM.A28R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | [17 empty] | NO
96 | BJBHT.A29R7 | BLECF_ | 216 | 0x | 00D8 | 6 | Arc | BYPLM.A29R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | [17 empty] | NO
97 | BJBHT.A30R7 | BLECF_ | 174 | 0x | 00AE | 6 | Arc | BYPLM.A30R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | [17 empty] | NO
98 | BJBHT.A31R7 | BLECF_ | 132 | 0x | 0084 | 6 | Arc | BYPLM.A31R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | [17 empty] | NO
99 | BJBHT.A32R7 | BLECF_ | 397 | 0x | 018D | 6 | Arc | BYPLM.A32R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | [17 empty] | NO
100 | BJBHT.A33R7 | BLECF_ | 2 | 0x | 0002 | 6 | Arc | BYPLM.A33R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | [17 empty] | NO
101 | BJBHT.A34R7 | BLECF_ | 486 | 0x | 01E6 | 6 | Arc | BYPLM.A34R7 | 7 | R |  |  |  | 7-9 | SR7 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | [17 empty] | NO |  | DAC Reset with HV not executed ( reset of card needed and need to be checked again)
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on D62: Ewald Effinger: was 412
Note on D67: Ewald Effinger: was 334
Note on D68: Ewald Effinger: was 031
Note on D69: Ewald Effinger: was 350
Note on D70: Ewald Effinger: was 081
Note on D71: Ewald Effinger: was 394
Note on AM74: Ewald Effinger: level was on, no o.k
Note on AN74: Ewald Effinger: level was on, no o.k
Note on AJ57: Ewald Effinger: evel was on, no o.k
Note on AK58: Ewald Effinger: evel was on, no o.k
Note on AL58: Ewald Effinger: evel was on, no o.k
Note on AL59: Ewald Effinger: evel was on, no o.k
Note on AN64: Ewald Effinger: evel was on, no o.k
Note on AI67: Ewald Effinger: evel was on, no o.k
Note on AJ67: Ewald Effinger: evel was on, no o.k
Note on AL67: Ewald Effinger: evel was on, no o.k
Note on AN67: Ewald Effinger: evel was on, no o.k
Note on AS69: Ewald Effinger: Problem with HV
Note on AS72: Ewald Effinger: HV connector was bad connected!
Note on AD83: Ewald Effinger: LF
Note on AS77: Ewald Effinger: Problem with HV
Note on AS81: Ewald Effinger: Problem with HV
Note on AD72: Ewald Effinger: LF
Note on AI81: Ewald Effinger: Noise on all channels)
Note on AC47: Ewald Effinger: LF
Note on AD47: Ewald Effinger: LF

## Point 8
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L8 | BLECF_ | 229 | 0x | 00E5 | 6 | Arc | BYPLM.A33L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
2 | BJBHT.A32L8 | BLECF_ | 237 | 0x | 00ED | 6 | Arc | BYPLM.A32L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
3 | BJBHT.A31L8 | BLECF_ | 292 | 0x | 0124 | 6 | Arc | BYPLM.A31L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
4 | BJBHT.A30L8 | BLECF_ | 257 | 0x | 0101 | 6 | Arc | BYPLM.A30L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
5 | BJBHT.A29L8 | BLECF_ | 215 | 0x | 00D7 | 6 | Arc | BYPLM.A29L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
6 | BJBHT.A28L8 | BLECF_ | 213 | 0x | 00D5 | 6 | Arc | BYPLM.A28L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
7 | BJBHT.A27L8 | BLECF_ | 275 | 0x | 0113 | 6 | Arc | BYPLM.A27L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
8 | BJBHT.A26L8 | BLECF_ | 277 | 0x | 0115 | 6 | Arc | BYPLM.A26L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
9 | BJBHT.A25L8 | BLECF_ | 247 | 0x | 00F7 | 6 | Arc | BYPLM.A25L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
10 | BJBHT.A24L8 | BLECF_ | 137 | 0x | 0089 | 6 | Arc | BYPLM.A24L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
11 | BJBHT.A23L8 | BLECF_ | 145 | 0x | 0091 | 6 | Arc | BYPLM.A23L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
12 | BJBHT.A22L8 | BLECF_ | 129 | 0x | 0081 | 6 | Arc | BYPLM.A22L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
13 | BJBHT.A21L8 | BLECF_ | 112 | 0x | 0070 | 6 | Arc | BYPLM.A21L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
14 | BJBHT.A20L8 | BLECF_ | 100 | 0x | 0064 | 6 | Arc | BYPLM.A20L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
15 | BJBHT.A19L8 | BLECF_ | 114 | 0x | 0072 | 6 | Arc | BYPLM.A19L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
16 | BJBHT.A18L8 | BLECF_ | 293 | 0x | 0125 | 6 | Arc | BYPLM.A18L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO |  | Counter reset 3.12.14 (to be checked after a view days)
17 | BJBHT.A17L8 | BLECF_ | 150 | 0x | 0096 | 6 | Arc | BYPLM.A17L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
18 | BJBHT.A16L8 | BLECF_ | 130 | 0x | 0082 | 6 | Arc | BYPLM.A16L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
19 | BJBHT.A15L8 | BLECF_ | 101 | 0x | 0065 | 6 | Arc | BYPLM.A15L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
20 | BJBHT.A14L8 | BLECF_ | 105 | 0x | 0069 | 6 | Arc | BYPLM.A14L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
21 | BJBHT.A13L8 | BLECF_ | 99 | 0x | 0063 | 6 | Arc | BYPLM.A13L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
22 | BJBHT.A12L8 | BLECF_ | 712 | 0x | 02C8 | 6 | Arc | BYPLM.A12L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
23 | BJBHT.C11L8 | BLECF_ | 95 | 0x | 005F | 6 | UA83 | BY05 | 8 | CL | 2 | 9 | C | 13-14 | SR8 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
24 | BJBHT.B11L8 | BLECF_ | 92 | 0x | 005C | 6 | UA83 | BY05 | 8 | CL | 2 | 8 | C | 11-12 | SR8 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
25 | BJBHT.A11L8 | BLECF_ | 86 | 0x | 0056 | 6 | UA83 | BY05 | 8 | CL | 2 | 7 | C | 9-10 | SR8 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 8 | 5 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
26 | BJBHT.A10L8 | BLECF_ | 67 | 0x | 0043 | 6 | UA83 | BY05 | 8 | CL | 2 | 6 | C | 7-8 | SR8 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
27 | BJBHT.A9L8 | BLECF_ | 33 | 0x | 0021 | 6 | UA83 | BY05 | 8 | CL | 2 | 5 | C | 5-6 | SR8 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
28 | BJBHT.C8L8 | BLECF_ | 224 | 0x | 00E0 | 6 | UA83 | BY05 | 8 | CL | 2 | 4 | C | 3-4 | SR8 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
29 | BJBHT.B8L8 | BLECF_ | 205 | 0x | 00CD | 6 | UA83 | BY05 | 8 | CL | 2 | 3 | C | 1-2 | SR8 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
30 | BJBHT.A8L8 | BLECF_ | 151 | 0x | 0097 | 6 | UA83 | BY05 | 8 | CL | 2 | 2 | B | 23-24 | SR8 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 7 | 4 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
31 | BJBHT.A7L8 | BLECF_ | 743 | 0x | 02E7 | 6 | UA83 | BY05 | 8 | CL | 2 | 1 | B | 21-22 | SR8 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
32 | BJBHT.B6L8 | BLECF_ | 239 | 0x | 00EF | 7 | UA83 | BY05 | 8 | CL | 1 | 10 | B | 19-20 | SR8 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
33 | BJBHT.A6L8 | BLECF_ | 231 | 0x | 00E7 | 1 | UA83 | BY05 | 8 | CL | 1 | 9 | B | 17-18 | SR8 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 7 | 4 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
34 | BJBHT.A5L8 | BLECF_ | 240 | 0x | 00F0 | 6 | UA83 | BY05 | 8 | CL | 1 | 8 | B | 15-16 | SR8 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [9 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
35 | BJBHT.C4L8 | BLECF_ | 246 | 0x | 00F6 | 8 | UA83 | BY05 | 8 | CL | 1 | 7 | B | 13-14 | SR8 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
36 | BJBHT.B4L8 | BLECF_ | 248 | 0x | 00F8 | 1 | UA83 | BY05 | 8 | CL | 1 | 6 | B | 11-12 | SR8 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
37 | BJBHT.A4L8 | BLECF_ | 243 | 0x | 00F3 | 2 | UA83 | BY05 | 8 | CL | 1 | 5 | B | 9-10 | SR8 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
38 | BJBHT.A3L8 | BLECF_ | 250 | 0x | 00FA | 8 | UA83 | BY05 | 8 | CL | 1 | 4 | B | 7-8 | SR8 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO |  | Ch 6 : BLMQI.07L8.B1E10_MQM Modulation signal small !!!
39 | BJBHT.A2L8 | BLECF_ | 261 | 0x | 0105 | 8 | UA83 | BY05 | 8 | CL | 1 | 3 | B | 5-6 | SR8 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO |  | Ch 8 : BLMEI.06L8.B2I11_TCLIB.6L8.B@ Modulation signal different to the other channels !!!
40 | BJBHT.B1L8 | BLECF_ | 269 | 0x | 010D | 5 | UA83 | BY05 | 8 | CL | 1 | 2 | B | 3-4 | SR8 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
41 | BJBHT.A1L8 | BLECF_ | 180 | 0x | 00B4 | 1 | UA83 | BY05 | 8 | CL | 1 | 1 | B | 1-2 | SR8 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO |  | Connector not porperly connected
42 | BJBHT.A1R8 | BLECF_ | 78 | 0x | 004E | 1 | UA87 | BY06 | 8 | CR | 1 | 1 | B | 1-2 | SR8 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
43 | BJBHT.B1R8 | BLECF_ | 75 | 0x | 004B | 5 | UA87 | BY06 | 8 | CR | 1 | 2 | B | 3-4 | SR8 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
44 | BJBHT.A2R8 | BLECF_ | 6 | 0x | 0006 | 8 | UA87 | BY06 | 8 | CR | 1 | 3 | B | 5-6 | SR8 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO |  | Wrong number -> cheking of card number in the tunnel
45 | BJBHT.A3R8 | BLECF_ | 528 | 0x | 0210 | 4 | UA87 | BY06 | 8 | CR | 1 | 4 | B | 7-8 | SR8 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
46 | BJBHT.A4R8 | BLECF_ | 71 | 0x | 0047 | 7 | UA87 | BY06 | 8 | CR | 1 | 5 | B | 9-10 | SR8 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | NO |  | CH2 : BLMEI.04R8.I2E20 (no modulation signal)
47 | BJBHT.B4R8 | BLECF_ | 287 | 0x | 011F | 4 | UA87 | BY06 | 8 | CR | 1 | 6 | B | 11-12 | SR8 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
48 | BJBHT.C4R8 | BLECF_ | 278 | 0x | 0116 | 1 | UA87 | BY06 | 8 | CR | 1 | 7 | B | 13-14 | SR8 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
49 | BJBHT.D4R8 | BLECF_ | 209 | 0x | 00D1 | 8 | UA87 | BY06 | 8 | CR | 1 | 8 | B | 15-16 | SR8 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
50 | BJBHT.A5R8 | BLECF_ | 233 | 0x | 00E9 | 6 | UA87 | BY06 | 8 | CR | 1 | 9 | B | 17-18 | SR8 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO |  | Channel 2 is blogged to negative! 1. Check connector of the CFC 2. disconnecting of cable
51 | BJBHT.A6R8 | BLECF_ | 267 | 0x | 010B | 6 | UA87 | BY06 | 8 | CR | 1 | 10 | B | 19-20 | SR8 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
52 | BJBHT.B6R8 | BLECF_ | 29 | 0x | 001D | 6 | UA87 | BY06 | 8 | CR | 2 | 1 | B | 21-22 | SR8 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
53 | BJBHT.C6R8 | BLECF_ | 27 | 0x | 001B | 6 | UA87 | BY06 | 8 | CR | 2 | 2 | B | 23-24 | SR8 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
54 | BJBHT.A7R8 | BLECF_ | 26 | 0x | 001A | 6 | UA87 | BY06 | 8 | CR | 2 | 3 | C | 1-2 | SR8 | BY05 | C | 10 | 1-2 | BY02 | R | 3 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
55 | BJBHT.A8R8 | BLECF_ | 24 | 0x | 0018 | 6 | UA87 | BY06 | 8 | CR | 2 | 4 | C | 3-4 | SR8 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
56 | BJBHT.B8R8 | BLECF_ | 25 | 0x | 0019 | 6 | UA87 | BY06 | 8 | CR | 2 | 5 | C | 5-6 | SR8 | BY05 | C | 10 | 5-6 | BY02 | C | 2 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
57 | BJBHT.C8R8 | BLECF_ | 127 | 0x | 007F | 6 | UA87 | BY06 | 8 | CR | 2 | 6 | C | 7-8 | SR8 | BY05 | C | 10 | 7-8 | BY02 | R | 3 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
58 | BJBHT.A9R8 | BLECF_ | 80 | 0x | 0050 | 6 | UA87 | BY06 | 8 | CR | 2 | 7 | C | 9-10 | SR8 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
59 | BJBHT.A10R8 | BLECF_ | 353 | 0x | 0161 | 6 | UA87 | BY06 | 8 | CR | 2 | 8 | C | 11-12 | SR8 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
60 | BJBHT.A11R8 | BLECF_ | 39 | 0x | 0027 | 6 | UA87 | BY06 | 8 | CR | 2 | 9 | C | 13-14 | SR8 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO |  | Wrong number -> cheking of card number in the tunnel
61 | BJBHT.B11R8 | BLECF_ | 113 | 0x | 0071 | 6 | UA87 | BY06 | 8 | CR | 2 | 10 | C | 15-16 | SR8 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO |  | Channel 2 have peaks of 400 BLMbits => peaks disappeared still some noise but this can come from the cable length
62 | BJBHT.C11R8 | BLECF_ | 300 | 0x | 012C | 6 | UA87 | BY06 | 8 | CR | 3 | 1 | C | 17-18 | SR8 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | NO |  | 1. Check connector of the CFC 2. disconnecting of cable
63 | BJBHT.A12R8 | BLECF_ | 272 | 0x | 0110 | 6 | Arc | BYPLM.A12R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO |  | Powersupply exchanged! The transmission was alternating between on/off! O.k. for now!
64 | BJBHT.A13R8 | BLECF_ | 51 | 0x | 0033 | 6 | Arc | BYPLM.A13R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
65 | BJBHT.A14R8 | BLECF_ | 54 | 0x | 0036 | 6 | Arc | BYPLM.A14R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
66 | BJBHT.A15R8 | BLECF_ | 30 | 0x | 001E | 6 | Arc | BYPLM.A15R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
67 | BJBHT.A16R8 | BLECF_ | 142 | 0x | 008E | 6 | Arc | BYPLM.A16R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 00 | [9 empty] | NO
68 | BJBHT.A17R8 | BLECF_ | 61 | 0x | 003D | 6 | Arc | BYPLM.A17R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
69 | BJBHT.A18R8 | BLECF_ | 148 | 0x | 0094 | 6 | Arc | BYPLM.A18R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 00 | [9 empty] | NO
70 | BJBHT.A19R8 | BLECF_ | 15 | 0x | 000F | 6 | Arc | BYPLM.A19R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
71 | BJBHT.A20R8 | BLECF_ | 294 | 0x | 0126 | 6 | Arc | BYPLM.A20R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
72 | BJBHT.A21R8 | BLECF_ | 286 | 0x | 011E | 6 | Arc | BYPLM.A21R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
73 | BJBHT.A22R8 | BLECF_ | 206 | 0x | 00CE | 6 | Arc | BYPLM.A22R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
74 | BJBHT.A23R8 | BLECF_ | 270 | 0x | 010E | 6 | Arc | BYPLM.A23R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
75 | BJBHT.A24R8 | BLECF_ | 131 | 0x | 0083 | 6 | Arc | BYPLM.A24R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
76 | BJBHT.A25R8 | BLECF_ | 227 | 0x | 00E3 | 6 | Arc | BYPLM.A25R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
77 | BJBHT.A26R8 | BLECF_ | 253 | 0x | 00FD | 6 | Arc | BYPLM.A26R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
78 | BJBHT.A27R8 | BLECF_ | 62 | 0x | 003E | 6 | Arc | BYPLM.A27R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
79 | BJBHT.A28R8 | BLECF_ | 153 | 0x | 0099 | 6 | Arc | BYPLM.A28R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO |  | Numbers inverted! 006A -> Checking of fiber in the surface -> Fibers mixed up in the BLETC
80 | BJBHT.A29R8 | BLECF_ | 106 | 0x | 006A | 6 | Arc | BYPLM.A29R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO |  | Numbers inverted! 0099 -> Checking of fiber in the surface -> Fibers mixed up in the BLETC
81 | BJBHT.A30R8 | BLECF_ | 252 | 0x | 00FC | 6 | Arc | BYPLM.A30R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
82 | BJBHT.A31R8 | BLECF_ | 104 | 0x | 0068 | 6 | Arc | BYPLM.A31R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
83 | BJBHT.A32R8 | BLECF_ | 271 | 0x | 010F | 6 | Arc | BYPLM.A32R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
84 | BJBHT.A33R8 | BLECF_ | 225 | 0x | 00E1 | 6 | Arc | BYPLM.A33R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
85 | BJBHT.A34R8 | BLECF_ | 297 | 0x | 0129 | 6 | Arc | BYPLM.A34R8 | 8 | R |  |  |  | 7-9 | SR8 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO |  | DAC Reset with HV not executed ( reset of card needed and need to be checked again)
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on O52: effinger: Fibre 2 : C22
Note on T52: effinger: Fibre 2: C13 10
Note on AD84: Ewald Effinger: E288 5/9/14 15:40 E288 8/9/14 17:00 o.k.
Note on AD70: Ewald Effinger: 8D11 5/9/14 15:40 8D11 8/9/14 17:30 o.k
Note on D87: Ewald Effinger: Number correct!
Note on D86: Ewald Effinger: Number correct!
Note on AT68: Ewald Effinger: In the new NES18 cable was a pin swap. Now it is o.k!
Note on AJ56: Ewald Effinger: was blogged to negative
Note on AS56: Ewald Effinger: HV was not o.k
Note on AD20: Ewald Effinger: ↑ increased slowly counter reset on the 3.2.14 11:30
Note on D36: Ewald Effinger: was BLECF_0221
Note on D46: Ewald Effinger: Was BLECF_0265
Note on AD92: Ewald Effinger: Fibers on the wrong input 3-4 instead of 1-2
Note on AS46: Ewald Effinger: HV missing
Note on AD38: Ewald Effinger: LF ↑↑↑ , fibers and CFC cleaned! O.k
Note on AD28: Ewald Effinger: LF ↑↑↑ , connector was badly plugged!
Note on AB36: Ewald Effinger: LF ↑↑↑, GOH checked
Note on AE36: Ewald Effinger: LF ↑↑↑, GOH checked
Note on AA20: Ewald Effinger: ↑ increased slowly
Note on AE20: Ewald Effinger: ↑ increased slowly

## Vertical Slice
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | other
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L2 | BLECF_ | 399 | 0x | 018F | 6 | Arc | BYPLM.A33L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16
2 | BJBHT.A32L2 | BLECF_ | 348 | 0x | 015C | 6 | Arc | BYPLM.A32L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8
3 | BJBHT.A31L2 | BLECF_ | 497 | 0x | 01F1 | 6 | Arc | BYPLM.A31L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16
4 | BJBHT.A30L2 | BLECF_ | 432 | 0x | 01B0 | 6 | Arc | BYPLM.A30L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8
5 | BJBHT.A29L2 | BLECF_ | 520 | 0x | 0208 | 6 | Arc | BYPLM.A29L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16
6 | BJBHT.A28L2 | BLECF_ | 439 | 0x | 01B7 | 6 | Arc | BYPLM.A28L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8
7 | BJBHT.A27L2 | BLECF_ | 342 | 0x | 0156 | 6 | Arc | BYPLM.A27L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16
8 | BJBHT.A26L2 | BLECF_ | 460 | 0x | 01CC | 6 | Arc | BYPLM.A26L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8
9 | BJBHT.A25L2 | BLECF_ | 388 | 0x | 0184 | 6 | Arc | BYPLM.A25L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16
10 | BJBHT.A24L2 | BLECF_ | 440 | 0x | 01B8 | 6 | Arc | BYPLM.A24L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8
11 | BJBHT.A23L2 | BLECF_ | 468 | 0x | 01D4 | 6 | Arc | BYPLM.A23L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16
12 | BJBHT.A22L2 | BLECF_ | 456 | 0x | 01C8 | 6 | Arc | BYPLM.A22L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8
13 | BJBHT.A21L2 | BLECF_ | 111 | 0x | 006F | 6 | Arc | BYPLM.A21L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16
14 | BJBHT.A20L2 | BLECF_ | 11 | 0x | 000B | 6 | Arc | BYPLM.A20L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8
15 | BJBHT.A19L2 | BLECF_ | 715 | 0x | 02CB | 6 | Arc | BYPLM.A19L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16
16 | BJBHT.A18L2 | BLECF_ | 366 | 0x | 016E | 6 | Arc | BYPLM.A18L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8
17 | BJBHT.A17L2 | BLECF_ | 427 | 0x | 01AB | 6 | Arc | BYPLM.A17L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16
18 | BJBHT.A16L2 | BLECF_ | 385 | 0x | 0181 | 6 | Arc | BYPLM.A16L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8
19 | BJBHT.A15L2 | BLECF_ | 444 | 0x | 01BC | 6 | Arc | BYPLM.A15L2 | 2 | L |  |  |  | 5-6 | SR2 | BY05 | F | 4 | 5-6 | BY02 | L | 1 | 9 | 6 | 9-16
20 | BJBHT.A14L2 | BLECF_ | 380 | 0x | 017C | 6 | Arc | BYPLM.A14L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8
21 | BJBHT.A13L2 | BLECF_ | 347 | 0x | 015B | 6 | Arc | BYPLM.A13L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16
22 | BJBHT.A12L2 | BLECF_ | 200 | 0x | 00C8 | 8 | Arc | BYPLM.A12L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8
23 | BJBHT.C11L2 | BLECF_ | 583 | 0x | 0247 | 6 | UA23 | BY02 | 2 | CL | 3 | 2 | C | 19-20 | SR2 | BY05 | C | 4 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16
24 | BJBHT.B11L2 | BLECF_ | 343 | 0x | 0157 | 6 | UA23 | BY02 | 2 | CL | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 10 | 7 | 9-16
25 | BJBHT.A11L2 | BLECF_ | 378 | 0x | 017A | 6 | UA23 | BY02 | 2 | CL | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 1-8
26 | BJBHT.B10L2 | BLECF_ | 134 | 0x | 0086 | 6 | UA23 | BY02 | 2 | CL | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 1-8
27 | BJBHT.A10L2 | BLECF_ | 50 | 0x | 0032 | 7 | UA23 | BY02 | 2 | CL | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 9-16
28 | BJBHT.A9L2 | BLECF_ | 517 | 0x | 0205 | 6 | UA23 | BY02 | 2 | CL | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 9-16
29 | BJBHT.C8L2 | BLECF_ | 357 | 0x | 0165 | 6 | UA23 | BY02 | 2 | CL | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 6 | 3 | 1-8
30 | BJBHT.B8L2 | BLECF_ | 333 | 0x | 014D | 6 | UA23 | BY02 | 2 | CL | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 9 | 6 | 1-8
31 | BJBHT.A8L2 | BLECF_ | 491 | 0x | 01EB | 6 | UA23 | BY02 | 2 | CL | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 9-16
32 | BJBHT.A7L2 | BLECF_ | 362 | 0x | 016A | 6 | UA23 | BY02 | 2 | CL | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 5 | 2 | 9-16
33 | BJBHT.C6L2 | BLECF_ | 189 | 0x | 00BD | 6 | UA23 | BY02 | 2 | CL | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 1-8
34 | BJBHT.B6L2 | BLECF_ | 344 | 0x | 0158 | 6 | UA23 | BY02 | 2 | CL | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8
35 | BJBHT.A6L2 | BLECF_ | 474 | 0x | 01DA | 6 | UA23 | BY02 | 2 | CL | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16
36 | BJBHT.A5L2 | BLECF_ | 328 | 0x | 0148 | 6 | UA23 | BY02 | 2 | CL | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16
37 | BJBHT.D4L2 | BLECF_ | 371 | 0x | 0173 | 8 | UA23 | BY02 | 2 | CL | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8
X | BJBHT.TEST | BLECF_ | 208 | 0x | 00D0 | 2 | UA23 | BY02 | 2 | CL | 3 | 10 | C | 23-24 | SR2 | BY05 | C | 4 | 11-12 | BY02 | V | 1 | 6 | 3 | 1-8
38 | BJBHT.C4L2 | BLECF_ | 181 | 0x | 00B5 | 1 | UA23 | BY02 | 2 | CL | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8
39 | BJBHT.B4L2 | BLECF_ | 320 | 0x | 0140 | 4 | UA23 | BY02 | 2 | CL | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16
40 | BJBHT.A4L2 | BLECF_ | 516 | 0x | 0204 | 7 | UA23 | BY02 | 2 | CL | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8
41 | BJBHT.A3L2 | BLECF_ | 324 | 0x | 0144 | 4 | UA23 | BY02 | 2 | CL | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16
42 | BJBHT.A2L2 | BLECF_ | 188 | 0x | 00BC | 8 | UA23 | BY02 | 2 | CL | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8
43 | BJBHT.B1L2 | BLECF_ | 510 | 0x | 01FE | 5 | UA23 | BY02 | 2 | CL | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16
44 | BJBHT.A1L2 | BLECF_ | 476 | 0x | 01DC | 1 | UA23 | BY02 | 2 | CL | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8
45 | BJBHT.A1R2 | BLECF_ | 706 | 0x | 02C2 | 1 | UA27 | BY03 | 2 | CR | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8
46 | BJBHT.B1R2 | BLECF_ | 606 | 0x | 025E | 5 | UA27 | BY03 | 2 | CR | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16
47 | BJBHT.A2R2 | BLECF_ | 675 | 0x | 02A3 | 8 | UA27 | BY03 | 2 | CR | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8
48 | BJBHT.A3R2 | BLECF_ | 641 | 0x | 0281 | 4 | UA27 | BY03 | 2 | CR | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16
49 | BJBHT.A4R2 | BLECF_ | 719 | 0x | 02CF | 5 | UA27 | BY03 | 2 | CR | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8
50 | BJBHT.B4R2 | BLECF_ | 738 | 0x | 02E2 | 2 | UA27 | BY03 | 2 | CR | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16
51 | BJBHT.C4R2 | BLECF_ | 652 | 0x | 028C | 1 | UA27 | BY03 | 2 | CR | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8
52 | BJBHT.D4R2 | BLECF_ | 665 | 0x | 0299 | 8 | UA27 | BY03 | 2 | CR | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8
53 | BJBHT.A5R2 | BLECF_ | 740 | 0x | 02E4 | 6 | UA27 | BY03 | 2 | CR | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16
54 | BJBHT.A6R2 | BLECF_ | 666 | 0x | 029A | 1 | UA27 | BY03 | 2 | CR | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16
55 | BJBHT.B6R2 | BLECF_ | 722 | 0x | 02D2 | 7 | UA27 | BY03 | 2 | CR | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 5 | 2 | 1-8
56 | BJBHT.A7R2 | BLECF_ | 658 | 0x | 0292 | 6 | UA27 | BY03 | 2 | CR | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16
57 | BJBHT.A8R2 | BLECF_ | 626 | 0x | 0272 | 6 | UA27 | BY03 | 2 | CR | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8
58 | BJBHT.B8R2 | BLECF_ | 406 | 0x | 0196 | 6 | UA27 | BY03 | 2 | CR | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16
59 | BJBHT.C8R2 | BLECF_ | 4 | 0x | 0004 | 6 | UA27 | BY03 | 2 | CR | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 1-8
60 | BJBHT.A9R2 | BLECF_ | 279 | 0x | 0117 | 6 | UA27 | BY03 | 2 | CR | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 10 | 7-8 | BY02 | R | 3 | 6 | 3 | 9-16
61 | BJBHT.A10R2 | BLECF_ | 97 | 0x | 0061 | 6 | UA27 | BY03 | 2 | CR | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 18 | 14 | 1-8
62 | BJBHT.B10R2 | BLECF_ | 721 | 0x | 02D1 | 6 | UA27 | BY03 | 2 | CR | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 7 | 4 | 1-8
63 | BJBHT.A11R2 | BLECF_ | 128 | 0x | 0080 | 6 | UA27 | BY03 | 2 | CR | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16
64 | BJBHT.B11R2 | BLECF_ | 627 | 0x | 0273 | 6 | UA27 | BY03 | 2 | CR | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8
65 | BJBHT.C11R2 | BLECF_ | 689 | 0x | 02B1 | 6 | UA27 | BY03 | 2 | CR | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 7 | 4 | 9-16
66 | BJBHT.A12R2 | BLECF_ | 681 | 0x | 02A9 | 8 | Arc | BYPLM.A12R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8
67 | BJBHT.A13R2 | BLECF_ | 687 | 0x | 02AF | 6 | Arc | BYPLM.A13R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16
68 | BJBHT.A14R2 | BLECF_ | 684 | 0x | 02AC | 6 | Arc | BYPLM.A14R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8
69 | BJBHT.A15R2 | BLECF_ | 484 | 0x | 01E4 | 6 | Arc | BYPLM.A15R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16
70 | BJBHT.A16R2 | BLECF_ | 686 | 0x | 02AE | 6 | Arc | BYPLM.A16R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8
71 | BJBHT.A17R2 | BLECF_ | 143 | 0x | 008F | 6 | Arc | BYPLM.A17R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16
72 | BJBHT.A18R2 | BLECF_ | 463 | 0x | 01CF | 6 | Arc | BYPLM.A18R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8
73 | BJBHT.A19R2 | BLECF_ | 379 | 0x | 017B | 6 | Arc | BYPLM.A19R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16
74 | BJBHT.A20R2 | BLECF_ | 382 | 0x | 017E | 6 | Arc | BYPLM.A20R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8
75 | BJBHT.A21R2 | BLECF_ | 451 | 0x | 01C3 | 6 | Arc | BYPLM.A21R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16
76 | BJBHT.A22R2 | BLECF_ | 404 | 0x | 0194 | 6 | Arc | BYPLM.A22R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8
77 | BJBHT.A23R2 | BLECF_ | 84 | 0x | 0054 | 6 | Arc | BYPLM.A23R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16
78 | BJBHT.A24R2 | BLECF_ | 418 | 0x | 01A2 | 6 | Arc | BYPLM.A24R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8
79 | BJBHT.A25R2 | BLECF_ | 390 | 0x | 0186 | 6 | Arc | BYPLM.A25R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16
80 | BJBHT.A26R2 | BLECF_ | 214 | 0x | 00D6 | 6 | Arc | BYPLM.A26R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8
81 | BJBHT.A27R2 | BLECF_ | 694 | 0x | 02B6 | 6 | Arc | BYPLM.A27R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16
82 | BJBHT.A28R2 | BLECF_ | 685 | 0x | 02AD | 6 | Arc | BYPLM.A28R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8
83 | BJBHT.A29R2 | BLECF_ | 725 | 0x | 02D5 | 6 | Arc | BYPLM.A29R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16
84 | BJBHT.A30R2 | BLECF_ | 195 | 0x | 00C3 | 6 | Arc | BYPLM.A30R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8
85 | BJBHT.A31R2 | BLECF_ | 690 | 0x | 02B2 | 6 | Arc | BYPLM.A31R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16
86 | BJBHT.A32R2 | BLECF_ | 360 | 0x | 0168 | 6 | Arc | BYPLM.A32R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8
87 | BJBHT.A33R2 | BLECF_ | 23 | 0x | 0017 | 6 | Arc | BYPLM.A33R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16
88 | BJBHT.A34R2 | BLECF_ | 724 | 0x | 02D4 | 6 | Arc | BYPLM.A34R2 | 2 | R |  |  |  | 7-9 | SR2 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on T23: effinger: Fibre 7 & 8 broken, spare fibre used
Note on F82: effinger: Checked-----> BLECF_0382 installed! Wrong number 0x017E (382) instead of 0x0188 (392) ; to be check in the tunnel; Card number checked: --> BLECF_0382 installed; sent number 382

## Overview
267 | BLECF_ | 457
551 | BLECF_ | #REF!
578 | BLECF_ | 757
435 | BLECF_ | 674
574 | BLECF_ | 753
596 | #REF! | 750
594 | BLECF_ | 749
492 | BLECF_ | 748
519 | BLECF_ | 744
502 | BLECF_ | 743
619 | BLECF_ | 742
422 | BLECF_ | 741
622 | BLECF_ | 739
615 | BLECF_ | 737
625 | BLECF_ | 736
632 | BLECF_ | 735
628 | BLECF_ | 50
616 | BLECF_ | 733
618 | BLECF_ | 732
633 | BLECF_ | 731
449 | BLECF_ | 729
613 | BLECF_ | 728
412 | BLECF_ | 727
597 | BLECF_ | 726
404 | BLECF_ | 725
409 | BLECF_ | 724
418 | BLECF_ | 717
448 | BLECF_ | 711
599 | BLECF_ | 710
64 | BLECF_ | 708
413 | BLECF_ | 705
630 | BLECF_ | 704
431 | BLECF_ | 703
415 | BLECF_ | 702
629 | BLECF_ | 701
612 | BLECF_ | 699
614 | BLECF_ | 698
601 | BLECF_ | 697
620 | BLECF_ | 695
402 | BLECF_ | 694
631 | BLECF_ | 693
623 | BLECF_ | 691
406 | BLECF_ | 690
621 | BLECF_ | 688
388 | BLECF_ | 687
391 | BLECF_ | 686
403 | BLECF_ | 685
389 | BLECF_ | 684
626 | BLECF_ | 683
651 | BLECF_ | 682
387 | BLECF_ | 681
559 | BLECF_ | 680
636 | #REF! | 679
638 | #REF! | 678
451 | BLECF_ | 677
285 | BLECF_ | 676
434 | BLECF_ | 574
609 | #REF! | 673
437 | BLECF_ | 672
595 | #REF! | 671
510 | BLECF_ | 670
426 | BLECF_ | 668
262 | BLECF_ | 667
444 | BLECF_ | 175
503 | BLECF_ | 664
598 | BLECF_ | 663
494 | BLECF_ | 662
634 | BLECF_ | 661
606 | BLECF_ | 660
512 | BLECF_ | 659
432 | BLECF_ | 657
603 | BLECF_ | 656
500 | #REF! | 655
433 | BLECF_ | 654
286 | BLECF_ | 653
516 | BLECF_ | 651
554 | BLECF_ | 650
640 | BLECF_ | 649
602 | #REF! | 648
294 | BLECF_ | 647
414 | BLECF_ | 646
421 | BLECF_ | 645
533 | BLECF_ | 57
513 | BLECF_ | 643
509 | BLECF_ | 642
572 | BLECF_ | 640
280 | BLECF_ | 639
440 | BLECF_ | 637
284 | BLECF_ | 636
279 | BLECF_ | 635
617 | BLECF_ | 634
661 | BLECF_ | 633
582 | BLECF_ | 632
501 | BLECF_ | 631
425 | BLECF_ | 720
442 | BLECF_ | 629
563 | BLECF_ | 625
417 | BLECF_ | 623
428 | BLECF_ | 622
610 | BLECF_ | 621
424 | BLECF_ | 620
410 | BLECF_ | 619
604 | BLECF_ | 383
624 | BLECF_ | 617
493 | BLECF_ | 616
507 | BLECF_ | 615
423 | BLECF_ | 614
499 | BLECF_ | 613
1 | BLECF_ | 612
469 | BLECF_ | 527
2 | BLECF_ | 610
436 | BLECF_ | 609
416 | BLECF_ | 607
608 | #REF! | 605
427 | BLECF_ | 604
611 | BLECF_ | 603
447 | BLECF_ | 602
411 | BLECF_ | 601
430 | BLECF_ | 600
511 | BLECF_ | 599
600 | BLECF_ | 598
273 | BLECF_ | 597
605 | BLECF_ | 596
276 | BLECF_ | 401
263 | BLECF_ | 594
277 | BLECF_ | 593
278 | BLECF_ | 592
508 | BLECF_ | 591
656 | BLECF_ | 590
565 | BLECF_ | 589
268 | BLECF_ | 588
644 | BLECF_ | 587
591 | BLECF_ | 586
537 | BLECF_ | 585
524 | BLECF_ | 584
549 | BLECF_ | 582
575 | BLECF_ | 581
120 | BLECF_ | 580
592 | BLECF_ | 579
514 | BLECF_ | 578
445 | BLECF_ | 577
439 | BLECF_ | 576
446 | BLECF_ | 575
438 | BLECF_ | 103
497 | BLECF_ | 572
3 | BLECF_ | 571
301 | BLECF_ | 570
4 | BLECF_ | 569
531 | BLECF_ | 568
297 | BLECF_ | 567
5 | BLECF_ | 566
6 | BLECF_ | 565
7 | BLECF_ | 564
8 | BLECF_ | 563
466 | BLECF_ | 562
9 | BLECF_ | 561
648 | BLECF_ | 638
639 | BLECF_ | 559
535 | BLECF_ | 558
10 | BLECF_ | 557
523 | BLECF_ | 235
11 | BLECF_ | 555
650 | BLECF_ | 360
505 | BLECF_ | 552
544 | #REF! | 551
265 | BLECF_ | 550
12 | BLECF_ | 549
13 | BLECF_ | 548
543 | BLECF_ | 547
566 | BLECF_ | 546
14 | BLECF_ | 545
637 | BLECF_ | 544
450 | #REF! | 543
15 | BLECF_ | 542
456 | #REF! | 541
16 | BLECF_ | 540
299 | BLECF_ | 539
17 | BLECF_ | 538
540 | BLECF_ | 537
443 | BLECF_ | 536
18 | BLECF_ | 534
19 | BLECF_ | 533
20 | BLECF_ | 532
530 | BLECF_ | 531
272 | BLECF_ | 530
21 | BLECF_ | 611
581 | BLECF_ | 525
476 | BLECF_ | 523
308 | BLECF_ | 522
585 | BLECF_ | 521
326 | BLECF_ | 520
483 | BLECF_ | 519
266 | BLECF_ | 518
349 | BLECF_ | 517
361 | BLECF_ | 516
305 | BLECF_ | 515
472 | BLECF_ | 514
642 | BLECF_ | 513
292 | BLECF_ | 512
23 | BLECF_ | 511
364 | BLECF_ | 510
525 | BLECF_ | 509
24 | BLECF_ | 508
564 | BLECF_ | 506
686 | BLECF_ | 505
470 | BLECF_ | 504
461 | #REF! | 502
283 | BLECF_ | 501
312 | BLECF_ | 500
25 | BLECF_ | #REF!
471 | BLECF_ | 499
324 | BLECF_ | 497
474 | BLECF_ | 496
318 | BLECF_ | 495
344 | BLECF_ | 583
473 | BLECF_ | 493
460 | BLECF_ | 709
352 | BLECF_ | 491
627 | BLECF_ | 490
485 | BLECF_ | 489
487 | BLECF_ | 488
454 | BLECF_ | 487
26 | BLECF_ | 486
479 | BLECF_ | 485
390 | BLECF_ | 484
680 | BLECF_ | 482
306 | BLECF_ | 481
465 | BLECF_ | 186
313 | BLECF_ | 479
314 | BLECF_ | 478
573 | BLECF_ | 477
365 | BLECF_ | 476
568 | BLECF_ | 475
356 | BLECF_ | 474
567 | BLECF_ | 473
468 | BLECF_ | 68
504 | BLECF_ | 471
683 | BLECF_ | 470
635 | BLECF_ | 469
332 | BLECF_ | 468
27 | BLECF_ | 467
29 | BLECF_ | 466
652 | #REF! | 265
681 | BLECF_ | 464
393 | BLECF_ | 463
641 | BLECF_ | 462
28 | BLECF_ | 461
329 | BLECF_ | 460
467 | BLECF_ | 459
333 | BLECF_ | 456
664 | BLECF_ | 455
30 | BLECF_ | 454
489 | BLECF_ | 453
491 | BLECF_ | 452
396 | BLECF_ | 451
31 | BLECF_ | 449
484 | BLECF_ | 448
557 | BLECF_ | 447
669 | BLECF_ | 446
659 | BLECF_ | 445
340 | BLECF_ | 444
663 | BLECF_ | 443
289 | BLECF_ | 442
690 | BLECF_ | 441
331 | BLECF_ | 440
327 | BLECF_ | 439
542 | BLECF_ | 438
288 | BLECF_ | 437
32 | BLECF_ | 436
546 | #REF! | 121
647 | BLECF_ | 434
679 | BLECF_ | 433
325 | BLECF_ | 432
666 | #REF! | 431
33 | BLECF_ | 334
34 | BLECF_ | 429
689 | BLECF_ | 428
338 | BLECF_ | 427
35 | BLECF_ | 426
36 | BLECF_ | #REF!
560 | BLECF_ | 424
310 | BLECF_ | 423
684 | BLECF_ | 700
670 | BLECF_ | 421
667 | BLECF_ | 420
561 | BLECF_ | 419
399 | BLECF_ | 418
320 | BLECF_ | 417
671 | BLECF_ | 415
555 | BLECF_ | 414
37 | BLECF_ | 413
668 | BLECF_ | 713
643 | BLECF_ | 411
464 | BLECF_ | 410
264 | BLECF_ | 409
515 | BLECF_ | 408
488 | BLECF_ | 407
541 | BLECF_ | 405
397 | BLECF_ | 404
534 | BLECF_ | 403
38 | BLECF_ | #REF!
486 | BLECF_ | 492
39 | BLECF_ | #REF!
322 | BLECF_ | 399
274 | BLECF_ | 398
40 | BLECF_ | 397
506 | BLECF_ | 395
677 | BLECF_ | 31
455 | BLECF_ | 393
41 | BLECF_ | 392
395 | BLECF_ | 382
42 | BLECF_ | 391
400 | BLECF_ | 390
550 | BLECF_ | 389
330 | BLECF_ | 388
478 | BLECF_ | 387
517 | BLECF_ | 386
339 | BLECF_ | 385
319 | BLECF_ | 384
290 | BLECF_ | 355
688 | BLECF_ | 381
341 | BLECF_ | 380
394 | BLECF_ | 379
346 | BLECF_ | 378
281 | BLECF_ | 377
287 | BLECF_ | 376
463 | BLECF_ | 375
275 | BLECF_ | 374
309 | BLECF_ | 373
658 | BLECF_ | 372
358 | BLECF_ | 371
587 | BLECF_ | 370
660 | BLECF_ | 369
672 | BLECF_ | 368
459 | BLECF_ | 367
337 | BLECF_ | 366
22 | BLECF_ | 234
348 | BLECF_ | 363
353 | BLECF_ | 362
300 | BLECF_ | 361
407 | BLECF_ | 82
682 | BLECF_ | 359
685 | BLECF_ | 358
350 | BLECF_ | 357
586 | #REF! | 18
649 | BLECF_ | 364
311 | BLECF_ | 352
665 | BLECF_ | 351
675 | BLECF_ | 228
298 | BLECF_ | 349
323 | BLECF_ | 348
342 | BLECF_ | 347
570 | BLECF_ | 346
687 | BLECF_ | 345
355 | BLECF_ | 344
345 | BLECF_ | 343
328 | BLECF_ | 342
457 | BLECF_ | 341
539 | BLECF_ | 340
458 | BLECF_ | 339
296 | BLECF_ | 338
482 | BLECF_ | 337
43 | BLECF_ | 336
552 | BLECF_ | 335
673 | BLECF_ | 192
351 | BLECF_ | 333
282 | BLECF_ | 331
553 | BLECF_ | 330
317 | BLECF_ | 329
357 | BLECF_ | 328
453 | BLECF_ | 326
293 | BLECF_ | 325
362 | BLECF_ | 324
291 | BLECF_ | 323
269 | BLECF_ | 322
321 | BLECF_ | 321
360 | BLECF_ | 320
419 | BLECF_ | 319
498 | BLECF_ | 21
307 | BLECF_ | 81
522 | BLECF_ | 316
538 | BLECF_ | 315
475 | BLECF_ | 314
584 | BLECF_ | 313
590 | BLECF_ | 42
303 | BLECF_ | 311
316 | BLECF_ | 310
480 | BLECF_ | 309
662 | BLECF_ | 308
556 | BLECF_ | 307
302 | BLECF_ | 306
429 | BLECF_ | 305
315 | BLECF_ | 317
304 | BLECF_ | 303
133 | BLECF_ | 301
44 | BLECF_ | 300
45 | BLECF_ | 299
46 | BLECF_ | #REF!
47 | BLECF_ | 297
48 | BLECF_ | 295
49 | BLECF_ | 294
50 | BLECF_ | 293
51 | BLECF_ | 292
52 | BLECF_ | 291
53 | BLECF_ | #REF!
54 | BLECF_ | 289
55 | BLECF_ | #REF!
56 | BLECF_ | 287
57 | BLECF_ | 286
58 | BLECF_ | 285
59 | BLECF_ | 284
60 | BLECF_ | 212
61 | BLECF_ | 282
62 | BLECF_ | 281
63 | BLECF_ | 280
65 | BLECF_ | 278
66 | BLECF_ | 277
67 | BLECF_ | 276
68 | BLECF_ | 275
69 | BLECF_ | 274
655 | #REF! | 273
70 | BLECF_ | 272
71 | BLECF_ | 271
72 | BLECF_ | 270
73 | BLECF_ | 269
74 | BLECF_ | 268
75 | BLECF_ | 267
76 | BLECF_ | 266
77 | BLECF_ | 180
78 | BLECF_ | 264
79 | BLECF_ | 263
80 | BLECF_ | 262
81 | BLECF_ | 261
82 | BLECF_ | #REF!
83 | BLECF_ | #REF!
84 | BLECF_ | 257
85 | BLECF_ | 256
86 | BLECF_ | 255
87 | BLECF_ | 253
88 | BLECF_ | 252
89 | BLECF_ | #REF!
90 | BLECF_ | 250
91 | BLECF_ | 249
92 | BLECF_ | 248
93 | BLECF_ | 247
94 | BLECF_ | 246
95 | BLECF_ | #REF!
96 | BLECF_ | #REF!
97 | BLECF_ | 243
98 | BLECF_ | #REF!
99 | BLECF_ | 241
100 | BLECF_ | 240
101 | BLECF_ | 239
102 | BLECF_ | 238
103 | BLECF_ | 237
104 | BLECF_ | 236
105 | BLECF_ | #REF!
106 | BLECF_ | 233
107 | BLECF_ | 232
108 | BLECF_ | 231
109 | BLECF_ | 730
110 | BLECF_ | 229
111 | BLECF_ | 528
112 | BLECF_ | 227
558 | BLECF_ | 226
113 | BLECF_ | 225
114 | BLECF_ | 224
115 | BLECF_ | 223
116 | BLECF_ | 743
117 | BLECF_ | 220
118 | BLECF_ | 219
119 | BLECF_ | 218
529 | BLECF_ | 217
121 | BLECF_ | 216
122 | BLECF_ | 215
401 | BLECF_ | 214
123 | BLECF_ | 213
441 | BLECF_ | 211
124 | BLECF_ | 210
125 | BLECF_ | 209
126 | BLECF_ | 62
127 | BLECF_ | 207
128 | BLECF_ | 206
129 | BLECF_ | 205
130 | BLECF_ | 204
131 | BLECF_ | 202
132 | BLECF_ | 201
343 | BLECF_ | 200
536 | BLECF_ | 304
134 | BLECF_ | 198
135 | BLECF_ | 197
569 | BLECF_ | 196
405 | BLECF_ | 195
270 | BLECF_ | 194
593 | #REF! | 193
548 | BLECF_ | 230
646 | BLECF_ | 191
579 | BLECF_ | 190
354 | BLECF_ | 189
363 | BLECF_ | 188
136 | BLECF_ | 187
520 | #REF! | 624
477 | BLECF_ | 185
481 | BLECF_ | 184
462 | BLECF_ | 183
452 | #REF! | 182
359 | BLECF_ | 181
137 | BLECF_ | 154
138 | BLECF_ | 179
139 | BLECF_ | 178
140 | BLECF_ | 177
141 | BLECF_ | 176
142 | BLECF_ | 174
143 | BLECF_ | 173
144 | BLECF_ | 172
145 | BLECF_ | #REF!
146 | BLECF_ | 170
147 | BLECF_ | #REF!
148 | BLECF_ | 168
149 | BLECF_ | 167
150 | BLECF_ | 166
151 | BLECF_ | 165
152 | BLECF_ | 164
153 | BLECF_ | 163
154 | BLECF_ | 162
155 | BLECF_ | 161
156 | BLECF_ | 283
157 | BLECF_ | 159
158 | BLECF_ | 158
518 | BLECF_ | 157
159 | BLECF_ | 156
160 | BLECF_ | 153
161 | BLECF_ | 152
162 | BLECF_ | 151
163 | BLECF_ | 150
607 | BLECF_ | 149
164 | BLECF_ | 148
165 | BLECF_ | 145
166 | BLECF_ | 144
392 | BLECF_ | 143
167 | BLECF_ | 142
168 | BLECF_ | 141
169 | BLECF_ | 140
170 | BLECF_ | #REF!
171 | BLECF_ | 138
172 | BLECF_ | 137
347 | BLECF_ | 98
173 | BLECF_ | 133
174 | BLECF_ | 132
175 | BLECF_ | 131
176 | BLECF_ | 130
177 | BLECF_ | 129
178 | BLECF_ | 127
179 | BLECF_ | 126
180 | BLECF_ | 125
577 | BLECF_ | 124
181 | BLECF_ | 123
182 | BLECF_ | 122
589 | BLECF_ | 72
183 | BLECF_ | 39
184 | BLECF_ | 716
645 | #REF! | 118
185 | BLECF_ | 117
186 | BLECF_ | 116
187 | BLECF_ | 115
188 | BLECF_ | 114
189 | BLECF_ | 113
190 | BLECF_ | 112
334 | BLECF_ | 111
191 | BLECF_ | #REF!
490 | BLECF_ | 109
192 | BLECF_ | 108
193 | BLECF_ | 107
194 | BLECF_ | 106
195 | BLECF_ | 105
196 | BLECF_ | 104
583 | BLECF_ | 535
197 | BLECF_ | 102
198 | BLECF_ | 101
199 | BLECF_ | 100
200 | BLECF_ | 99
201 | BLECF_ | 712
653 | #REF! | 96
202 | BLECF_ | 95
654 | BLECF_ | 94
203 | BLECF_ | 93
204 | BLECF_ | 92
205 | BLECF_ | 91
206 | BLECF_ | 88
207 | BLECF_ | 87
208 | BLECF_ | 86
295 | BLECF_ | 85
398 | BLECF_ | 55
545 | BLECF_ | 83
209 | BLECF_ | 480
676 | BLECF_ | 245
210 | BLECF_ | 80
211 | BLECF_ | 79
212 | BLECF_ | 78
213 | BLECF_ | 77
496 | BLECF_ | 76
214 | BLECF_ | 75
215 | BLECF_ | 74
216 | BLECF_ | 844
217 | BLECF_ | 71
218 | BLECF_ | 6
219 | BLECF_ | 69
220 | BLECF_ | 472
221 | BLECF_ | 67
222 | BLECF_ | 66
223 | BLECF_ | 65
224 | BLECF_ | 63
225 | BLECF_ | 61
420 | BLECF_ | 745
226 | BLECF_ | 59
227 | BLECF_ | 58
532 | BLECF_ | 57
228 | BLECF_ | 56
229 | BLECF_ | 54
230 | BLECF_ | #REF!
231 | BLECF_ | 64
232 | BLECF_ | 51
233 | BLECF_ | 49
571 | BLECF_ | 48
234 | BLECF_ | 47
235 | BLECF_ | #REF!
236 | BLECF_ | 45
237 | BLECF_ | 44
238 | BLECF_ | 43
562 | BLECF_ | 34
239 | BLECF_ | #REF!
547 | BLECF_ | 40
240 | BLECF_ | 23
495 | BLECF_ | 38
241 | BLECF_ | #REF!
242 | BLECF_ | 36
243 | BLECF_ | 35
244 | BLECF_ | 400
245 | BLECF_ | 33
674 | BLECF_ | 221
246 | BLECF_ | 30
247 | BLECF_ | 29
248 | BLECF_ | 242
249 | BLECF_ | 27
250 | BLECF_ | 26
251 | BLECF_ | 25
252 | BLECF_ | 24
408 | BLECF_ | 751
253 | BLECF_ | 22
254 | BLECF_ | 754
255 | BLECF_ | 20
256 | BLECF_ | 19
257 | BLECF_ | 353
258 | BLECF_ | 17
259 | BLECF_ | 15
576 | BLECF_ | 14
678 | BLECF_ | 13
260 | BLECF_ | 12
335 | BLECF_ | 11
527 | BLECF_ | 10
580 | BLECF_ | 9
526 | BLECF_ | 7
521 | BLECF_ | 5
588 | BLECF_ | 3
261 | BLECF_ | 2
657 | BLECF_ | 244
271 | BLECF_ | 88
336 | BLECF_ | 715
366 | BLECF_ | 706
367 | BLECF_ | 606
368 | BLECF_ | 675
369 | BLECF_ | 641
370 | BLECF_ | 719
371 | BLECF_ | 738
372 | BLECF_ | 652
373 | BLECF_ | 665
374 | BLECF_ | 740
375 | BLECF_ | 666
376 | BLECF_ | 722
377 | BLECF_ | 658
378 | BLECF_ | 626
379 | BLECF_ | 406
380 | BLECF_ | 4
381 | BLECF_ | 279
382 | BLECF_ | 97
383 | BLECF_ | 721
384 | BLECF_ | 128
385 | BLECF_ | 627
386 | BLECF_ | 689
528 | BLECF_ | 119