| LHC- - - 1999-09-22 |
| --- |
| procedures for changes of BLM system settings during LHC running period including Specification of the database structure and the software needed to configure AND MANAGE the Beam Loss Monitor system in the LHC |
| Abstract The Beam Loss Monitor system is made of about 4000 detectors spread around LHC ring. For each detector a set of operational parameters is prepared. Beam-abort thresholds are an example of safety-critical parameters. These parameters are stored in databases, modified and send to electronics when needed. This document describes the database structure and software needed to operate Beam Loss Monitor system and defines the procedures to be followed should the BLM system need to undergo changes during the LHC running period. |
| Prepared by : Mariusz Sapinski, Christos Zamantzas, Bernd Dehning, Eva Barbara Holzer, Laurette Ponce, Chris Roderick, Pascal Le Roux, Julien Mariethoz | Checked by : |  |
| History of Changes |
| Rev. No. | Date | Pages | Description of Changes |
| 0 1 2 | 2008-02-29 2008-06-04 2009-06-10 | 12 14 | Initial version Included comments from Christos and Bernd Add procedures |

Table of Contents (not compulsory, can be removed)

1.	Introduction	5

2.	System parameters	6

3.	database structure	7

4.	use cases	7

5.	software	8

5.1	lsa trim application	13

5.1.1	what the software should do	13

5.1.2	software input	13

5.1.3	performance	13

5.1.4	attributes	13

5.2	lsa expert application	14

5.2.1	what the software should do	14

5.2.2	software input	14

5.2.3	performance	14

5.2.4	attributes	14

5.2.5	procedure	15

5.3	LSA EXpert application for electronics settings	15

5.3.1	what the software should do	15

5.3.2	software input	16

5.3.3	ATTRIButes	16

5.4	laYout expert application	16

5.4.1	what the software should do	16

5.4.2	performance	16

5.4.3	attributes	16

6.	references	17

7.	Appendix	17

7.1	Parameter List for BLm system	17

7.2	database details	23

7.2.1	History of changes	23

7.2.2	Structure in layout database	23

7.2.3	structure in lsa	23

# Introduction

The main purpose of the Beam Loss Monitor (BLM) system is to protect the LHC machine components from effects induced by beam particles hitting the walls of the vacuum chamber. The major effect is damage of accelerator elements, but the system will also protect from magnet transition from superconducting to normal state (referred as ‘quench’). For fast losses [1] the BLM system is the only system protecting the superconducting magnets against damages which makes it one of the critical elements for the LHC operation.

The majority of the detectors of the BLM system will be ionisation chambers (IC). For regions with very high loss rates (e.g. beam dump line, collimation regions), secondary emission monitors (SEM) are installed in addition.

The BLM system can treat up to 6400 monitors –this is the total number of slots in the electronic crates. The number of detectors foreseen for the protection of the LHC elements is about 4000. The system has a direct connection to the Beam Interlock System (BIS) and it will trigger the beam dump if the radiation level is higher than a given threshold. Most of the channels are treated as safety-critical, i.e. the beam will not be allowed to be circulating if they are not working correctly. Additional monitors, referred as mobile monitors, can be used for radiation measurements in any region of the accelerator but will not provide input for the Beam Interlock System. From the machine-protection point of view they are considered as empty slots.

For the system to be functional, a number of parameters need to be loaded on the electronics. The components storing the settings are: threshold comparator (BLETC) cards, combiner & survey (BLECS) cards and VME crates themselves. The parameters can be divided, according to their functionality into two groups. The first group are the parameters used to perform internal checks of electronics, like checking if the correct equipment and firmware version is being used (serial numbers) or setting up the tests. The second group determines the level of radiation at which the beam dump is triggered.

In total there are almost 400 numerical parameters per monitor plus additional settings per crate and per BLETC and BLECS card. 384 of these parameters are beam-abort thresholds (32 beam energy levels and 12 signal integration times). These parameters are stored in CERN databases.

The beam-abort thresholds depend on a particular configuration i.e. position of the monitor with respect to expected loss locations (typically a worst-case loss scenario is chosen), as well as the protected element. For instance for every Main Quadrupoles (MQ) on the arc the monitors which protect these quadrupoles are placed in the same locations with respect to the magnet elements. For each of these monitors the beam loss position is expected to be localized in a similar distance from the monitor. Therefore, the whole group of monitors will have the same threshold values. Such a group is referred to as a family. It is important for the BLM experts to be able to set and tune the BLM threshold values by families.

This document describes the database structure used to store the BLM system settings and it specifies the requirements for the software needed to perform specific operations on these data. It also defines the procedures to be followed should changes become necessary to the BLM system during an LHC running period. The procedures shall guarantee the system safety and integrity and not compromise the operational availability.

The information flow in the whole system is presented in Figure 1.

Figure 1: Information flow in the BLM system.

# System parameters

The parameters of the BLM system can be divided into four groups:

- operational thresholds,

- electronics parameters (testing, system state),

- parameters steering the data acquisition (logging flags),

- general system description.

The full set of parameters is shown in Appendix Error! Reference source not found..

# database structure

There are three databases in which the BLM-system related information are stored.

MTF (Manufacturing and Test Folders travellers) database,

LHC Layout database,

LSA (LHC Software Architecture) database.

The data stored in MTF database concern every element of the system (including the spare elements). This database is foreseen to store information about the history of every element, results of the tests performed on it, where it has been installed, etc.

The data in the Layout database describe the currently installed system configuration. Some information needed to operate the BLM system exists in both databases. In the normal situation, when both databases are available, the Layout database is synchronized with information in MTF.  This synchronization must happen every time the configuration of electronic connection or the monitors themselves change. The changes in MTF are made manually (web interface) and the synchronization scripts are run on request. In case of simple change of a parameter value the synchronization between MTF and Layout databases takes about 15 minutes.

The initial filling of the databases (procedure also applicable when adding a new monitor to the system) starts with Layout where the slots for the new equipment are created. These slots are then propagated to MTF where the information concerning electronic connectivity is added. This information is propagated back to Layout database. The last step is performed because the information stored in Layout database will be used to set up the BLM system and initially this database does not contain information about electronic connections.

The Layout database does not provide history of changes therefore a synchronization of changes in Layout database with MTF will be provided to keep track of the changes in the system.

The LSA database contains the actual settings of the LHC elements. It is used to store the BLM thresholds as well as the other safety-critical settings. They are stored in so called MASTER tables (one per crate). It includes also some not safety-critical information which are intended for the Logging data concentrator. The data concerning the configuration of electronic channels is synchronised from the Layout to the LSA database. The synchronization scripts are run on demand. As LSA is the only database which will always be operational during LHC running, a set of tools will be provided which will allow performing all changes directly in LSA.

It can be concluded that the information concerning electronic connection of a monitor is stored in Layout database, while operational settings of the system are stored in LSA.

# use cases

There are different situations in which intervention in the parameter structure of the BLM system will be needed.

The interventions are of two types:

Hardware change, typically change of electronics (exchange of CFC card, switching monitor to another channel, etc.)

Change of  threshold settings

| Nr | action | Master Table | Fields affected in other tables |
| --- | --- | --- | --- |
| 1 | disable/remove monitor | Affected | BLM_INFO: connected |
| 2 | adding a monitor | Affected | BLM_INFO: monitor_name monitor_exp_name connected family_name |
| 3 | change of monitor location or electronic channel | Affected | Goes via Layout BLM_INFO: dcum crate_name crate_card_idx crate_card_chan_idx blecf_id |
| 4 | Change of monitor factor (Cm) | Only Applied table | LSA Parameter table - trim |
| 5 | Change family Master Thresholds (Tf) | Affected | family_thresholds: 12 integration times (32 records will be touched) |
| 6 | Deleting a family | Affected | Remove records from: BLM_INFO, BLM_FAMILIES and BLM_FAMILY_THRESHOLDS tables |
| 7 | Creating a family | Affected | After deleting a family |
| 8 | Splitting existing family without threshold change | Affected | Actually operations 6 and 7 must be performed in order to split family |
| 9 | Adding mobile monitor | not affected | Only BLM_MOBILE_DISPLAY_NAMES table affected |

# procedures

Here the procedures needed to perform changes, described in the use cases above, are described.

Abbreviations:

MI		manual input of data

MP		manually triggered procedure

AP		automatic procedure

M&L		measurement and logging DB

FD		fixed display

CON		concentrator

BLM rep.	BLM representative

BLM representative:

The role of the BLM representative will be filled 24/7 by one person, according to a pre-defined rota, reachable by one “BLM telephone line” (forwarded to the mobile of the BLM representative).

Agreement of BLM representative:

A written and documented agreement of the BLM representative is needed for:

- Any change in the structure of a database (MTF, Layout, LSA, M&L DB)

- Any change in RBAG (concerning BLM)

The changes shall be requested by an EMDS document using EDMS approval procedure.

Execution of procedures:

The procedures described below will be executed by the BLM representative on request from the MPP representative (for procedure xxx); or on request of the BLM representative after agreement of MPP representative/machine supervisor/ ???

The procedures can only be executed from the CCC by the BLM representative together with the machine supervisor (two persons for cross checking).

Persons involved in the update:

| Pascal | Pascal Le Roux | BE/CO | 163500 | Layout |
| --- | --- | --- | --- | --- |
| Greg | Grzegor Kruk | BE/CO | 162435 | LSA |
| Chris | Chris Roderick | BE/CO | 162084 | LSA, M&L |
| Sonia | Sonia Mallon Amerigo | GS/ASE | 78616 | MTF |
| Slava | Slava Grishin | BE/BI/BL | 163132 | MTF input |
| Julien | Julien Mariethoz | BE/CO |  | Layout |
| Mariusz | Mariusz Sapinski | BE/BI/BL | 162211 |  |
| Marek | Marek Misiowiec | BE/CO | 165333 | concentrator |
| Jonathan | Jonathan Emery | BE/BI/BL | 162193 | LSA settings |
| Christos | Christos Zamantzas | BE/BI/BL | 162170 |  |
| Responsible for applications: |
| Fabio | Fabio Follin | BE/OP | 163429 | Online display and TRIM application |

## Procedure to introduce a new monitor in the BLM system

Time from request to end of procedure: 1 day

Procedure:

| 1 | MI | Layout | Introduce new slot( DCUM, expert name (?)) | Pascal |
| --- | --- | --- | --- | --- |
| 2 |  | email | To Sonia with request for update | Pascal |
| 3 | MP | Layout to MTF | Update | Sonia |
| 4 | MI | MTF | MTF specific data | Sonia |
| 5 | MI | MTF | Settings of BLM and electronics equipment | Slava |
| 6 | AP | MTF to Layout | Update slot-equipment link and error detection (nightly) | Julien or automatic (nightly) |
| 7 |  | email | To Slava, in case of error | Julien |
| 8 | MI | Layout | Introduce missing slots | Julien |
| 9 | MP | Layout to M&L | Scripts looking for changes in Layout to update M&L, to be done before CON restart | Chris |
| 10 | MP | Layout to LSA | Update | Mariusz |
| 11 | MI | LSA | Input Thresholds and settings | Mariusz |
| 12 | MI | LSA | Input HV modulation test thresholds and settings | Jonathan |
| 13 | MP | LSA, LSA to FEC | F1, F3, F4, F5 and F6 from TRIM application (Generation of parameter space and settings, propagation to HW and online check) | Mariusz |
| 14 | MP |  | Test HV modulation (if test failed –> back to step 12) | Jonathan |
| 16 | MP | CON | Restart (from DIAMON application in CCC or by scripts) reads ‘BLM info’ from LSA and variable names from Logging DB | OP or Christos or Marek |
| 17 | MP | FD | Restart (execute after restart of CON!) | OP ?? |
| 19 | ?? | M&L to CON | Comparison of variables | ?? |

## Procedure to disable a monitor in the BLM system

Time from request to end of procedure: 1 hour

Description: A monitor (which can still be physically present) is removed from the input to the beam interlock system (BIS).

Procedure:

| 1 | MI | MTF | Setting disable flag | Slava |
| --- | --- | --- | --- | --- |
| 2 | AP or MP | MTF to Layout | Update disable flag; nightly or manually by Julien | Auto or Julien |
| 3 | MP | Layout to LSA | update | Mariusz |
| 4 |  | LSA | Verify (with trim app.) that LSA updated successfully | Mariusz |
| 5*) | MI | LSA | Input HV modulation test thresholds and settings | Jonathan |
| 6 | MP | LSA,LSA to FEC | F4, F5 and F6 from TRIM application | Mariusz |
| 7*) | MP |  | Test HV modulation (if test failed –> back to step 5) | Jonathan |
| 8*) | MP | CON | Restart (from DIAMON application in CCC or by scripts) to get ‘BLM info’ from LSA | OP or Christos or Marek |
| 9*) | MP | FD | Restart | OP |

*) Steps 5 and 7 are necessary, if the HW modulation online test of the monitor to be disables had failed (e.g. problems with the hardware).

**) Step 9 will become necessary, if the FD will implement the functionality to distinguish disabled channels. Then, it needs to be seen if step 8 is necessary to be executed before step 9.

## Procedure to remove a monitor from the database

Time from request to end of procedure: 1 day

Procedure:

| 1 | MI | MTF | Setting disable flag | Slava |
| --- | --- | --- | --- | --- |
| 2 | AP or MP | MTF to Layout | Update disable flag; nightly or manually by Julien | Auto or Julien |
| 3 | MP | Layout to LSA | Update | Mariusz |
| 4 | MI | LSA | Input HV modulation test thresholds and settings | Jonathan |
| 5 | MI | LSA | Input generic name | Mariusz |
| 6 | MP | LSA, LSA to FEC | F2, F3, F4, F5 and F6 from TRIM app. (Generation of parameter space and settings, propagation to HW and online check) | Mariusz |
| 7 | MP |  | Test HV modulation (if test failed –> back to step 4) | Jonathan |
| 9 same as 9? | MP | CON | Restart (from DIAMON application in CCC or by scripts) reads ‘BLM info’ from LSA and variable names from Logging DB | OP or Christos or Marek |
| 10 | MP | FD | Restart (execute after restart of CON!) | OP ?? |
| 11 | ?? | M&L to CON | Comparison of variables | ?? |

## Procedure after change of electronics or electronic connection (excluding moving a monitor from one crate to another)

Time from request to end of procedure: 1 day (test required, to make the procedure faster)

Procedure:

| 1 | MI | MTF | Settings of BLM and electronics equipment | Slava |
| --- | --- | --- | --- | --- |
| 2 | AP | MTF to Layout | Update slot-equipment link and error detection (nightly) | Julien |
| 3 |  | email | To Slava, in case of error | Julien |
| 4 | MI | Layout | Change structure | Julien |
| 5 | MP | Layout to LSA | Update | Mariusz |
| 6*) | MI | LSA | Input HV modulation test thresholds and settings | Jonathan |
| 7 | MA | LSA to FEC | F4, F5 and F6 From TRIM application (Re-generation of parameter space and settings, propagation to HW and online check) | Mariusz / BLM rep. ?? |
| 8*) | MP |  | Test HV modulation (if test failed –> back to step 6) | Jonathan |

*) Steps 6 and 8 are necessary in case of changes in electronics channels. They are not necessary for exchange of electronics cards.

## Procedure after moving a monitor from one crate to another

Time from request to end of procedure: 1 day (test required, to make the procedure faster)

Procedure:

| 1 | MI | MTF | Settings of BLM and electronics equipment | Slava |
| --- | --- | --- | --- | --- |
| 2 | AP | MTF to Layout | Update slot-equipment link and error detection (nightly) | Julien |
| 3 |  | email | To Slava, in case of error | Julien |
| 4 | MI | Layout | Change structure | Julien |
| 5 | MP | Layout to LSA | Update | Mariusz |
| 6 | MI | LSA | Input HV modulation test thresholds and settings | Jonathan |
| 7 | MA | LSA to FEC | F3, F4, F5 and F6 From TRIM application (Re-generation of parameter space and settings, propagation to HW and online check) | Mariusz / BLM rep. ?? |
| 8 | MP |  | Test HV modulation (if test failed –> back to step 6) | Jonathan |

## Procedure to change MASTER thresholds

Time from request to end of procedure: 1 hour

Procedure:

| 1 | MI | LSA | Thresholds | Mariusz |
| --- | --- | --- | --- | --- |
| 2 | MP | LSA to FEC | F4, F5 and F6 From TRIM application | BLM rep. ?? |

## Procedure to change the family structure

Time from request to end of procedure: 1 day

Procedure:

| 1 | MI | LSA | Family settings | Mariusz |
| --- | --- | --- | --- | --- |
| 2 | MP | LSA to FEC | F4, F5 and F6 From TRIM application | BLM rep. ?? |

# software

A program or a set of programs able to perform tasks (use cases) mentioned in the previous chapter will be created.

The possible software tasks have been divided into three programs, although in the future more or less separate programs with similar functionality can be created.

The first kind is LSA trim application, working with LSA parameters which are monitor factors (Cm in Figure 1). The second is LSA-expert application which main purpose will be setting the threshold table. The third kind of application will change other system parameters in case of changes in electronic channels. It is called here Layout-expert application because it should principally operate on Layout database, although in case of Layout/MTF databases unavailability it will operate on directly LSA.

## lsa trim application

### what the software should do

The first task of LSA trim application is changing the values of monitor factor. The trim mechanisms provide history of changes. There will be 2 modes of trimming. In the default mode a group of monitors should be trimmed together. Initially this group will correspond to BLM family, but a flexibility to define a group from any monitors should be kept.  In order to make this feature work, a new table, containing monitor grouping, will be added in LSA. In this case the trim application should also take care about management of this table, that is creating, populating and deleting groups of monitors.

In the second mode it will be possible to trim one single monitor. In this case a special flag should be stored in trim history saying that a monitor has been treated in a non-standard way (this might be a specific string added automatically to the comment to trim action).

### software input

The input should be a new value of monitor factor or a shift or factor relative to present value. By default the monitor factor should be set to recommended value (0.1). This default value should define suggested trim level. For convenience it should be the same for all monitors.

### performance

The trim software should allow to trim a group of monitors or a single monitor in a time between LHC fills.

### attributes

The trim will be allowed only when beam permit flag is off. The trimming procedure will not change the beam permit flag.

The access to the trim application will be protected by RBAC (Role-Based Access) protocol [2].

The values of monitor factor are allowed in the range 0.01-1.

## lsa expert application

This application manages BLM families and sets the values of the thresholds in LSA tables. It communicates with LSA staging tables. Before each change made on staging tables they should be populated with the values actually existing in final tables using the LSA script called LSA_BLM.POPULATE_STAGE_FROM_LSA(). The argument of this script might limit the procedure to a chosen family, group of monitors filtred on expert name or IP, for instance:

LSA_BLM.POPULATE_STAGE_FROM_LSA(p_blm_exp_name_like=>’BLMDI.%’)

will populate only monitors with expert name corresponding to the pattern ‘BLMDI.%’.

### what the software should do

The program should:

- produce a table of thresholds (32x12),

- generate a list of monitors in a family based on SQL query performed on monitor expert name,

- create a new family,

- delete a family,

- produce some kind of status containing a statistical information (for instance list families with numbers of monitors in them),

- read the existing threshold tables and visualize them,

- disable/enable a monitor or change maskable flag (only in emergency case when Layout database is not accessible – this might belong to Layout expert application).

- access (viewing and modification) of the BLECS card parameters.

### software input

The program input is:

the SQL query (executed on expert name) which defines monitors belonging to a family,

the method (set of parameterisations) to expand a seed into a table (this will change!),

the family seed (one number, in Greys, obtained from Geant4/FLUKA simulations)

### performance

Expert LSA application is not foreseen to be used during the operation; therefore the important changes in the system structure are not time-critical. The only time-critical function might be disabling a monitor, which in some cases could be done between the fills.

### attributes

This application can only be used by expert. The access to LSA tables is password-protected.

The changes in the MASTER table can be made only if the beam permit flag is off but the program can still operate on staging tables. (beam permit check on LSA level?).

The software exists in form of root scripts which communicate with LSA staging tables using as input the ASCII files defining the families. Example of such a file is presented in Figure 4.

The scripts are the following:

- LSAWrite_family.C which writes new family into the STAGE_BLM_FAMILIES and STAGE_BLM_FAMILY_THRESHOLD tables, argument is family name

- LSAAttribute_monitors_to_family.C which moves monitors from family ‘DEFAULT’ the the family used as argument,

- LSADelete_family.C which deletes the family from LSA and moves all monitors which were belonging to it into ‘DEFAULT’ family.

Figure 2. Example of  a family configuration file.

### procedure

In case of MASTER table update the changes made to stage tables by the LSA Expert Application must be completed by the following procedure:

- inspection by a special program visualizing the differences between stage and final tables,

- after accepting the changes an LSA script to send changes to final tables must be launched: LSA_BLM.MERGE_STAGE_DATA(),

- manual procedure of updating the LSA devices table by AB-CO experts must be launched (call Greg).

## LSA EXpert application for electronics settings

Because a group of system parameters are only used to control electronics settings and self-test procedures, a distinct application is foreseen for setting of these parameters.

### what the software should do

This application should allow settings of four kind of electronic-specific parameters: global, IP-specific, crate-specific and monitor-specific. They are listed below:

- Global parameters are:

- IP-specific parameters are:

- Crate-specific parameters are:

- Monitor-specific parameters are:

### software input

Application should read existing parameters from LSA database and allow to change them in a simple text mode.

### ATTRIButes

Similarly to the first LSA EXPERT application this one should be protected by the RBAC system.

## laYout expert application

This application manages the electronic structure of the BLM system. In the normal operational scenario these changes are done in MTF and then propagated to Layout and later to LSA. In order not to be blocked in machine operation in case of unavailability of MTF and/or Layout this application should be able to perform all necessary operations on LSA only. In this case a back-propagation of changes to Layout and MTF will be made.

### what the software should do

The software should allow to make changes in electronic channel of every monitor. The typical examples are:

- exchange CFC card (serial number propagated to LSA),

- connect a BLM to different channel on the same or different CFC card (channel Number and slot number propagated to LSA),

- add a new monitor to the system.

During the normal operation the Layout expert application will perform changes in the Layout/MTF database. Consecutively the changes are automatically propagated to LSA. The synchronisation between MTF and Layout exists and in this moment is in the responsibility of AB-CO. The synchronization from Layout to LSA is provided by LSA_BLM.SYNCH_LAYOUT script run by BLM experts from LSA [3].

### performance

It should be possible to make this changes in one day (including hardware intervention). In emergency case the intervention will be short (direct change of the parameters in LSA). The long case is adding a new monitor. This is a complex operation which includes creation of a new slot in Layout, propagation to MTF, back-propagation of electronic connection information to Layout and synchronization with LSA.

### attributes

The access to the application should be restricted (password protected?) to the BLM experts group.

The automatic propagation of changes from Layout to LSA staging tables should be allowed when beam permit flag is on, but the update of final tables will be allowed when beam permit flag is off.

# references

[1] Quench levels and transient beam losses in LHC magnets, J.B. Jeanneret, D.Leroy, L.Oberli, T.Trenkler, LHC Project Report 44

[2] Role-Based Access Control for the Accelerator Control System at CERN,

S.Gysin, A.D. Petrov, P.Charrue, W.Gajewski, V.Kain, K.Kostro, G. Kruk, S. Page, M. Peryt proceedings of ICALEPCS 2007

[3] Chris Roderick presentation:

http://ab-div-bdi-bl-blm.web.cern.ch/ab-div-bdi-bl-blm/Audit/audit_afternoon/BLM_Audit_LSA_Chris_roderick_DB_20080610.pps

# Appendix

## Parameter List for BLm system

The Table summarizes all the variables stored in databases and used to run the BLM system. Red fields mean that information is stored in a given database.

| Version 2008.02.07 |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Name | Value (Example) | Unit | MTF | Layout | LSA | Constraint | history required | LSA Final Table (prefix with 'LSA_' from LSA_STAGE_LOADING DB account) |  | Comments |
|  |  |  |  |  |  |  | min. | max. | MTF | LSA |  | LSA Column |  |
| Nr | Slot Identifier | BLMQI.B8L6 |  |  |  | c |  |  |  | 1 |  | MONITOR_NAME |  |
| 2 | Other Identifier | BLMQI.8L6.B2I1_MQML |  |  |  | c |  |  |  | 1 |  | MONITOR_EXP_NAME |  |
| 3 | Monitor DCUM | 16359.857 | m |  |  | c |  |  |  |  |  | DCUM |  |
| 4 | Monitor Type | IC |  |  |  |  |  |  |  |  |  |  |  |
| 5 | Monitor Identificator | HCBLM_I001-05000346 |  |  |  |  |  |  |  |  |  |  |  |
| 6 | HV Patch | BJBHT.A20L8 |  |  |  |  |  |  |  |  |  |  |  |
| 7 | BJBAP Input |  |  |  |  |  |  |  |  |  |  |  |  |
| 8 | Tunnel Mini Rack id. | HCBYPLM001_CR000009 |  |  |  |  |  |  |  |  |  |  |  |
| 9 | LC(offset) w/cable&M-r | 0.1 | pA |  |  |  |  |  |  |  |  |  |  |
| 10 | Signal via source w/cable | 107.3 | pA |  |  |  |  |  |  |  |  |  |  |
| 11 | BJBAP Cable Number |  |  |  |  |  |  |  |  |  |  |  |  |
| 12 | AnalogFE:Ra_Cr_CFC_In | BYPLM.A20L8_CFC_1 |  |  |  | c |  |  |  |  | BLM_INFO | CRATE_CARD_IDX | reformated in LSA |
| 13 | BLECF Identificator | HCBLM_I001-05000100 |  |  |  | c |  |  |  |  | BLM_INFO | BLECF_ID | reformated in LSA |
| 14 | Opt Link AFE:Cr_CFC_Out | CFC_A1;CFC_A2 |  |  |  |  |  |  |  |  |  |  |  |
| 15 | OL Patch Tunnel:Ra_Pa_In | BYPLM.A20L8_7;BYPLM.A20L8_8 |  |  |  |  |  |  |  |  |  |  |  |
| 16 | OLP Surface:Ra_Pa_Sl_Out | SR8=BY05_G_4_7;SR8=BY05_G_4_8 |  |  |  |  |  |  |  |  |  |  |  |
| 17 | OLD Acq:Ra_VME_Sl_DAB_In | SR8=BY02_1_13_9_1;SR8=BY02_1_13_9_2 |  |  |  | c |  |  |  |  | BLM_INFO | CRATE_CARD_CHAN_IDX | reformated in LSA |
| 18 | VME BLETC Id. | BLETC_057 |  |  |  |  |  |  |  |  |  |  |  |
| 19 | VME CPU Host Name | CFV-SR8-BLML |  |  |  | c |  |  |  |  | BLM_INFO | CRATE_CMW_NAME | reformated in LSA |
| 20 | VME Crate Id. | 71720015 |  |  |  |  |  |  |  |  |  |  |  |
| 21 | VME Fantray Id. | 71300069 |  |  |  |  |  |  |  |  |  |  |  |
| 22 | VME Power Supply Id. | 71470017 |  |  |  |  |  |  |  |  |  |  |  |
| 23 | VME CPU Id. | CH42501584 |  |  |  |  |  |  |  |  |  |  |  |
| 24 | VME BOBR Id. | 9 |  |  |  |  |  |  |  |  |  |  |  |
| 25 | VME BLECS Id. |  |  |  |  |  |  |  |  |  |  |  |  |
| 26 | VME BLETC Serial | 0E00002B |  |  | n |  |  |  |  | ? | BLM_INFO | BLETC_SERIAL |  |
| 27 | VME BLECS Serial | 0C00000F |  |  | n |  |  |  |  | ? | BLM_INFO | BLECS_SERIAL |  |
| 28 | Monitor Function | M |  |  |  | c |  |  |  |  | BLM_INFO | MASK | reformated in LSA |
| 29 | Monitor Status | 1 |  |  |  | c |  |  |  |  | BLM_INFO | STATUS |  |
| 30 | VME Fantray Hostname | CFVM-SR8-BLML |  |  |  |  |  |  |  |  |  |  |  |
| 31 | DAB card Id. | BLM_C06421114 |  |  |  |  |  |  |  |  |  |  |  |
| 32 | Mezzanine card Id. | MEZZ_110 |  |  |  |  |  |  |  |  |  |  |  |
| 33 | CIBU Id. | HCCIBUS-ID000238 |  |  |  |  |  |  |  |  |  |  |  |
| 34 | FPGA firmware vers(BLETC) |  |  |  |  |  |  |  |  | ? | BLM_INFO | BLETC_FIRM_V |  |
| 35 | FPGA firmware vers(BLECS) |  |  |  |  |  |  |  |  | ? | BLM_INFO | BLECS_FIRM_V |  |
| 36 | Test:100pA/84 sec run.sum |  | pA |  |  |  |  |  |  |  |  |  |  |
| 37 | Test:10pA/84 sec run.sum |  | pA |  |  |  |  |  |  |  |  |  |  |
| 38 | Set voltage:peak-peak |  | V |  |  |  |  |  |  |  |  |  |  |
| 39 | Set frequency |  | Hz |  |  |  |  |  |  |  |  |  |  |
| 40 | Resulting phase |  | ° |  |  |  |  |  |  |  |  |  |  |
| 41 | Resulting Gain |  |  |  |  |  |  |  |  |  |  |  |  |
| 42 | Current signal:peak-peak |  | pA |  |  |  |  |  |  |  |  |  |  |
| 43 | Signal Current |  | pA |  |  |  |  |  |  |  |  |  |  |
| 44 | Threshold/Ch Matrix Test |  |  |  |  |  |  |  |  |  |  |  |  |
| 45 | TTable User Beam Permit T |  |  |  |  |  |  |  |  |  |  |  |  |
| 46 | UBP Check BLETC to BLECS |  |  |  |  |  |  |  |  |  |  |  |  |
| 47 | UBP Check BLECS to CIBU |  |  |  |  |  |  |  |  |  |  |  |  |
| 48 | Test EMC/40 µs run.sum |  | pA |  |  |  |  |  |  |  |  |  |  |
| 49 | Test EMC/1.3 sec run.sum |  | pA |  |  |  |  |  |  |  |  |  |  |
| 50 | EDMS Procedure | 877031 |  |  |  |  |  |  |  |  |  |  |  |
| 51 | Modulation 1 Voltage | 128 | Volt_Bit |  |  |  | 25 | 250 |  |  | BLM_IP_PARAMETERS | MOD_1_VOLTAGE | Unique per IP |
| 52 | Modulation 1 Period | 4883 | Per_Bit |  |  |  | 488 | 48828 |  |  | BLM_IP_PARAMETERS | MOD_1_PERIOD | Unique per IP |
| 53 | Modulation 1 Gain minumum | 5 |  |  |  |  | 2 | 30 |  |  | BLM_INFO | MOD_1_GAIN_MIN | Unique per monitor |
| 54 | Modulation 1 Gain maximum | 50 |  |  |  |  | 20 | 2000 |  |  | BLM_INFO | MOD_1_GAIN_MAX | Unique per monitor |
| 55 | Modulation 1 Phase minumum | 10 | ° |  |  |  | 10 | 100 |  |  | BLM_INFO | MOD_1_PHASE_MIN | Unique per monitor |
| 56 | Modulation 1 Phase maximum | 100 | ° |  |  |  | 100 | 300 |  |  | BLM_INFO | MOD_1_PHASE_MAX | Unique per monitor |
| 57 | Modulation 2 Voltage | 64 | Volt_Bit |  |  |  | 25 | 250 |  |  | BLM_IP_PARAMETERS | MOD_2_VOLTAGE | Unique per IP |
| 58 | Modulation 2 Period | 16275 | Per_Bit |  |  |  | 488 | 48828 |  |  | BLM_IP_PARAMETERS | MOD_2_PERIOD | Unique per IP |
| 59 | Modulation 2 Gain minumum | 5 |  |  |  |  | 2 | 30 |  |  | BLM_INFO | MOD_2_GAIN_MIN | Unique per monitor |
| 60 | Modulation 2 Gain maximum | 50 |  |  |  |  | 20 | 2000 |  |  | BLM_INFO | MOD_2_GAIN_MAX | Unique per monitor |
| 61 | Modulation 2 Phase minumum | 10 | ° |  |  |  | 5 | 100 |  |  | BLM_INFO | MOD_2_PHASE_MIN | Unique per monitor |
| 62 | Modulation 2 Phase maximum | 100 | ° |  |  |  | 50 | 250 |  |  | BLM_INFO | MOD_2_PHASE_MAX | Unique per monitor |
| 63 | Modulation Voltage conversion factor | 11.72/300 | V/Volt_Bit |  |  |  | 0.023 | 0.118 |  |  | BLM_COMMON_PARAMETERS | CONV_FACTOR_MOD_VOLTAGE | Common to all |
| 64 | Modulation Period conversion factor | 2048/10**6 | s/Per_Bit |  |  |  | 0.002 | 0.003 |  |  | BLM_COMMON_PARAMETERS | CONV_FACTOR_MOD_PERIOD | Common to all |
| 65 | Modulation Gain conversion factor |  |  |  |  |  |  |  |  |  | BLM_COMMON_PARAMETERS | CONV_FACTOR_MOD_GAIN | Common to all |
| 66 | Modulation Phase conversion factor |  |  |  |  |  |  |  |  |  | BLM_COMMON_PARAMETERS | CONV_FACTOR_MOD_PHASE | Common to all |
| 67 | HV normal operation | 32768 | HV_Bit |  |  |  | 30580 | 34960 |  |  | BLM_IP_PARAMETERS | NORMAL_OP_HV | Unique per IP |
| 68 | HV1 - HV2 normal operation | 437 | HV_Bit |  |  |  | 218 | 2180 |  |  | BLM_IP_PARAMETERS | NORMAL_OP_HV1_HV2_DIFF | Unique per IP |
| 69 | HV for Modulation test | 34500 | HV_Bit |  |  |  | 33860 | 38230 |  |  | BLM_IP_PARAMETERS | MOD_TEST_HV | Unique per IP |
| 70 | HV for HVCFC test | 34952 | HV_Bit |  |  |  | 33860 | 38230 |  |  | BLM_IP_PARAMETERS | HVCFC_TEST_HV | Unique per IP |
| 71 | HV for RDAC reset | 37137 | HV_Bit |  |  |  | 37680 | 42060 |  |  | BLM_IP_PARAMETERS | RDAC_RESET_HV | Unique per IP |
| 72 | HV for RGOH reset | 39342 | HV_Bit |  |  |  | 41500 | 44570 |  |  | BLM_IP_PARAMETERS | RGOH_RESET_HV | Unique per IP |
| 73 | HV peak for HVCFC test | 36000 | HV_Bit |  |  |  | 33860 | 39330 |  |  | BLM_IP_PARAMETERS | HVCFC_TEST_HV_PEAK | Unique per IP |
| 74 | HV peak for RDAC reset | 38000 | HV_Bit |  |  |  | 37680 | 43150 |  |  | BLM_IP_PARAMETERS | RDAC_RESET_HV_PEAK | Unique per IP |
| 75 | HV peak for RGOH reset | 41000 | HV_Bit |  |  |  | 41500 | 44570 |  |  | BLM_IP_PARAMETERS | RGOH_RESET_HV_PEAK | Unique per IP |
| 76 | HV conversion factor | 3000/2**16 | V/HV_Bit |  |  |  | 0.04 | 0.05 |  |  | BLM_COMMON_PARAMETERS | CONV_FACTOR_HV | Common to all |
| 77 | threshold table (12 values) |  | BLM_bit |  |  |  |  |  |  |  | BLM_FAMILY_THRESHOLDS | TIME_RS*_THRESHOLD |  |
| 78 | BEAM_ENERGY_POS | 1..32 |  |  |  |  |  |  |  | NA | BLM_FAMILIES | BEAM_ENERGY_POS |  |
| 79 | FAMILY_NAME | THRI.11...B1.2_MQ |  |  |  |  |  |  |  |  | BLM_FAMILIES | FAMILY_NAME |  |
| 80 | CONV_FACTOR_BLMBIT_GY_S | 3.6200E-09 |  |  |  |  |  |  |  |  | BLM_FAMILIES | CONV_FACTOR_BLMBIT_GY_S |  |
| 81 | CONV_FACTOR_GY_S_AMPS | .000054 |  |  |  |  |  |  |  |  | BLM_FAMILIES | CONV_FACTOR_GY_S_AMPS |  |
| 82 | MONITOR_FACTOR | 0.0...1.0 |  |  |  |  |  |  |  | trim | SETTING_SCALARS | TARGET | in LSA parameter table |
| 83 | IP | 01/08/08 |  |  |  |  |  |  |  | 1 | BLM_INFO | IP |  |
| 84 | LAST_MODIFICATION | 01/08/08 |  |  |  |  |  |  |  |  | BLM_INFO | LAST_MODIFICATION |  |
| 85 | LOG_RS01 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS01 |  |
| 86 | LOG_RS02 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS02 |  |
| 87 | LOG_RS03 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS03 |  |
| 88 | LOG_RS04 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS04 |  |
| 89 | LOG_RS05 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS05 |  |
| 90 | LOG_RS06 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS06 |  |
| 91 | LOG_RS07 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS07 |  |
| 92 | LOG_RS08 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS08 |  |
| 93 | LOG_RS09 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS09 |  |
| 94 | LOG_RS10 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS10 |  |
| 95 | LOG_RS11 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS11 |  |
| 96 | LOG_RS12 |  |  |  |  |  |  |  |  |  | BLM_INFO | LOG_RS12 |  |
|  | Colour Coding: |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | YELLOW: Variable not needed in database |  |  |  |  |  |  |  |  |  |  |  |
|  |  | RED: Variable available in database |  |  |  |  |  |  |  |  |  |  |  |
|  | n | RED + N: New variable to be introduced into database |  |  |  |  |  |  |  |  |  |  |
|  | Legend: |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 1 – change only if adding new normal (not mobile) monitor to the system |  |  |  |  |  |  |  |  |  |  |  |

## database details

### History of changes

The MTF database keeps a history of changes in order to provide a tool for tracking the modifications of every element. The Layout database has no built-in management of the history of changes therefore if the changes have to be made first in Layout (for instance in case of MTF unavailability) they are propagated back to MTF.  This is important for the operational phase when changes will be made only in Layout database (note that this is not true in the cases where those changes concern equipment replacements or additions. For example, if a new BLECF card is introduced then a new record must be created in MTF).

The history of the parameters stored in LSA is managed in two ways. History of monitor factors is followed by mechanisms built into the trim application while a change of any other parameter (except of Logging flags) affects the MASTER table and in this case a snapshot of the current situation in the database is made. The procedure to reconstruct the state of the database from snapshot and trim history depends on the kind of changes made in LSA. Normally, the last snapshot taken before the required date is loaded to LSA and then the trim history is re-applied until the required state of LSA is reached. If a new monitor had been added after the last LSA snapshot, the state of this monitor is not defined in the historical data. Therefore, it must be redefined manually.

### Structure in layout database

(here mainly text from Pascal and Julien…)

The main table (?) of Layout database contains information about BLM locations and corresponding expert names. It contains also monitor status (connected/unconnected), maskable/unmaskable flag and number of the channel  on BLETC card.

A special table called BLM_DATA_FOR_LSA_MV has been created in Layout DB in order to prepare data for transfer to LSA.

### structure in lsa

The parameters are stored in LSA database in a specific structure, illustrated in Figure 3.

On the left part of Figure 3 there are four tables. The LSA setting table is a table where the monitor factors are declared as LSA parameters. The three other tables are called Monitor, Family and Threshold table are used to produce the MASTER table. The APPLIED table is produced from the MASTER and from trim parameters.

The LSA structure uses data staging. That is, the tables presented in Figure 3 exist as STAGE tables and as FINAL tables. The STAGE tables are foreseen to load new parameters and perform consistency tests on them before sending them to the FINAL tables where they become accessible by the electronics. This data push from stage to final tables is triggered manually by BLM experts once they checked that the changes they have made can be send to LSA.

The default value of monitor factor is 0.1. The threshold levels in MASTER table are about 3 times above quench levels, therefore the APPLIED table is at 30% of quench level. The minimum value of monitor factor is 0.01. In case of warm elements the MASTER table stays always below the damage level.

The names used for the tables in LSA are:

Monitor table: stage_blm_monitor and blm_info

Family table: stage_blm_families and blm_families

Threshold table: stage_blm_family_thresholds and blm_family_thresholds

The content of the staging tables is presented in Figure 4.

Figure 3: Structure of LSA tables.

Figure 4: Structure of staging tables

All possible available electronic channels, i.e. 6400 channels, are defined in the BLM_INFO table. Naturally these channels are divided into 3 groups:

- normal monitors – connected to Beam Interlock System, they have expert names explained in [3]

- “mobile” monitors, used for measurements and not connected to BIS,

- empty slots, where either or both frond-end and back-end electronics are not installed.

From the safety point of view, mobile monitors and empty slots are the same. Therefore, they look the same in the LSA structure described above. Their expert name is generated automatically. They start with BLMM prefix and the rest of the expert name contains the card index and the the electronic channel index.

Additional tables in LSA used by the BLM software are:

- BLM_RS_TIMES which contains values of the integration times,

- LHC_IPS which contains list of possible interaction points,

- STAGE_BLM_COMMON_PARAMETERS and BLM_COMMON_PARAMETERS which contains BLECT card parameters common to all monitors

- STAGE_BLM_IP_PARAMETERS and BLM_IP_PARAMETERS which contain values of parameters specific to a given IP.

- STAGE_BLM_MOBILE_DISPLAY_NAMES and BLM_MOBILE_DISPLAY_NAMES which contains an attribution of non-generic names to mobile monitors, what allows to register data from mobile monitors in LSA/Measurement database without making any changes in safety-critical part of LSA.

Figure 5:  Final tables

# Footnotes

There is a small group of monitors on the dump line with expert names starting with BLMPI or BLMPS which are not connected to BIS but to dump kicker magnets.