| LHC-BCTDC-TP-0001.00 rev0.1 AB-BI 751016 2007-06-12 |
| --- |
| Test Procedure |
| The commissioning of beam instrumentation in the LHC Sectors INDIVIDUAL SYSTEM TESTS OF THE LHC Beam Loss [BLM] Monitors |
| Abstract The tests are meant to validate the correct functioning of the LHC Beam Loss Monitors. This document covers: The phase of hardware commissioning following installation campaign First-time debugging SW tests without beam Self-checks during the consecutive regular cold check-outs of the complete system The commissioning with beam following machine start-ups. The purpose of these tests is to assess and insure: The entire acquisition chain from the monitor up to the front end computers The data processing and communication using public interfaces The quality of our measurement for first operations Upon termination of these commissioning phases, the LHC Beam Loss monitors should be fully qualified for first operational use. |
| Prepared by : Bernd Dehning Eva Barbara Holzer Michael Ludwig Jean-Jacques Gras Stephen Jackson AB-BI | Checked by : | Approval Leader : Roland Garoby |
| Approval List: |
| History of Changes |
| Rev. No. | Date | Pages | Description of Changes |
| 0.1 | 20.2.2008 | All | Bring the document up to standards for commissioning, enforce chapter structure |

Table of Contents

(not compulsory, can be removed)

1.	Introduction	5

2.	Purpose	5

3.	The standard infrastructure components	5

3.1	The BI VME cRAtes & Racks	6

3.2	Local WIFI ethernet	6

3.3	LOCAL Linux PC & consoles	6

3.4	THE BI ServerS	6

3.5	Error reporting infrastructure	6

3.6	The Standard Software TOOLS	6

4.	GENERAL conditions required for testing	8

4.1	Infrastructures	8

4.2	The BI VME Crates	8

5.	General Safety Issues	8

5.1	The risks induced by the tests & the compensatory measures	8

5.2	safety panels and Signs	9

5.3	Machine Protection	9

5.4	Access Conditions	9

6.	Beam Instrumentation Commissioning Planning	9

7.	Dealing with technical Problems	9

8.	Beam Loss Monitors:  SyStem Architecture	10

8.1	The monitors	10

8.2	Tunnel equipment (ARC Front-end electronics and cabling)	11

8.3	Front-end equipment for LSS monitors	11

8.4	Back-end equipment (surface electronics)	12

8.5	software	12

8.6	complete instrument	12

8.7	Software interface usage	13

9.	Definition of the tests	13

9.1	notation	13

9.2	Procedures in the MTF [in progress]	14

9.3	Scope and Limits of the procedures	14

10.	tests for the BLM	15

10.1	monitors	15

10.2	tunnel equipment (tunnel and alcoves)	15

10.3	back-end equipment	16

10.4	Software public interfaces	18

10.5	software dedicated interfaces	18

10.6	Software runtime	19

10.7	complete instrument	20

10.8	Software interface usage conformity	23

11.	ProcedureS for BLM commissioning	23

11.1	procedure for initial first-time commissioning	24

11.2	procedure for Cold checkout	26

11.3	Commissioning with Beam	27

11.4	procedure for quick self-test	28

# Introduction

Upon installation of the BLM in LHC, a series of commissioning tests will have to be performed by the instrument specialists. The objectives of these tests are:

- Update and verify documentation in MTF

- Eliminate problems as early as possible in order to minimize any impact on LHC schedules

- Establish repeatable test procedures for the hardware, the software, the whole instrument without beam and constraints and the whole instrument under operational conditions with beam

- Ensure a minimal quality and usage under nominal conditions, and a first calibration where appropriate

- Ensure required safety standards

- Quantify the conditions needed for any MD time

After successful commissioning, the instrument is delivered to operations for first usage.

The commissioning does cover instrument performance assessment the instrument is working like defined in the specification. It does not cover the optimisation, upgrades, instrument physics MDs and any cross-calibration: these issues are addressed in a separate “performance assessment”. This document will regularly be updated in the EDMS folder, so please check in EDMS  https://edms.cern.ch/document/877031/0.1 that your document version is up to date.

# Purpose

For the different stages of commissioning (i.e. hardware commissioning/cold check out, offline system commissioning and commissioning with beam), this document

Lists the system components which will be the object of the tests.

Defines the conditions required to start testing, those required during the tests, those which determine the successful completion and the infrastructure required for these tests

Describes in detail the procedures which will be applied for these tests and their sequence

Defines the expected results

This document is the reference document for the BLM MTF-checklist which will be used during these commissioning tests. By consulting the MTF the current hardware commissioning status of the instruments can be retrieved at any time.

The interfaces to the machine protection system are tested, but the overall integration of the BLMs into the MP system is covered in a separate document [1], which also in the case of the BLMs covers the performance assessment with beam.

# The standard infrastructure components

The full functionality of the standard infrastructure components is a prerequisite for instrument commissioning, and the testing of these components is out of scope of this document. Nevertheless they are explicitly listed, and for the purpose of this document are assumed to be fully functional.

## The BI VME cRAtes & Racks

Each instrument is controlled by a BI front end computer (FEC) installed in a VME crate. The standard functionality of these FECs will already have been commissioned following the procedure described in https://edms.cern.ch/document/750985 .

Timing & telegram must be functional, but their detailed configurations are part of the specific instrument.

The infrastructure provided by the racks (power, mechanical mounting, cooling & ventilation) is described in the LHC-BI-wiki ( http://bdidev1.cern.ch/bdisoft/operational/abbdilhctb_wiki/ ).

## Local WIFI ethernet

In all LHC surface buildings and underground areas a WIFI infrastructure is available allows verifying a large part of the tests foreseen in this document. It is recommended that all installations taking place around the LHC be done with a portable PC with WIFI capabilities. To allow accessing the systems on the LHC technical network it is recommended to access the BI servers via Exceed or remote desktop connection.

## LOCAL Linux PC & consoles

In addition to this WIFI infrastructure, some local Linux PCs (carrying specific SW and interfaces) and standard CO consoles (remote terminal screens) are also installed in BI strategic locations. These locations are described in http://bdidev1.cern.ch/bdisoft/operational/abbdilhctb_wiki/Ctrl/InstEtagere.

## THE BI ServerS

The BI Linux server bdisrv1 and Windows Application Server bits01.cern.ch will be used to run BI applications and store our monitoring acquisitions during the remote diagnostic tests.

## Error reporting infrastructure

The infrastructure should become functional with FESA version10, but details are still unknown.

### Post-mortem

The standard post-mortem triggering and data retrieval mechanisms are used, which are described at http://bdidev1.cern.ch/bdisoft/operational/abbdilhctb_wiki/Ctrl/PostMort .

## The Standard Software TOOLS

These tools are part of the standard software infrastructure, which is provided by CO and BI-SW in the context of the accelerator control system at CERN.

### SSH, FEC startup and monitoring

SSH will be used to log into a Linux-workstation or FEC, in order to execute programs and to monitor and control process activity without a graphical interface. It can be used from Linux (ssh user@system) or Windows ( PuTTY )

The Linux command:

/mcr/reset/rem-reset <hostname>

will be used to reboot the FEC <hostname> and monitor the boot sequences through the build-in remote console feature of the fan-tray (there is no equivalent from Windows).

### timtest

Timtest is a program provided by AB-CO-HT which can be run on the FEC to investigate and monitor machine timing reception. Only some simple features will be used by the instrument specialist to allow a quick timing diagnosis, the usage is described in the BI-SW-wiki at:

http://bdidev1.cern.ch/bdisoft/operational/abbdisw_wiki/FAQ/Timtest

### THE FESA Navigator

The FESA navigator will be used to control, monitor and diagnose the different FESA servers running in the FECs. These are mainly the LTIM (machine timing) server and specific instrument classes. The navigator can be efficiently used when navigator-automations for specific tasks on instruments are loaded and executed.

### The BLM concEntrator

The BLM concentrator is a gateway provided by CO specifically for the BLM system, in order to provide proper scaling behaviour and safety relevant stability and behaviour. This concentrator, provided by LSA, will be the only means to access the 1Hz 'Logging' data..  The concentrator will subscribe to each front-end to get data for 256 channels, republishing all valid monitors to the final clients.  In the case where a front end has less than 256 monitors installed (where some channel identifiers are marked as SPARE), the concentrator will discard the unused channels and not republish them.  This will have the effect of reducing the 6400 (256*25) possible channels, to the ~4000 which actually have monitors connected on them.  In addition to republishing the data, the concentrator will also cross-check that the channel index to monitor name mapping supplied from the front-ends, matches that defined in the LSA database.

### The ROSALI Application

The ROSALI (Rapid Online Software ALgorithm Implementation:  http://bdidev1/bdisoft/operational/applauncher.php?launch=ROSALIGUI ) application is an application program running on a Linux or Windows workstation, which permits to execute pre-defined data-acquisition and analysis sequences within the accelerator control system. This tool is provided by AB-BI-SW and it will be used to assess the quality of the instruments.

### The Beam Synchronous Timing Expert Application

The Beam Synchronous Timing expert application will allow the monitoring and diagnostic of the operational BOBR cards and their bunch and turn synchronisation.

### The AB-BI-SW web site

The LHC chapter of the AB-BI-SW section web site will contain all the necessary information to develop, run and maintain the corresponding instrument software. It will in particular

- provide a link towards this document

- provide links to the public instrument interface

- provide documentation about the software

# GENERAL conditions required for testing

## Infrastructures

Nominal operating conditions in the concerned underground and surface areas are required for:

The electrical distribution system

Ethernet

The General Machine Timing

The Beam Synchronous Timing.

All the standard CO central services (boot sequence servers, nfs and terminal servers, CCC, CCR network, database FEC configuration, FEC build chain,) necessary to boot, access and investigate any problems on the FECs are vital for the success of the tests.

We hope to be able to rely on the TS GTPM application (Gestion des Pannes Majeures) to know the current state of the standard services (Access, 220V, general timing distribution and Ethernet) in every related surface and underground zone as soon as it will be made available.

Furthermore the conditions required to start and perform these tests are:

- Interconnection of cryostats performed in the sector

- Installation of front-end mini-racks, crates, fan Unit supervised by AB/BI/BL.

- Optical fibre terminations supervised by AB/BI/BL

- BLM detector installed by supervised by AB/BI/BL

- ‘daisy-chain’ high-voltage and signal cables supervised by AB/BI/BL

- Installation of BLM electronics supervised by AB/BI/BL

## The BI VME Crates

The BI VME crates under test have to be commissioned (see https://edms.cern.ch/document/750985 for details) and operational.

# General Safety Issues

The physical interventions corresponding to these tests are covered by dedicated task descriptions (mainly task 1, 2, 6 and 7 in the hardware commissioning context) in the BI Overall Health and Safety Plan ( https://edms.cern.ch/file/713428/1/PPSPS_BI_Latest.doc ) and the corresponding ‘Authorisation d’Ouverture de Chantier’ ( https://edh.cern.ch/Document/AOC/2541612 up to 06/08/07 ).

## The risks induced by the tests & the compensatory measures

### Work in the tunnel

For the equipment installed in the underground areas, every participant will have the standard security equipments (Biocell, security shoes, helmet and lamp) and follow the orders given in the corresponding task descriptions.

### Security Training

The necessary training for the corresponding ‘Habilitation Electrique’ has been defined and requested for every participant.

## safety panels and Signs

In general, no specific panels or signs are required during these tests. Exceptions will be described in the steps concerned.

## Machine Protection

Great care must be taken during testing with beam with respect to machine protection, if the instrument implements an interface to MP. The specific conditions are laid down in written together with MP experts, and at certain testing stages the presence of a MP-specialist might be required.

## Access Conditions

In general, no access limitations are required during these tests. Exceptions will be described in the steps concerned.

# Beam Instrumentation Commissioning Planning

BI follows as much as possible the LHC planning to deploy and commission its instruments in the different sectors of the machine. The time estimation of our commissioning interventions is given in the procedure descriptions.

But, due to contingencies (HW, SW or people availability…), some steps in the procedures may have to be done after the dedicated installation or commissioning slot. Late interventions will also be necessary to deal with sudden hardware failures, or as part of problem follow-up.

All these late interventions will be made in agreement with the HWC team and when the conditions required for these interventions are met.

# Dealing with technical Problems

Technical problems appearing during installation and commissioning phases will be solved during this test period whenever possible, in an ad-hoc manner.

Faulty boards, short cables, crates… will be replaced by correct ones etc.

For long cables, spare connections will be used whenever possible

Software and communication related problems should be detected at an earlier stage if possible, so that testing the entire instrument is feasible from the start. General infrastructure problems must be well distinguished from instrument SW problems. For infrastructure problems the support of AB-CO during the tests is required.

Problems which have an impact on day-to-day operation planning are discussed ad-hoc with the engineer in charge (EIC) and/or operator in the CCC.

In case of faulty equipment, traces of the problem will be kept in the equipment MTF whenever necessary. The evolution of the commissioning can be followed with the step reports on our instruments: (https://edms.cern.ch/asbuilt/plsql/mtf_urep.start_urep?p_ure_code=48).

In case of faulty infrastructure, traces of the problem will be kept in the commissioning MTF as a non conformity whenever necessary (for instance in case of a bad optic fibre).

It will most probably not be possible to solve some of these problems during the commissioning phase. In these cases, later interventions will be necessary (see previous chapter), which will be discussed and prepared in coordination with the CWG and other relevant committees.

# Beam Loss Monitors:  SyStem Architecture

The BLM system is comprised of some 4000 monitors (3700 Ionisation Chambers and 280 Secondary Emission detector) distributed around the machine. The principle of the system is to measure outside the cryostat or the vacuum chamber the secondary shower particles created by protons or ions impacting on the vacuum chamber.  When at least one signal is over a pre-defined threshold, the BLM system triggers a beam dump to protect against damage of equipment and prevent quenches of magnet coils.

When the LHC is in “safe beam” mode some of these signals can be masked in order to avoid early beam dumps, to get higher operational availability.

The derivation of the thresholds is based on a combination of several simulations (loss maps, secondary shower development, quench level, detector response) and verification measurements.

To ensure the specified reliability level of the system, some of the tests presented here have to be done regularly (see G. Guaglio, PH-D thesis, CERN-THESIS-2006-012, CERN-LIB), namely before every fill. An overview of the system architecture can be found at [2].

## The monitors

These can be grouped into 2 distinct monitor types: ICs [3] and SEMs [4]:

Bernd: put references as footnotes

Table 8.1: The monitors of the BLM system

| Monitor type | Nb of monitors | Description |
| --- | --- | --- |
| IC | 3700 | Ionisation chambers |
| SEM | 270 | Secondary emission monitor |

- BLMQI and BLMQS: monitors at the regular places of the quadrupoles. The last letter gives the monitor type: I for IC, and S for SEM.

- BLMEI: IC installed at irregular (exceptional) places.

- BLMDI and BLMDS: IC or SEM installed in the two dump lines.

- BLMM: generic name (placeholder) for a mobile monitor (not installed). 2450 channels for mobile monitors are available.

- For the monitors which are installed in the LSS the front-end electronics is regrouped in the UA, RR, UJ or USC areas, and connected via coaxial and multi-wire cables.

There is no bakeout.

## Tunnel equipment (ARC Front-end electronics and cabling)

In the ARC the front-end electronics is placed beneath the SSS. It is comprised of one yellow mini-rack containing:

1 3U electronics crate (shared between BPM and BLM)

1 fan unit

1 fibre-optic patch panel

1 protection cover

Low voltage power supply per 3U crate

For the HV (~1500V) power supply of the monitors, there is one wall mounted patch-box per mini-rack (up to 8 monitors of the same kind IC or SEM).

In the broad sense there is no intermediate electronics.

### cabling towards monitors

- Coaxial signal cable from each monitor to arc mini-rack.

- Coaxial power cable from patch box to the first monitor.

- Coaxial power cable daisy chain between up to 8 monitors.

- Coaxial power cable from last monitor to arc mini-rack (test of HV).

### electronics

- Charge to frequency converter

- Multiplexer

- Optical data transmission to surface electronics

A group of up to 8 nearby monitors, either ICs or SEMs, is digitized, multiplexed and transmitted through a laser fibre link, using current to frequency conversion (this converter is named BLECF).

The multiplexing inside the front-end electronics takes place in the implementation, but inside the BLETC the data is digitally de-multiplexed again to obtain the data functionally per single monitor.

The data stream is transmitted to the surface threshold comparators (BLETC), using redundant optical fibre links. Two groups of each 8 monitors feed into one threshold comparator BLETC.

## Front-end equipment for LSS monitors

The LSS front-end electronics in UA, RR and UJ is grouped in the BI rack BY02. This rack contains 3U electronics crates (similar to the arc crates), a corresponding number of fan units and a fibre-optic patch panel.

For each octant two high voltage power supplies are used to supply the monitors with bias voltage, one being redundant.

### cabling towards tunnel equipment

- Coaxial signal cable from each monitor to a wall mounted patch box

- Twisted pair multi-wire signal cable from patch box to the LSS front-end electronics in UA, RR and UJ.

- Coaxial power cable from patch box to the first monitor, between the monitors and back to the LSS front-end electronics.

### electronics

Same as 8.2.2

## Back-end equipment (surface electronics)

- VME crates

- Threshold comparator cards with optical receiver and de-multiplexer

- Combiner card

- CIBU (interface to BIC)

### cabling towards tunnel equipment

Optical transmission via redundant lines from front-end electronics. The connection of the individual BLM monitor channels to the corresponding VME crates is documented in detail in the BLM slot MTF. All channels are transmitted to the collector, but spare channels which are unused are masked by software, see chapter 3.7.4.

### electronics

All digital acquisition electronics is regrouped in the SR building (SX at IR4) at each LHC access point. In these areas the electronics are grouped in the BI rack BY02. One SR building covers the corresponding octant, i.e. a half sector L and R. Hardware commissioning of a complete sector therefore requires the use of 2 operational SR infrastructures.

## software

The software interfaces, real-time behaviour and instrument usage specifications are described in the LHC Interface Software Specification (LIDS) pages at:

http://bdidev1.cern.ch/bdisoft/development/BDI-Domains/bdeyelids/bdeyelids.php?currentSelection=DE&currentDomain=LHC&currentInstrument=BLM

### device – monitor mapping

There is exactly one FESA software device for each FEC.

### synchronisation

Not applicable

### real-time performance

The real time behaviour is defined by the FESA framework.  It is foreseen to have a single process (no split real-time and communications processes) and rely on the FESA thread model for concurrency between real-time actions and servicing of clients’ requests.

## complete instrument

### calibration

A possible baseline drift of the electronics is constantly measured and compensated by adding a bias current of up to 100 pA to the chamber signal.

### cross comparison

Not applicable

### arbitration

In order to overcome single event effects and the low reliability of laser transmissions, a tripling of the digital CFC electronics with voting and a double optical link is used. CRC comparisons for the transmitted data allow a increase of availability if one link is not functioning.

## Software interface usage

The system will be limited to 4 clients.  For the 1hz data (MAX DATA), a concentrator will have sole access to the FECs.  All other interested clients (including the fixed display applications and the logging clients) will pass via this concentrator to get the data.  For the XPOC data, OP will provide a client which will GET the data (no subscription) following a beam dump trigger.  For the Capture data, a dedicated client also provided by OP will have access .  Finally, the collimation data will be sent over a private UDP channel between the front end and the collimation supervisory system.

### fixed display

The fixed display will be fed data from the concentrator (i.e. no direct contact with the FEC).

### data logging server

The logging client will be fed data from the concentrator (i.e. no direct contact with the FEC).

### operational clients

There will be 4 operational clients. The generic post mortem application, an XPOC client, a client for analysing the capture data, and the fixed displays.  In addition, the 1Hz data will be made available to other clients via the logging client.

### machine protection

These issues are described in the EDMS https://edms.cern.ch/document/893660/1  “MPS aspects of the beam loss monitoring system commissioning”

# Definition of the tests

## notation

We define four commissioning scenarios:

- initial first-time commissioning “10HWC”

- cold check-out  			  “20COLD”

- commissioning with beam 	  “30BEAM”

- quick self-check			  “40QUICK”

which contain each a sequence of tests to be performed.

Each test is identified with a label, and contains a list of steps to be executed. The label contains a sequence number, a blank space (for separation), a few descriptive keywords followed by sequence information within curly brackets. An example is:

20 Interlock Check {M, 1h, T U S, HWC COLD}

- step 1

- step 2

- ...

where the label is composed of:

- “20” is the sequence number which is used for sorting and sequence purposes

- “Interlock Check” are descriptive keywords to characterize the test

- {M, 1h, T U S,  HWC COLD} is the sequence information, describing whether the test is

      - Per Monitor (M) or per FEC (F)

      - Time needed for the test:

hours(h) or minutes(m) or seconds(s) per F or M

      - Access conditions required (tunnel= T, UA, UJ, UX= U, surface= S)

      - Used in commissioning phase (HWC, COLD, BEAM, QUICK)

The above example thus defines an Interlock Check which is done for each monitor, takes 1h per monitor, needs access to the tunnel, underground and surface locations and is performed during initial hardware commissioning and at cold checkouts.

## Procedures in the MTF [in progress]

In order to clearly distinguish the tests for each procedure the tests are prefixed with a string, and become “MTF-labels”. The prefixes are for:

- initial first-time commissioning “10HWC-”

- cold check-out  			 “20COLD-”

- commissioning with beam 	 “30BEAM-”

- quick self-check			 “40QUICK-”

Like this the four procedure types appear grouped together in the MTF, where the jobs are listed in alphanumerical ascending order with respect to their MTF-labels (which are the descriptions). The procedures for initial first-time commissioning, cold check-out and commissioning with beam will be performed by the instrumentation specialist and recorded in the MTF, the procedure for quick self-check is defined in this document but is performed by the CCC sequencer in an automatic way and the recording and management is done ad-hoc by the CCC crew.

Tests which are performed on a per-FEC basis are registered in the equipment MTF of the corresponding FEC, tests which are performed per -Monitor are in the MTF of the corresponding monitor.

## Scope and Limits of the procedures

At any time it might be necessary to suspend actual routine testing or operational exploitation of an instrument, and start specialist debugging. Imposing a procedural approach to a debugging activity generally does not lead to efficient working, since debugging means “collecting knowledge about a yet unsolved and potentially unknown problem”, and a procedure “lays down how to deal with known problems” only. During debugging the relevant tests in each architectural layer will be repeated several times, until satisfactory results are obtained. If the problem is then well identified and can be well isolated within its layer, there is normally no need to repeat all of the tests which are sequentially following. Depending on the situation, some test might become mandatory nevertheless (i.e. a recompilation of the software might entail a standard interface check, and a repaired connector might entail a signal transmission check).

As a consequence of well-targeted debugging the result status of a test procedure will usually become “non sequential”: all tests are OK and passed, but they have not been performed strictly in the required order. This un-sequential testing can be easily detected by looking at the time-stamps of the test results in the corresponding MTF folders. Therefore the result of the whole procedure has less value, in the sense that the statement:  “the instrument is working properly”, has less confidence. This is a potentially dangerous situation, but a systematic re-testing of all higher-sequence steps cannot be enforced using the MTF, and might also turn out impractical. It must be left to the project leader, the instrument specialists and the management to decide on a case by case basis if the obtained confidence is sufficient for machine operation at a given point in time, or prescribe parts or all of a test-procedure. For major repairs nevertheless the full procedure should be performed so that all results are in the prescribed sequential order, giving the optimum confidence and quality of the test result.

If some of the tests are not passed for known reasons, this negative result obviously has to be recorded in the corresponding MTF folder. Within the scope of this activity this should be taken as a compliment for thorough testing, and not as a criticism for a badly performing instrument.

# tests for the BLM

The following tests will be performed during the hardware commissioning of the BLM system.

One engineer with at least one technician shall be engaged in this test program, for which the overall responsibility lies with AB/BI. Differences arise from the different distances between the detector locations and the data acquisition crates. Installations in the arcs allow a direct conversation between the person who manipulates the cabling and the person who observes the crate measurement instrument. Installations in the long straight sections and in the dump lines require telephone contact for these tasks.

The LHC beam loss monitors are foreseen to be hardware commissioned during the electronics installation phases. Independent tests of each monitor and each channel will be made by AB/BI/BLM to validate the connectivity and the topology of the system. All these tests are referenced in the MTF database.

All testing steps needed for initial first-time commissioning, cold check out and quick self-testing are listed and defined in this chapter. There are no tests with beam. The structure of this chapter follows the functionality and architecture of the instrument, so that the tests corresponding to each functional layer are grouped together.

## monitors

## tunnel equipment (tunnel and alcoves)

### monitors and cabling towards monitors

These tests have to be done per [ Bernd: is this correct ? ] yellow mini rack, and the scope covers also all monitors at the same time.

210 Tunnel Cables and Functional Test {M, 5m, T, HWC}

- Cable identification check (TS/EL sticker + AB/BI specific sticker)

- Cable continuity check; particularly verify if there are no conductors swapped

- Test of insulation, correct interconnection and current limits

- The crate shall be equipped with a test card to record the measurements.

- Measure the leakage current (offset) with the cable and the monitor connected. The measured current on all channels shall be below 0.5 pA.

- Measured Properties:

- Cable number

- Leakage current offset

211 Tests of insulation, correct interconnections and current limits {M, 5m, T, HWC}

- At the location of the detector the signal cable shall be disconnected and a current source shall be connected to each signal cable. The correct channel labelling shall be checked by observing channel by channel the current signal at the test card

- Any inversion of cables shall be corrected

- In case the measured currents are outside the limits the corresponding cable shall be changed

- In the event that the measured current is still outside the prescribed limits after a change of cable, a CERN specialist has to be informed to correct the installation

- Measured Properties:

- Current signal per channel on the test card

### electronics

The electronics is tested together with the cabling (see previous chapter).

## back-end equipment

### intermediate electronics

### cabling towards Tunnel equipment

### electronics

330 Threshold Channel and Matrix {F, 5m, -, HWC QUICK}

- before each fill

- Threshold table, channel parameters (e.g. electronics card number) and channel mapping (matrix for maskable/unmaskable and connected/unconnected) are compared to the one stored in the data base

331 Beam Permit Transmission Comparator-Combiner {F, 5m, -, HWC QUICK}

- [ Bernd: the formulation of this test is not clear. What exactly needs to be done ? ]

- Sequencer or SIS demands combiner to initiate test

- Combiner addresses the TC in sequence and records test results

- from all threshold comparators to the last combiner card

- before each fill

- for maskable and un-maskable channels

- Beam permit transmission from all threshold comparators to the last combiner card

332 Beam Permit Transmission Combiner-CIBU {F, 5m, -, HWC QUICK}

- Responsibility of BIS!

- CS sets first line A beam permit OK, line B no beam permit and then the other way round. CIBU detects the pattern.

- before each fill

- for maskable and un-maskable channels

- Beam permit transmission from each combiner card to the “controls interlock beam user” (the interface to the BIC in the VME crate).

### machine protection interface

340 Disabling of User Beam Permit {F, 5m, -, HWC}

- Tested all 16 channels, all 12 running sums and all 32 energy values of one threshold comparator, each time a new firm ware is installed – test of firmware.

- On each channels, each running sum and at each energy in the threshold comparator is brought above the abort threshold (increasing the signal and/or lowering the threshold value).

- Check if firmware correctly identifies the channel above threshold.

### Energy reception interface

350 LHC Energy Information {F, 5m, -, HWC QUICK}

- The LHC energy information is sent via the “slow timing” to the CTRV (VME card) and from there to the BLECS (which distributes it to all TC in parallel)

- The details (including the starting point) of this test is not yet defined

- The BLECS verifies continuously that it has received a) a new energy value (toggle bit) every 100ms and the CRC control word.

### vme services

The FEC is rebooted remotely and the whole instrument must become operational after the start up delay. Any self-tests must be completed successfully, and without external interventions. Verify the FEC configuration. Traces for future reference are to be recorded: Boot-log from the terminal and messages from the system.

360 Front End System Test {F, 1h, -, HWC COLD}

- Verify FEC configuration, snapshot of rc.local

- Remote-reboot the VME crate, check no errors are present, record trace

- Check central timing (use FEC timing diagnostics, record results)

- Check BST and external conditions

- Measured Properties:

- Rc.local/transfer.ref

- Boot trace of each FEC ( /mcr/reset/rem_reset <fec> | tee <fec>.boot )

- Timing trace dump

361 Front End Software {F, 1h, -, HWC COLD}

- Check that all relevant software services are running stable, record a trace (“ps axt”)

- Check that acquisition is alive by reading 10 pA data.

- Check FEC timing and synchronisation behaviour: measure RT execution times using Fesa 2.10 tools

- Check the physical installation of the DAB cards using the ‘status’ fields in the Acquisition property.  It is possible to have a bad installation of a card, resulting in the data being corrupted as it is transferred from the CPU (normally an address line problem).  Re-inserting the card can often solve this problem

- Where possible, check the capture, XPOC and post mortem data triggering by send the appropriate timing events

- Measured Properties:

- Activity trace of each FEC ( ps axt | tee <fec>.activity )

- RT activity trace

## Software public interfaces

All public SW interfaces are tested from the application software layer, in a standard operation-like manner (no stress test). The instrument must react consistently and repeatable, without requiring external intervention.

400 Public Interface {M, 1h, -, HWC COLD}

- Check basic functionality (GET/SET) of all public interfaces using the FESA Navigator or ROSALI. Check that values are within their legal range.

- check that data rates as specified are kept under normal operational conditions for all public interfaces

## software dedicated interfaces

### Post mortem

After the reception of a post-mortem trigger the PM-data buffer of the public interface has to be available, frozen and consistent. A post mortem dump has to be recorded for future reference.

510 Post Mortem {M, 1h, -, HWC COLD}

- Generate a testing post-mortem trigger, either manually or remotely.

- Check the post-mortem acquisition (BLM expert application)

- Check the post mortem acquisition (PM system)

- Logging continues without interruption during PM data buffer readout

### Data Logging

520 Data logging {M, 1h, -, HWC COLD}

- Together with DB specialist, check that data logging is configured correctly

- Check database logged values with a running instrument (can be simulated values)

- Logging continues without interruption

- Invoke a capture data and XPOC data request and monitor the effect on the logged data

### XPOC

530 XPOC {M, 1h, -, HWC COLD}

- Generate a testing XPOC trigger, either manually or remotely.

- Check the XPOC acquisition (BLM expert application)

- Check the XPOC acquisition (XPOC system)

- Logging continues without interruption during XPOC readout (?)

### Study data

540 Study Data {M, 1h, -, HWC COLD}

- Generate a testing “study data” trigger, either manually or remotely.

- Check the study data acquisition (BLM expert application)

- Check the study data acquisition (OP application)

- Logging continues without interruption during study data buffer readout

- Test behaviour of clients if 2 conflicting capture data configurations are sent

### collimation data

550 Collimation Data {M, 1h, -, HWC COLD}

- Generate a testing collimation trigger, via the CSS.

- Check the collimation data that was sent over UDP (specific property in the API used for debugging only)

- Check the collimation data acquisition (collimation system)

- Logging continues without interruption during collimation data readout

- Test that the collimation data is sent during a capture data request and a post mortem data request

## Software runtime

### synchronisation

610 MTG machine timing {F, 1h, -, HWC COLD}

- Central machine timing up and running

- [ Stephens: can you be more specific please ]

- LEDs channel xxx flashing

- Hardware timings correctly coming out and values OK

- SW interrupts arriving

### real-time performance

620 FEC runtime load {F, 1h, -, HWC COLD}

- Under operating conditions, using several clients:

- Check that the average CPU load is below 0.5

- Check that memory reserves are sufficient

- Check run-time stability of the whole FEC over several days (timescale larger than 1 complete LHC filling + collisions)

- Monitor the status of the 16 DAB cards.  They should all read similar status values.

### settings and persistency

Once the calibration is done and all relevant settings are determined, they are entered into the system, using the public software interfaces. These settings must not be lost in case of a system failure or reboot, and they are needed for future reference and comparison. Therefore a backup of the initial instantiation data for the whole instrument is taken. A persistent data snapshot should be stored as well, where appropriate.

630 System Settings and Persistency {F, 1h, -, HWC COLD}

- Enter all HW and SW specialist settings.

- Make a backup of instance and persistency data

- Reboot the systems remotely

- Measured Properties:

- Persistency file dump for the device

- Instantiation file dump for the device

### error reporting

An out-of-nominal condition will be reported by the instrument on at least WARNING level. Generally, most failures occur during initialisation or out-of-nominal conditions, and a stable and precise reporting is essential for quick problem solution. Also it is vital that the instrument itself is robust and stays “alive” even in abnormal conditions, so that at least the reporting will continue to work. It is not at all trivial to consistently test error behaviour, because this would imply producing all possible errors beforehand. Nevertheless a few basic hardware errors can be produced easily and should be reported by the software: disconnecting and reconnecting certain signals on the VME electronics should be sufficient.

640 Error Conditions {M, 1h, -, HWC}

- Check Error reporting:  disconnect or simulate out-of-nominal acquisitions and check error-reporting and stability. (Use bad specialist settings to produce errors: The point is to see errors reported. The optimum solution would be to put the instrument SW into a special state where it spits out all errors it has defined in a sequence)

- Check instrument software survey tools / LIDS

- Reboot and recover reference settings (restore backup file)

## complete instrument

### Full acquisition chain

710 High Voltage Modulation Test {F, 1m, -, HWC QUICK}

- Check the connectivity up to the surface card before every fill

- Modulate the high tension with a sine wave

- Acquire modulated monitor signal, and check

- Measured Properties:

- modulated monitor signal

711 Acquisition Chain Test Via Test Signal {F, 1m, -, HWC QUICK}

- P1:100pA/84 sec run.sum [ Bernd: formulation is unclear ]

- Performed in sequence with the HV modulation test

- The bias current is increased by 100 pA and the 84 second running sum is measured and checked. Bias current reduced to previous value.

- Measured Properties:

- 84 second running sum

This test is run continuously in the CCC to monitor the system.

712 Monitoring 10 pA {F, 1m, -, HWC QUICK}

- [ Bernd: formulation is a bit unclear ] (was MTF test 30-BLM, P2:10pA/84 sec run.sum)

- continuous check of the tunnel electronics : each minute the integrated injected current offset is acquired through the standard acquisition chain and the 84 second running sum is verified

- The bias current is regulated in a control loop (possibility to correct for eventual offset in the electronics – e.g. ageing effects) so that a 10 pA offset is measured.

- the combiner BLECS verifies signal levels

- Measured Properties:

- Bias current

### cross comparison, indirect calibration

720 Functional Test With Radioactive Source {M, 5m, T, HWC}

- The procedure for this last test will be described in a dedicated document made in collaboration with TIS. The purpose is to create a signal on the chamber with the RA source and check its presence in the corresponding DAB card channels

- the event that the measured parameters are outside the prescribed limits, a CERN specialist shall be informed to correct the installation

- Create a signal on the ionisation chamber with the RA source

- Check the presence of signal in the corresponding DAB card

- Measured Properties:

- RA signals

### emc test

730 EMC Test {F, 2h, -, HWC}

- The 40 us and 1.3 s running sum are measured with possible interference sources switched ON and OFF

- Possible interference sources are:

- Collimator jaw motors

- Kicker magnets

- …

### Machine protection: provoked quench without beam

[ Bernd: I added this from the document MPS-BLM_20070930 . check ]

740 Provoked Quench by BLM IR6 {F, 1d, -, BEAM}

- Test machine protection system functionality of the BLM to provoke a non-damaging quench by manipulating a BLM  (AB-CO)

- The test has to be done for one beam only

- Test the direct interlocked BLM in IR6

### test with beam: Pilot of 5(109 p+ at 450 Gev

[ Bernd: I added this from the document MPS-BLM_20070930 . check ]

750 MPS functionality of the BLM NO QUENCH {F, 1d, -, BEAM}

- Test machine protection system functionality of the BLM (without quenching) (AB-CO)

- Decrease thresholds to very low value (AB-BI)

- Create local bump to force a beam dump request with very low level of losses (AB-BT)

- Measure delay between signal over threshold and the beam dump : the whole chain from the chamber to the LBDS, via the BIC, is tested

### test with beam: 43 bunches of 4(1010 p+ at 450 Gev (below damage value)

[ Bernd: I added this from the document MPS-BLM_20070930 . check ]

These two tests are dangerous, and there is a certain possibility that the corresponding beam losses occur at certain moments during LHC commissioning, without being provoked explicitly. If the BLM and MPS perform as expected, the corresponding tests is considered OK. Nevertheless if corresponding beam losses do not occur or cannot be consistently detected, it might be necessary to carefully perform these tests within the context of machine protection commissioning.

760 Provoked Quench with Transient Losses {F, 1d, -, BEAM}

- Test machine protection system functionality of the BLM to provoke a non-damaging quench (transient losses)  (AB-CO)

- The test has to be done for one beam only

- Steer the beam in the chosen magnet (AB-BI)

- Check level of losses and BLM readings at the quench level

- Repeat for MQM, MQTL, MQ and MB

761 Provoked Quench with Steady State Losses {F, 1d, -, BEAM}

- Test machine protection system functionality of the BLM to provoke a non-damaging quench (steady state losses)  (AB-CO)

- The test has to be done for one beam only

- Create a local bump

- Check level of losses and BLM readings at the quench level

## Software interface usage conformity

They will follow the general LHC commissioning schedule, and sometimes can be done on an ad-hoc basis parasitically. The objective is to ensure the correct usage of our SW interfaces in order to fully exploit the functionality ensuring the required stability. The software interfaces have to be used correctly, specifically only public interfaces are used in the CCC, RBAC restrictions are defined applied correctly, subscriptions are handled as specified in LIDS.

These tests will require a small amount of parasitic or even dedicated MD-time, since the instrument might be unreliable for other users during testing.

### fixed display

### data logging server

### operational clients

The BLM concentrator serves the clients: XPOC, study data, post mortem, collimator server, CCC fixed display, data logging

830 Concentrator Server {M, 2h, -, HWC COLD BEAM}

- Based on a list of main clients provided by the CCC:

- Check that all clients automatically interpret the arbitration flag correctly

- Check the update frequency for each client

- Data consistency for each client

- Error reporting for each client

### XPOC

### machine protection

### CCC Sequencer

860 CCC Sequencer Integration {M, 2h, -, HWC COLD}

- Check that the sequencer correctly implements the BLM QUICK test

# ProcedureS for BLM commissioning

The procedures list the sequence of steps, which are defined in the previous chapter, the prerequisites, the access conditions, and the estimated test duration under normal conditions, for

- Initial first-time commissioning

- Cold check-out procedure

- Commissioning with beam

- Quick Self-Test

However, due to the complexity of some situations and their unpredictability, the given procedures might still be incomplete, and will be subject to updating and changes following the experiences.

To improve readability an “(M)” (for MTF per monitor) or “(F)” (for MTF per FEC) is added to the column “test duration” in the procedure tables.

The MTF equipment folder for the monitors can be found at:

https://edms.cern.ch/asbuilt/plsql/mtf_urep.start_urep?p_ure_code=48

->“Beam Instrumentation”->”BI Beam loss Monitor”->

(continue)->(select monitor)->Tab “Installation and Commissioning”

The MTF equipment folder for the FECs can be found at:

https://edms.cern.ch/asbuilt/plsql/mtf_urep.start_urep?p_ure_code=48

->“Beam Instrumentation”->”BI VME”->(continue)->”BVCRA.SX4.BCTDC<X>” [ MTF seems down ]

->Tab “Installation and Commissioning”

## procedure for initial first-time commissioning

The initial first-time commissioning procedure is a complete test including hardware, cables, electronics, functionality, calibration, machine protection, data transmission and interpretation. This test will be required after major installations or upgrades.

The LHC Beam Loss Monitors are foreseen to be hardware commissioned during the  installation phase.

### Entry conditions

Before commencing hardware commissioning the following infrastructure has to be in place and, where applicable, operational:

- VME chassis installed and tested. This is under the responsibility of AB-BI (ref. LHC-B-TP-0001.00).

- VME modules such as CPU, BLETC, BLECS, CTRV and BOBR have to be installed

- 220V and Ethernet operational

[ Bernd: please confirm that all these tests for the HWC are OK ]

Table11.1: Initial first-time commissioning procedure

| MTF-label | Tunnel | UA/UJ | Surface | Total duration for all monitors, FECs [h] | Extra Prerequisite |
| --- | --- | --- | --- | --- | --- |
| 10HWC-210 Tunnel Cables and Functional Test | X | - | - | 2weeks(1) |  |
| 10HWC-211 Tests of insulation, correct interconnections and current limits | X | - | - | 2weeks(1) |  |
| 330 Threshold Channel and Matrix | - | - | - | 5m(2) |  |
| 331 Beam Permit Transmission Comparator-Combiner | - | - | - | 5m(2) |  |
| 332 Beam Permit Transmission Combiner-CIBU | - | - | - | 5m(2) |  |
| 340 Disabling of User Beam Permit | - | - | - | 2h | New firmware |
| 350 LHC Energy Information | - | - | - | 5m | Telegram information ? |
| 360 Front End System Test | - | - | - | 3h(3) |  |
| 361 Front End Software | - | - | - | 3h(3) |  |
| 400 Public Interface | - | - | - | 3h(3) |  |
| 510 Post Mortem | - | - | - | 3h(3) |  |
| 520 Data logging | - | - | - | 3h(3) |  |
| 530 XPOC | - | - | - | 3h(3) |  |
| 540 Study Data | - | - | - | 3h(3) |  |
| 550 Collimation Data | - | - | - | 3h(3) |  |
| 610 MTG machine timing | - | - | - | 3h(3) |  |
| 620 FEC runtime load | - | - | - | 3h(3) |  |
| 630 System Settings and Persistency | - | - | - | 3h(3) |  |
| 640 Error Conditions | - | - | - | 3h |  |
| 710 High Voltage Modulation Test | - | - | - | 10m(3) |  |
| 711 Acquisition Chain Test via test signal | - | - | - | 10m(3) |  |
| 712 Monitoring 10 pA | - | - | - | 10m(3) |  |
| 720 Functional Test With Radioactive Source | X | - | - | 3weeks |  |
| 730 EMC Test | - | - | - | 3h |  |
| 830 Concentrator Server | - | - | - | 2h |  |
| 860 CCC Sequencer Integration | - | - | - | 2h |  |

- these tests are grouped together and are performed per geographic location

- these tests are grouped together and are performed on a global level for the whole instrument

- these tests will be executed with the help of specific expert programs

### The status after the first-time commissioning

After the successful installation and hardware commissioning, the status of the BLM will be as follows:

The hardware is fully installed

The cables are connected to the monitors.

The tunnel electronics cards are working under test conditions

The surface electronics cards are working under test conditions

The digital acquisition electronics are operational

The software runs

The BLMs thresholds are verified

The BLM signals can be read as noise in the CCC

The status of the BLM after this test is “operational for the first time”.

## procedure for Cold checkout

The cold checkout procedure includes testing of the whole instrument on all architectural layers: monitor, front-end electronics, cables, VME-electronics, timing, data treatment, publishing, network transmission and machine protection interface.  This test will be required after each machine cool-down, and at least once after each shutdown. This test is done without beam.

If the “initial first-time commissioning” procedure has been successfully passed before within a reasonable time-window, the corresponding tests which appear in both procedures may be skipped.

### entry conditions

Table11.2: Cold-Checkout Commissioning procedure

| MTF-label | Tunnel | UA/UJ | Surface | Total duration for all monitors, FECs [h] | Extra Prerequisite |
| --- | --- | --- | --- | --- | --- |
| 360 Front End System Test | - | - | - | 3h(3) |  |
| 361 Front End Software | - | - | - | 3h(3) |  |
| 400 Public Interface | - | - | - | 3h(3) |  |
| 510 Post Mortem | - | - | - | 3h(3) |  |
| 520 Data logging | - | - | - | 3h(3) |  |
| 530 XPOC | - | - | - | 3h(3) |  |
| 540 Study Data | - | - | - | 3h(3) |  |
| 550 Collimation Data | - | - | - | 3h(3) |  |
| 610 MTG machine timing | - | - | - | 3h(3) |  |
| 620 FEC runtime load | - | - | - | 3h(3) |  |
| 630 System Settings and Persistency | - | - | - | 3h(3) |  |
| 830 Concentrator Server | - | - | - | 2h |  |

### The status after the COLD check-out tests

After the successful installation and hardware commissioning, the status of the BLM will be as follows:

The tunnel electronics cards are working under test conditions

The surface electronics cards are working under test conditions

The digital acquisition electronics are operational

The software runs

The BLMs thresholds are verified

The BLM signals can be read as noise in the CCC

The data delivered to the machine protection system is correct

The status of the BLM after this test is “operational after cool-down”.

## Commissioning with Beam

The commissioning of the BLM with beam has the aim to verify the system under real-world conditions, and to verify the correctness of the integration for machine operations. Initial comparisons and cross-calibrations are done, in order to gain confidence in the instrument. The quantification of the precision of the intensity reading for all nominal operation scenarios, to identify performance limiting factors, is described in a separate document.

### entry conditions required

Beam of sufficient intensity must circulate in the LHC, in the ring corresponding to the monitor.

Table11.3: Commissioning with Beam Procedure

| MTF-label | Tunnel | UA/UJ | Surface | Total duration for all monitors, FECs [h] | Extra Prerequisite |
| --- | --- | --- | --- | --- | --- |
| 740 Provoked Quench by BLM IR6 | - | - | - | 1d(1) | MPS must be commissioned |
| 750 MPS functionality of the BLM NO QUENCH | - | - | - | 1d(1) |  |
| 760 Provoked Quench with Transient Losses | - | - | - | 1d(1) | Dangerous tests: written agreement from LHC project leader |
| 761 Provoked Quench with Steady State Losses | - | - | - | 1d(1) | Dangerous tests: written agreement from LHC project leader |

### The status after the commissioning with beam

The [ Bernd: please write this paragraph]	Further “performance assessment procedures” will structure the assessment of the measurement quality within the framework of a post-commissioning activity, which will be lead and coordinated by BI for this instrument.

## procedure for quick self-test

The self-test procedure includes testing of calibration, machine protection and data transmission. This test will be required before the LHC injects beam above the destruction threshold. The aim is to perform this test automatically as a part of a self-test sequence, which is initiated by the CCC. This procedure is defined here, but it will not be recorded by the MTF, but instead be automatically treated by the CCC-sequencer and the CCC-crew on an ad-hoc basis (no MTF-labels present).

### Entry conditions

The CCC-sequencer must be fully functional and integrated with this instrument, and there must be no beam. This check can be triggered before each first injection, and requires no manual intervention.

Table11.4: Quick Self-Test procedure

| Name | Duration per monitor [seconds] | Pre Requisite |
| --- | --- | --- |
| 330 Threshold Channel and Matrix | 5m(1) | CCC sequencer fully integrated |
| 331 Beam Permit Transmission Comparator-Combiner | 5m(1) | CCC sequencer fully integrated |
| 332 Beam Permit Transmission Combiner-CIBU | 5m(1) | CCC sequencer fully integrated |
| 350 LHC Energy Information | 5m(1) | CCC sequencer fully integrated |
| 710 High Voltage Modulation Test | 1m | Expert application |
| 711 Acquisition Chain Test Via Test Signal | 1m | Expert application |
| 712 Monitoring 10 pA | 1m | Expert application |

- All FECs are concerned for a global test

### The status after the Quick self-test

The self-test status will be recorded ad-hoc by the CCC. After successful completion of this procedure is:

- BLM is working

- BLM interface to Machine protection is working

- BLM monitoring is working and shows OK

The status of the BLM is “ready for injection of beam”.

# Footnotes

http://bdidev1.cern.ch/bdisoft/development/BDI-Domains/bdeyelids/bdeyelids.php?currentSelection=DE&currentDomain=LHC&currentInstrument=General