## HCBLM_I001
Beam Loss Monitor: Ionization Chamber
Identification |  |  |  |  | Step 1 (): Assembly | [6 empty] | Step 1: Non-Conformities |  | Properties |  |  | Step 2 (): Vacuum stand | [6 empty] | Step 2: Non-Conformities |  | Properties | [7 empty] | Step 3 (): Leakage current measurement at IHEP by IHEP | [6 empty] | Step 3: Non-Conformities |  | Properties |  |  | Step 4 (): Leakage current measurement at IHEP by CERN | [6 empty] | Step 4: Non-Conformities |  | Properties |  | Step 5 (): Shipping | [6 empty] | Step 5: Non-Conformities |  | Properties |  |  | Step 6 (): Leakage current measurement at CERN | [6 empty] | Step 6: Non-Conformities |  | Properties |  | Step 7 (): Source test | [6 empty] | Step 7: Non-Conformities |  | Properties | [13 empty] | Step 8 (): Source test at tunnel | [6 empty] | Step 8: Non-Conformities |  | Properties |  |  | Equipment Non-Conformities |  | Additional document 1 | Additional document 2 | Additional document 3
Equipment ID | Other ID | Manufacturer | Status | Comment | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Leakage Current (Assemb) | Head Producer | High Voltage (Assembly) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Batch Number | Gas Bottle Number | Spectrum (before filling) | Temperature (filling) | Pressure (filling) | Temperature (heating) | Duration (heating) | Spectrum (after filling) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Leakage Current_f | Leakage Current_v | High Voltage (Leak-I-I) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Leakage Current (IHEP) | High Voltage (Leak-I-C) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Lorry Number | IHEP Departure Date | CERN Arrival Date | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | High Voltage (Leak-C-C) | Leakage Current (CERN) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | High Voltage1 | Current1 | High Voltage2 | Current2 | High Voltage3 | Current3 | High Voltage4 | Current4 | High Voltage5 | Current5 | High Voltage6 | Current6 | Dose Rate | Source Activity (Surface) | Status | Result | Standard Step ID | Completed On | Executed by | Results doc. | Comments | Doc. Number | Type | Current | High Voltage (Tunnel) | Source Activity (Tunnel) | Doc. Number | Type | Doc. Number | Doc. Number | Doc. Number
[14 empty] | [pA] | [] | [V] | [9 empty] | [] | [] | [] | [°C] | [mbar] | [°C] | [h] | [] | [9 empty] | [pA] | [pA] | [V] | [9 empty] | [pA] | [V] | [9 empty] | [] | [] | [] | [9 empty] | [V] | [pA] | [9 empty] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [V] | [pA] | [millisievert par heure] | [GBq] | [9 empty] | [pA] | [V] | [GBq]
(CERN ID) | [42 empty] | <1 pA
BL01 | HCBLM_I001 | Beam Loss Monitor: Ionization Chamber |  |  | MTF013476 | [7 empty] | 3 | P11483 | P11446 | P11448 | MTF013477 | [7 empty] | 8 | P11450 | P11468 | P11453 | P11451 | P11452 | P11454 | P11456 | P11455 | MTF013478 | [7 empty] | 3 | P11458 | P11457 | P11459 | MTF013479 | [7 empty] | 2 | P11449 | P11460 | MTF013480 | [7 empty] | 3 | P11461 | P11462 | P11463 | MTF013481 | [7 empty] | 2 | P11464 | P11465 | MTF013482 | [7 empty] | 14 | P11471 | P11477 | P11472 | P11478 | P11473 | P11479 | P11474 | P11480 | P11475 | P11481 | P11476 | P11482 | P11467 | P11469 | MTF013483 | [7 empty] | 3 | P11470 | P11466 | P11447
| 0 | 8 |  | 3
HCBLM_I001-05000008 | IHEP0008 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.4 | 0.1 | 2000 | Done | Ok | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.8 | 1800 | Done | Ok | 5 | 2006-09-28 | Grishin |  |  |  |  | 1 | 2006-06-16 | 2006-06-22 | Done | Ok | 6 | 2006-06-23 | G.Ferioli |  |  |  |  | 2000 | 0.1 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.3 | [10 empty] | 4.7 | 98 | Pending |  | 8
HCBLM_I001-05000009 | IHEP0009 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.7 | 0.9 | 2000 | Done | Ok | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.7 | 1800 | Done | Ok | 5 | 2006-09-28 | Grishin | [7 empty] | Done | Ok | 6 | 2006-06-23 | G.Ferioli |  |  |  |  | 2000 | 0.7 | Done | Ok | 7 | 2006-06-27 | G.Ferioli |  |  |  |  | 1500 | 54.2 | [10 empty] | 4.7 | 98 | Pending |  | 8
HCBLM_I001-05000010 | IHEP0010 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 150 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000011 | IHEP0011 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.2 | 0.9 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 1 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000012 | IHEP0012 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000013 | IHEP0013 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.8 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000014 | IHEP0014 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 2.5 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.48 | 1 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000015 | IHEP0015 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 3 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.1 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.7 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000016 | IHEP0016 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.4 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 11 | 1800 | [9 empty] | 2 | 25.07.2006
HCBLM_I001-05000017 | IHEP0017 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000018 | IHEP0018 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 4 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.24 | 0.1 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000019 | IHEP0019 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 5 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Mishagin |  |  |  |  | 0.4 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000020 | IHEP0020 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-24 | Mishagin |  |  |  |  | 7 | IHEP | 2000 | Done | Ok | 2 | 2006-04-25 | Larionov |  |  |  |  | 2 | 1 | lar002_1;lar002_2 | 13.5 | 1074 | 240 | 6 | lar002_3;lar002_4 | Done | Ok | 3 | 2006-04-25 | Tikhonov |  |  |  |  | 0.24 | 0.2 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000021 | IHEP0021 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.2 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.7 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000022 | IHEP0022 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 2.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 1.4 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000023 | IHEP0023 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.4 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000024 | IHEP0024 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 0.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000025 | IHEP0025 | IHEP | Manufacturing |  | Done | Not Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 2.5 | IHEP | 2000 | Cancelled | Cancelled | 2 | [14 empty] | Cancelled | Not Ok | 3 | [11 empty] | 4 | 2006-06-01 | G.Ferioli
HCBLM_I001-05000026 | IHEP0026 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.5 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 1 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000027 | IHEP0027 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Cancelled | Not Ok | 3 | [11 empty] | 4 | 2006-06-01 | G.Ferioli
HCBLM_I001-05000028 | IHEP0028 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 4 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.3 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000029 | IHEP0029 | IHEP | Manufacturing |  | Done | Not Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.5 | IHEP | 2000 | Cancelled | Not Ok | 2 | [14 empty] | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.73 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.3 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000030 | IHEP0030 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.73 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli | [5 empty] | 1800
HCBLM_I001-05000031 | IHEP0031 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.7 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.3 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000032 | IHEP0032 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.8 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000033 | IHEP0033 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.3 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000034 | IHEP0034 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.9 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | -1.3 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000035 | IHEP0035 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000036 | IHEP0036 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-25 | Seleznev |  |  |  |  | 1 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.7 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | -0.3 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000037 | IHEP0037 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-27 | Mishagin |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.73 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 1 | 1800 | [9 empty] | 2 | 25.07.2006
HCBLM_I001-05000038 | IHEP0038 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.7 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli | [5 empty] | 1800
HCBLM_I001-05000039 | IHEP0039 | IHEP | Manufacturing |  | Done | Not Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.4 | IHEP | 2000 | Cancelled | Cancelled | 2 | [14 empty] | Cancelled | Not Ok | 3 | [11 empty] | 4 | 2006-06-01 | G.Ferioli
HCBLM_I001-05000040 | IHEP0040 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.3 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000041 | IHEP0041 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-27 | Mishagin |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.73 | 0.7 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.5 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000042 | IHEP0042 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.6 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000043 | IHEP0043 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli | [5 empty] | 1800 | [9 empty] | 2 | 25.07.2006
HCBLM_I001-05000044 | IHEP0044 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.1 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.49 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000045 | IHEP0045 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.4 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.6 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000046 | IHEP0046 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-26 | Mishagin |  |  |  |  | 0.2 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.9 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000047 | IHEP0047 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-04-27 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Done | Ok | 2 | 2006-05-03 | Larionov |  |  |  |  | 3 | 1 | lar003_1;lar003_2 | 18-20 | 1095 | 240 | 6 | lar003_3;lar003_4 | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.73 | 0.74 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.6 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000048 | IHEP0048 | IHEP | Manufacturing |  | Done | Not Ok | 1 | 2006-04-27 | Tikhonov |  |  |  |  | 0.6 | IHEP | 2000 | Cancelled | Cancelled | 2 | [14 empty] | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.48 | 0.74 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.6 | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000049 | IHEP0049 | IHEP | Manufacturing |  | Done | Not Ok | 1 | 2006-04-27 | Tikhonov |  |  |  |  | 0.4 | IHEP | 2000 | Cancelled | Cancelled | 2 | [14 empty] | Done | Ok | 3 | 2006-05-03 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | 1800
HCBLM_I001-05000050 | IHEP0050 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.9 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar002_8 | 19-21 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-18 | Tikhonov |  |  |  |  | 0.24 | 0.48 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli | [5 empty] | 1800 | [9 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000051 | IHEP0051 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar002_8 | 19-21 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-18 | Tikhonov |  |  |  |  | 0.24 | 0.24 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli |  |  |  |  | 0.4 | [10 empty] | 1 | 2006-06-16 | 2006-06-22
HCBLM_I001-05000052 | IHEP0052 | IHEP | Manufacturing |  | Done | Ok | 1 | 2006-05-12 | Tikhonov |  |  |  |  | 0.24 | IHEP | 2000 | Done | Ok | 2 | 2006-05-18 | Larionov |  |  |  |  | 4 | 1 | lar004_7;lar002_8 | 19-21 | 1096 | 240 | 6 | lar004_11;lar004_12 | Done | Ok | 3 | 2006-05-18 | Tikhonov |  |  |  |  | 0.73 | 0.97 | 2000 |  |  | 4 | 2006-06-01 | G.Ferioli | [15 empty] | 1 | 2006-06-16 | 2006-06-22

## 2.Standard Steps
Standard Steps
Beam Loss Monitor: Ionization Chamber
Standard Step ID | Other identifier | Sequence number | Description | Provided by | Expected by | Responsible | Aplicable standard | Chapter | Comment
MTF013476 |  | 1 | Assembly | IHEP | CERN | Asanov,Seleznev |  |  | Head's welding procedure; Cleaning;Assembly;Current measurement
MTF013477 |  | 2 | Vacuum stand | IHEP | CERN | Larionov |  |  | Vacuum procedure; 36 IC per batch
MTF013478 |  | 3 | Leakage current measurement at IHEP by IHEP | IHEP | CERN | Seleznev |  |  | CERN Picometer noice +/- 0.3 pA
MTF013479 |  | 4 | Leakage current measurement at IHEP by CERN | CERN | CERN | Ferioli |  |  | CERN Picometer noice +/- 0.3 pA
MTF013480 |  | 5 | Shipping | IHEP,CERN | CERN | Seleznev,Grishin |  |  | 360 IC per month
MTF013481 |  | 6 | Leakage current measurement at CERN | CERN,IHEP | CERN | Ferioli,Sleptsov |  |  | CERN Picometer noice +/- 0.3 pA
MTF013482 |  | 7 | Source test | CERN,IHEP | CERN | Ferioli,Sleptsov |  |  | For 6 HV values measure 6 Current values
MTF013483 |  | 8 | Source test at tunnel | CERN,IHEP | CERN | Ferioli,Sleptsov |  |  | archive:measure the parameters of IC time to time

## 3.Properties
Properties
Description | Code | Unit | Type | Category | Default | Comment
Leakage Current (Assemb) | P11483 | pA | NUM
Head Producer | P11446 |  | CHAR |  |  | DG Technology(Italy),Aerovac(UK),Alca(Italy),IHEP(Russia)
High Voltage (Assembly) | P11448 | V | NUM
Leakage Current (IHEP) | P11449 | pA | NUM
Batch Number | P11450 |  | NUM |  |  | Number of batch on the vacuum stand
Gas Bottle Number | P11468 |  | NUM
Spectrum (before filling) | P11453 |  | CHAR
Temperature (filling) | P11451 | °C | NUM
Pressure (filling) | P11452 | mbar | NUM
Temperature (heating) | P11454 | °C | NUM
Duration (heating) | P11456 | h | NUM
Spectrum (after filling) | P11455 |  | CHAR
Leakage Current_v | P11457 | pA | NUM |  |  | Measurement after vacuum stand
Leakage Current_f | P11458 | pA | CHAR |  |  | Measurement after el. box soldering
Lorry Number | P11461 |  | NUM |  |  | Number of lorry from Protvino
IHEP Departure Date | P11462 |  | CHAR
CERN Arrival Date | P11463 |  | CHAR
High Voltage1 | P11471 | V | NUM
Current1 | P11477 | pA | NUM
High Voltage2 | P11472 | V | NUM
Current2 | P11478 | pA | NUM
High Voltage3 | P11473 | V | NUM
Current3 | P11479 | pA | NUM
High Voltage4 | P11474 | V | NUM
Current4 | P11480 | pA | NUM
High Voltage5 | P11475 | V | NUM
Current5 | P11481 | pA | NUM
High Voltage6 | P11476 | V | NUM
Current6 | P11482 | pA | NUM
Current | P11470 | pA | NUM |  |  | Rad source current measurements
Dose Rate | P11467 | millisievert par heure | NUM
Source Activity (Surface) | P11469 | GBq | NUM
High Voltage (Leak-I-I) | P11459 | V | NUM
High Voltage (Leak-C-C) | P11464 | V | NUM
High Voltage (Leak-I-C) | P11460 | V | NUM
High Voltage (Tunnel) | P11466 | V | NUM
Leakage Current (CERN) | P11465 | pA | NUM
Source Activity (Tunnel) | P11447 | GBq | NUM

## Personal Information
Personal Information
EDMSlogin | JJGRAS
EDMSGroup | AB-BI-BLM-MTF
Created On | 2005-01-01