| X-ray photoelectron spectroscopy (XPS) report |
| --- |
| Client: B. Holzer | Tel.: |
| Job No.: 1009-XPS-TS-MME-SC-2006-05-03 |
| Job opened by: A. Reginelli | Date: 2006-05-03 |
| Number of samples: 16 Sample codification: 1009 Inox samples, 1009-a: Inox without Al protection #1 1009-b: Inox without Al protection #2 1009-c: Inox with Al protection #1 1009-d: Inox with Al protection #2 Aluminium samples, 1009-e: Al without Al protection #1 1009-f: Al without Al protection #2 1009-g: Al with Al protection #1 1009-h: Al with Al protection #2 Copper samples, 1009-i: Cu without Al protection #1 1009-j: Cu without Al protection #2 1009-k: Cu with Al protection #1 1009-l: Cu with Al protection #2 Nature of samples: Stainless steel, Aluminium and Cu samples; cleaned at IHEP Prodvino. Two types of samples: packed in plastic bag with and without Al protection. Sample storage: air Aim of analysis: Check the surface sample cleanness and find out if the contamination is due to the packaging. Results Stainless steel samples: Figure 1 presents the so-called surface concentrations as obtained by XPS for the stainless steel samples. The samples without the Al foil protection are quite dirty; the cleanness of these surfaces does not satisfies the criterion fixed for UHV applications (C < 40 at. %). Si, Ca contamination are present on all the samples. For the two samples packaged in Al foil, the situation is different. The first one has low C atomic concentration (25.9 %), but the second one is highly contaminated. Si and Ca contamination are again present on all the samples, possibly due to an insufficient rinsing. The conclusion is that the quality of cleaning is not reproducible. Aluminium: Figure 2 presents the so-called surface concentrations as obtained by XPS for the aluminium samples The two samples without the Al foil do not satisfy the criterion fixed for UHV, (we recall that acceptance level for UHV insertion being 29 at % C for Al), while the two with the aluminium foil are acceptable. Si is present again at especially high amount (more than 10%). Copper samples: Figure 3 shows the so-called surface concentrations as obtained by XPS for the copper samples. All the samples are completely contaminated (C atomic concentration > 90%). The contamination overtakes the limit for UHV application is fixed at 44 at. %. Packaging them in aluminium foil does not change anything. By visual inspection it is also possible to see various halos on the samples. Conclusion Packaging the sample in Aluminium foils can help a bit, but the main problem is in the cleaning procedure, which is insufficient to remove the contamination and in the rinsing. Figure 1: So-called relative atomic surface concentrations calculated from the photoelectron peak intensities in the XPS spectra – Stainless steel sample; Figure 2: So-called relative atomic surface concentrations calculated from the photoelectron peak intensities in the XPS spectra – Aluminum samples; Figure 3: So-called relative atomic surface concentrations calculated from the photoelectron peak intensities in the XPS spectra – Copper sample |