| X-ray photoelectron spectroscopy (XPS) report |
| --- |
| Client: Berndt Dehning | Tel.: |
| Job No.: 1058-XPS-TS-MME-SC-2006-09-12 |
| Job opened by: A. Reginelli | Date: 2006-09-12 |
| Number of samples: 10 Sample codification: 1058 1058-a: Inox Plate cleaned at IHEP Protvino #1 (23/07/2006) 1058-b: Inox Plate cleaned at Protvino #2 (23/07/2006) 1058-c: Inox Plate cleaned at Protvino #3 (23/07/2006) 1058-d: Inox Plate cleaned at Protvino #4 (23/07/2006) 1058-e: Inox Plate cleaned at Protvino #5 (23/07/2006) 1058-f: Al Plate cleaned at Protvino #1 (23/07/2006) 1058-g: Al Plate cleaned at Protvino #2 (23/07/2006) 1058-h: Al Plate cleaned at Protvino #3 (23/07/2006) 1058-i: Al Plate cleaned at Protvino #4 (23/07/2006) 1058-j: Al Plate cleaned at Protvino #5 (23/07/2006) Nature of samples: Stainless steel and Aluminium plates cleaned at Protvino. Sample storage: air Aim of analysis: Verify the cleanness. Results The results obtained from the X-rays photoelectron spectroscopy analysis are summarized in fig. 1 and 2. The Stainless steel samples (fig. 1) are clean; the C atomic concentration is between 14 and 17.5, well below the acceptance limit for UHV utilization (40 at. %). A quite big amount of Si is detected ~ 10 at. % and also traces of Mo and Ca. A possible origin of silicon might be the residue of silicates from the cleaning process. For the Al sample the situation is similar (fig. 2); C amount is between 14.2 and 21.1 at. %, well below the acceptance limit (29 at.%). Si is still present in a quite big quantity (~14 at. %) and traces of Ca are detected. Conclusions: All the samples are sufficiently clean to be used in UHV. Figure 1: So-called relative atomic surface concentrations calculated from the photoelectron peak intensities in the XPS spectra – Inox plates Figure 2: So-called relative atomic surface concentrations calculated from the photoelectron peak intensities in the XPS spectra – Al plates |