| X-ray photoelectron spectroscopy (XPS) report |
| --- |
| Client: Barbara Holzer, AB | Tel.: |
| Job No.: 2013-XPS-TS-MME-CCS-2007-04-24 |
| Job opened by: M.Taborelli, D.Letant-Delrieux | Date: 2007-04-26 |
| Number of samples: 4 Sample codification: 2013 2013-a-1: Stainless steel platelet contaminated at CERN and cleaned at IHEP (20/03/07) 2013-b-1: Stainless steel platelet contaminated at CERN and cleaned at IHEP (20/03/07) 2013-c-1: Aluminium platelet contaminated at CERN and cleaned at IHEP (20/03/07) 2013-d-1: Aluminium platelet contaminated at CERN and cleaned at IHEP (20/03/07) Nature of samples: Plates received wrapped in Aluminium foil (16/04/07) Sample storage: air. Aim of analysis: Verify the cleanness. Results Two samples were analyzed for each metal. The results obtained from the X-rays photoelectron spectroscopy analysis are summarized in fig. 1. We recall that the acceptance level of the amount of carbon contamination on stainless steel and aluminium samples is fixed at 40 at. % and 29 at. %, respectively. The stainless steel and aluminium samples are clean. Minor traces of P are detected on the stainless steel surface. For both type of metal a rather large signal of Si is measured; the latter is very likely coming from silicates. Probably rinsing was insufficient to eliminate the silicates from the cleaning detergent bath. The quality of rinsing is also responsible for the presence of Ca on the surface. Conclusions: The stainless steel and aluminium samples are sufficiently clean to be used in UHV. Figure 1: So-called relative atomic surface concentrations calculated from the photoelectron peak intensities in the XPS spectra - |