## Merged sheet
MEMORY MAPPING BLECS | [9 empty] | pDev->blecsHVLFRESSETUP1.set((const long int *)pAcq->blecsHVLFRESSETUP1,256);
MERGE | 700000
Logging DB
Name | Alternative name | Database-name | Page | Start | End | Bit position | data type | Bit field | struct entry (automatic)) | FESA Field | Property DataFIeldRef | Property DataFIeldRef logged | struct to Fesa Field | data to field (not logged) | DWORD Description | Comments | STATUS logged | concentrator | total in bits |  | info | Calculation for display | MUX FPGA | Additional decoding information | additional comments | VHDL address decoding
LOGGING
SYS_OP | SysOperational | SYS_OPERATIONAL | OVERVIEW | 700000 | 700004 | 31 | bool | :1 | bool blecsSYS_OP:1; | <field name="blecsSYS_OP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>System is operational (View by the BLECS)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsSYS_OP"/> | pDev->blecsSYS_OP.set(pAcq->blecsSYS_OP); |  | System level | System is operational (View by the BLECS) | 1 |  | 1 |  |  |  | 4 |  |  | if (vme_addr[22..0] >= h"700000") and (vme_addr[22..0] < h"700004") then mux_sel[] = 4; end if; --System level / comments: System is operational (View by the BLECS)
TEST | SysUnderTest | SYS_UNDER_TEST | OVERVIEW |  |  | 30 | bool | :1 | bool blecsTEST:1; | <field name="blecsTEST" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>System is under test (All BLM crate of the IP are involved but only this crate is tested)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsTEST"/> | pDev->blecsTEST.set(pAcq->blecsTEST); |  |  | System is under test (All BLM crate of the IP are involved but only this crate is tested) | 1 |  | 1
T_REQ | TstRequ | TEST_PEND_GLOBAL | OVERVIEW |  |  | 29 | bool | :1 | bool blecsT_REQ:1; | <field name="blecsT_REQ" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test requested (while take place when the conditions are fillfull)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsT_REQ"/> | pDev->blecsT_REQ.set(pAcq->blecsT_REQ); |  |  | Test requested (while take place when the conditions are fillfull) | 1 |  | 1
FIRM_SERIAL_FAULT |  | FIRM_SERIAL_FAULT | OVERVIEW |  |  | 28 | bool | :1 | bool blecsFIRM_SERIAL_FAULT:1; | <field name="blecsFIRM_SERIAL_FAULT" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Mismatch between flash setup and FPGA firmware or board serial number</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsFIRM_SERIAL_FAULT"/> | pDev->blecsFIRM_SERIAL_FAULT.set(pAcq->blecsFIRM_SERIAL_FAULT); |  |  | Mismatch between flash setup and FPGA firmware or board serial number | 1 |  | 1
HV1 | FaultHighV1Comp | HV_1_COMP_FAULT | OVERVIEW |  |  | 27 | bool | :1 | bool blecsHV1:1; | <field name="blecsHV1" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High voltage 1 fault (survey by comparators)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsHV1"/> | pDev->blecsHV1.set(pAcq->blecsHV1); |  |  | High voltage 1 fault (survey by comparators) | 1 |  | 1
HV2 | FaultHighV2Comp | HV_2_COMP_FAULT | OVERVIEW |  |  | 26 | bool | :1 | bool blecsHV2:1; | <field name="blecsHV2" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High voltage 2 fault (survey by comparators)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsHV2"/> | pDev->blecsHV2.set(pAcq->blecsHV2); |  |  | High voltage 2 fault (survey by comparators) | 1 |  | 1
HV12ADC | FaultHighVAdc | HV_ADC_FAULT | OVERVIEW |  |  | 25 | bool | :1 | bool blecsHV12ADC:1; | <field name="blecsHV12ADC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High voltage 1 or 2 fault (survey by ADC)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsHV12ADC"/> | pDev->blecsHV12ADC.set(pAcq->blecsHV12ADC); |  |  | High voltage 1 or 2 fault (survey by ADC) | 1 |  | 1
LVCC | FaultLowVComp | LV_COMP_FAULT | OVERVIEW |  |  | 24 | bool | :1 | bool blecsLVCC:1; | <field name="blecsLVCC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Low voltage fault with (survey by comparators)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsLVCC"/> | pDev->blecsLVCC.set(pAcq->blecsLVCC); |  |  | Low voltage fault with (survey by comparators) | 1 |  | 1
LVA | FaultLowVAdc | LV_ADC_FAULT | OVERVIEW |  |  | 23 | bool | :1 | bool blecsLVA:1; | <field name="blecsLVA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Low voltage fault (survey by ADC)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsLVA"/> | pDev->blecsLVA.set(pAcq->blecsLVA); |  |  | Low voltage fault (survey by ADC) | 1 |  | 1
BE | FaultBmEngy | BEAM_ENGY_FAULT | OVERVIEW |  |  | 22 | bool | :1 | bool blecsBE:1; | <field name="blecsBE" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Energy fault (From CISV card AB/CO)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsBE"/> | pDev->blecsBE.set(pAcq->blecsBE); |  |  | Beam Energy fault (From CISV card AB/CO) | 1 |  | 1
TCKF | FaultTurnClk | TURN_CLK_FAULT | OVERVIEW |  |  | 21 | bool | :1 | bool blecsTCKF:1; | <field name="blecsTCKF" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Turn Clock fault (From BOBR card)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsTCKF"/> | pDev->blecsTCKF.set(pAcq->blecsTCKF); |  |  | Turn Clock fault (From BOBR card) | 1 |  | 1
DUMP | DumpOccured | DUMP_OCCURED | OVERVIEW |  |  | 20 | bool | :1 | bool blecsDUMP:1; | <field name="blecsDUMP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Dump occurred</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsDUMP"/> | pDev->blecsDUMP.set(pAcq->blecsDUMP); |  |  | Dump occurred | 1 |  | 1
UST | TstPendUsrSysTest | TEST_PEND_USR_SYS_TST | TESTS |  |  | 19 | bool | :1 | bool blecsUST:1; | <field name="blecsUST" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User System test pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsUST"/> | pDev->blecsUST.set(pAcq->blecsUST); |  |  | User System test pending | 1 |  | 1
UCT | TstPendUsrConsist | TEST_PEND_USR_CONSIST | TESTS |  |  | 18 | bool | :1 | bool blecsUCT:1; | <field name="blecsUCT" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User Consistency test pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsUCT"/> | pDev->blecsUCT.set(pAcq->blecsUCT); |  |  | User Consistency test pending | 1 |  | 1
UTBPL | TstPendUsrThresBpl | TEST_PEND_USR_THRES_BPL | TESTS |  |  | 17 | bool | :1 | bool blecsUTBPL:1; | <field name="blecsUTBPL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User Threshold to BPL pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUTBPL"/> |  | pDev->blecsUTBPL.set(pAcq->blecsUTBPL); | if (data.itemAvailable("blecsUTBPL") == true)blecsAcq.blecsUTBPL = data.blecsUTBPL.get(); |  | User Threshold to BPL pending
UET | TstPendUsrEngy | TEST_PEND_USR_ENGY | TESTS |  |  | 16 | bool | :1 | bool blecsUET:1; | <field name="blecsUET" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User Energy Test pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUET"/> |  | pDev->blecsUET.set(pAcq->blecsUET); | if (data.itemAvailable("blecsUET") == true)blecsAcq.blecsUET = data.blecsUET.get(); |  | User Energy Test pending
UBPBIS | TstPendUsrBmPermBis | TEST_PEND_USR_BP_BIS | TESTS |  |  | 15 | bool | :1 | bool blecsUBPBIS:1; | <field name="blecsUBPBIS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User BPBIS (BLECS to Beam interlock) pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsUBPBIS"/> | pDev->blecsUBPBIS.set(pAcq->blecsUBPBIS); |  |  | User BPBIS (BLECS to Beam interlock) pending | 1 |  | 1 |  | second
EST | TstPendExpSysTest | TEST_PEND_EXP_SYS_TST | TESTS |  |  | 14 | bool | :1 | bool blecsEST:1; | <field name="blecsEST" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert System test pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsEST"/> | pDev->blecsEST.set(pAcq->blecsEST); |  |  | Expert System test pending | 1 |  | 1 |  | every
ECT | TstPendExpConsist | TEST_PEND_EXP_CONSIST | TESTS |  |  | 13 | bool | :1 | bool blecsECT:1; | <field name="blecsECT" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Consistency test pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsECT"/> | pDev->blecsECT.set(pAcq->blecsECT); |  |  | Expert Consistency test pending | 1 |  | 1 |  | Read it
ETBPL | TstPendExpThresBpl | TEST_PEND_EXP_THRES_BP | TESTS |  |  | 12 | bool | :1 | bool blecsETBPL:1; | <field name="blecsETBPL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Threshold to BPL pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsETBPL"/> | pDev->blecsETBPL.set(pAcq->blecsETBPL); |  |  | Expert Threshold to BPL pending | 1 |  | 1
EET | TstPendExpEngy | TEST_PEND_EXP_ENGY | TESTS |  |  | 11 | bool | :1 | bool blecsEET:1; | <field name="blecsEET" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Energy Test pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsEET"/> | pDev->blecsEET.set(pAcq->blecsEET); |  |  | Expert Energy Test pending | 1 |  | 1
EBPBIS | TstPendExpBmPermBis | TEST_PEND_EXP_BP_BIS | TESTS |  |  | 10 | bool | :1 | bool blecsEBPBIS:1; | <field name="blecsEBPBIS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert BPBIS (BLECS to Beam interlock) pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsEBPBIS"/> | pDev->blecsEBPBIS.set(pAcq->blecsEBPBIS); |  |  | Expert BPBIS (BLECS to Beam interlock) pending | 1 |  | 1
EBPTC | TstPendExpBmPermTc | TEST_PEND_EXP_BP_TC | TESTS |  |  | 9 | bool | :1 | bool blecsEBPTC:1; | <field name="blecsEBPTC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert BPTC (Beam Permit BLETC to BLECS) pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsEBPTC"/> | pDev->blecsEBPTC.set(pAcq->blecsEBPTC); |  |  | Expert BPTC (Beam Permit BLETC to BLECS) pending | 1 |  | 1
EHVLF | TstPendExpLowFreqModul | TEST_PEND_EXP_LF_MOD | TESTS |  |  | 8 | bool | :1 | bool blecsEHVLF:1; | <field name="blecsEHVLF" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVLF test pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsEHVLF"/> | pDev->blecsEHVLF.set(pAcq->blecsEHVLF); |  |  | Expert HVLF test pending | 1 |  | 1
EHVCFC | TstPendExpCfcTest | TEST_PEND_EXP_CFC_TST | TESTS |  |  | 7 | bool | :1 | bool blecsEHVCFC:1; | <field name="blecsEHVCFC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVCFC (100 pA test) pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsEHVCFC"/> | pDev->blecsEHVCFC.set(pAcq->blecsEHVCFC); |  |  | Expert HVCFC (100 pA test) pending | 1 |  | 1
EHVRDAC | TstPendExpRstDac | TEST_PEND_EXP_RST_DAC | TESTS |  |  | 6 | bool | :1 | bool blecsEHVRDAC:1; | <field name="blecsEHVRDAC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVRDAC (Dac Reset) pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsEHVRDAC"/> | pDev->blecsEHVRDAC.set(pAcq->blecsEHVRDAC); |  |  | Expert HVRDAC (Dac Reset) pending | 1 |  | 1
EHVRGOH | TstPendExpRstGoh | TEST_PEND_EXP_RST_GOH | TESTS |  |  | 5 | bool | :1 | bool blecsEHVRGOH:1; | <field name="blecsEHVRGOH" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVRGOH (GOH Reset) pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsEHVRGOH"/> | pDev->blecsEHVRGOH.set(pAcq->blecsEHVRGOH); |  |  | Expert HVRGOH (GOH Reset) pending | 1 |  | 1
EMC | TstPendExpManCtrl | TEST_PEND_EXP_MAN_CTRL | TESTS |  |  | 4 | bool | :1 | bool blecsEMC:1; | <field name="blecsEMC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Manual control Request pending</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsEMC"/> | pDev->blecsEMC.set(pAcq->blecsEMC); |  |  | Expert Manual control Request pending | 1 |  | 1
TST | TstPendExpSysTest | TEST_PEND_TMR_SYS_TST | TESTS |  |  | 3 | bool | :1 | bool blecsTST:1; | <field name="blecsTST" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Timer System Test pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsTST"/> |  | pDev->blecsTST.set(pAcq->blecsTST); | if (data.itemAvailable("blecsTST") == true)blecsAcq.blecsTST = data.blecsTST.get(); |  | Timer System Test pending
TSTRN | TstPendTimerSysTestNorm | TEST_REQUEST_BY_TIMER_NORM | TESTS |  |  | 2 | bool | :1 | bool blecsTSTRN:1; | <field name="blecsTSTRN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Internal timer is requesting a test of the system with normal priority</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsTSTRN"/> | pDev->blecsTSTRN.set(pAcq->blecsTSTRN); |  |  | Internal timer is requesting a test of the system with normal priority | 1 |  | 1
TSTRP | TstPendTimerSysTestPrio | TEST_REQUEST_BY_TIMER_PRIO | TESTS |  |  | 1 | bool | :1 | bool blecsTSTRP:1; | <field name="blecsTSTRP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Internal timer is requesting a test of the system with high priority (BeamPermit away next Dump)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsTSTRP"/> | pDev->blecsTSTRP.set(pAcq->blecsTSTRP); |  |  | Internal timer is requesting a test of the system with high priority (BeamPermit away next Dump) | 1 |  | 1
LBCIBU | LastCombBfrCibus |  | OVERVIEW |  |  | 0 | bool | :1 | bool blecsLBCIBU:1; | <field name="blecsLBCIBU" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>BLECS (Combiner) is the last one before the CIBU interface / controls the HV</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsLBCIBU"/> |  | pDev->blecsLBCIBU.set(pAcq->blecsLBCIBU); | if (data.itemAvailable("blecsLBCIBU") == true)blecsAcq.blecsLBCIBU = data.blecsLBCIBU.get(); |  | BLECS (Combiner) is the last one before the CIBU interface / controls the HV
NSTR | TimeBfrNxtSysTst | TIME_BFR_NXT_SYS_TST | TESTS | 700004 | 700008 |  | unsigned | :32 | unsigned blecsNSTR:32; | <field name="blecsNSTR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Time remained before normal system test [s]</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsNSTR"/> |  | pDev->blecsNSTR.set(pAcq->blecsNSTR); | if (data.itemAvailable("blecsNSTR") == true)blecsAcq.blecsNSTR = data.blecsNSTR.get(); | Next system test request | Time remained before normal system test [s] |  | 1 | 32 |  |  |  | 5 |  |  | if (vme_addr[22..0] >= h"700004") and (vme_addr[22..0] < h"700008") then mux_sel[] = 5; end if; --Next system test request / comments: Time remained before normal system test [s]
NCSTR | TimeBfrNxtSysTstCritical | TIME_BFR_NXT_SYS_TST_CRIT | TESTS | 700008 | 70000C |  | unsigned | :32 | unsigned blecsNCSTR:32; | <field name="blecsNCSTR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Time remained before cricical system test [s]</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsNCSTR"/> |  | pDev->blecsNCSTR.set(pAcq->blecsNCSTR); | if (data.itemAvailable("blecsNCSTR") == true)blecsAcq.blecsNCSTR = data.blecsNCSTR.get(); | Next critical system test request | Time remained before cricical system test [s] |  | 1 | 32 |  |  |  | 6 |  |  | if (vme_addr[22..0] >= h"700008") and (vme_addr[22..0] < h"70000C") then mux_sel[] = 6; end if; --Next critical system test request / comments: Time remained before cricical system test [s]
AST | TstActiveSysTest | TEST_ACTIVE_SYS_TST | TESTS | 70000C | 700010 | 31 | bool | :1 | bool blecsAST:1; | <field name="blecsAST" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsAST"/> | pDev->blecsAST.set(pAcq->blecsAST); |  | Active tests overview | Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested) | 1 |  | 1 |  |  |  | 7 |  |  | if (vme_addr[22..0] >= h"70000C") and (vme_addr[22..0] < h"700010") then mux_sel[] = 7; end if; --Active tests overview / comments: Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested)
ACT | TstActiveConsist | TEST_ACTIVE_CONSIST | TESTS |  |  | 30 | bool | :1 | bool blecsACT:1; | <field name="blecsACT" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : CONSISTENCY (threshold table control)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsACT"/> | pDev->blecsACT.set(pAcq->blecsACT); |  |  | Active test : CONSISTENCY (threshold table control) | 1 |  | 1
ATBPL | TstActiveThresBpl | TEST_ACTIVE_THRES_BP | TESTS |  |  | 29 | bool | :1 | bool blecsATBPL:1; | <field name="blecsATBPL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : THRESHOLD to BPL</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsATBPL"/> | pDev->blecsATBPL.set(pAcq->blecsATBPL); |  |  | Active test : THRESHOLD to BPL | 1 |  | 1
AET | TstActiveEngy | TEST_ACTIVE_ENGY | TESTS |  |  | 28 | bool | :1 | bool blecsAET:1; | <field name="blecsAET" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : ENERGY TEST</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsAET"/> | pDev->blecsAET.set(pAcq->blecsAET); |  |  | Active test : ENERGY TEST | 1 |  | 1
ABPBIS | TstActiveBmPermBis | TEST_ACTIVE_BP_BIS | TESTS |  |  | 27 | bool | :1 | bool blecsABPBIS:1; | <field name="blecsABPBIS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsABPBIS"/> | pDev->blecsABPBIS.set(pAcq->blecsABPBIS); |  |  | Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 1 |  | 1
ABPTC | TstActiveBmPermTc | TEST_ACTIVE_BP_TC | TESTS |  |  | 26 | bool | :1 | bool blecsABPTC:1; | <field name="blecsABPTC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsABPTC"/> | pDev->blecsABPTC.set(pAcq->blecsABPTC); |  |  | Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 1 |  | 1
AHVLF | TstActiveLowFreqModul | TEST_ACTIVE_LF_MOD | TESTS |  |  | 25 | bool | :1 | bool blecsAHVLF:1; | <field name="blecsAHVLF" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : HVLF (Modulation of the high voltage and analysis)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsAHVLF"/> | pDev->blecsAHVLF.set(pAcq->blecsAHVLF); |  |  | Active test : HVLF (Modulation of the high voltage and analysis) | 1 |  | 1
AHVCFC | TstActiveCfcTest | TEST_ACTIVE_CFC_TST | TESTS |  |  | 24 | bool | :1 | bool blecsAHVCFC:1; | <field name="blecsAHVCFC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsAHVCFC"/> | pDev->blecsAHVCFC.set(pAcq->blecsAHVCFC); |  |  | Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card) | 1 |  | 1
AHVRDAC | TstActiveRstDac | TEST_ACTIVE_RST_DAC | TESTS |  |  | 23 | bool | :1 | bool blecsAHVRDAC:1; | <field name="blecsAHVRDAC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : HVRDAC (Activation DAC reset mode of the tunnel card)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsAHVRDAC"/> | pDev->blecsAHVRDAC.set(pAcq->blecsAHVRDAC); |  |  | Active test : HVRDAC (Activation DAC reset mode of the tunnel card) | 1 |  | 1
AHVRGOH | TstActiveRstGoh | TEST_ACTIVE_RST_GOH | TESTS |  |  | 22 | bool | :1 | bool blecsAHVRGOH:1; | <field name="blecsAHVRGOH" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : HVRGOH (Activation GOH reset mode of the tunnel card)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsAHVRGOH"/> | pDev->blecsAHVRGOH.set(pAcq->blecsAHVRGOH); |  |  | Active test : HVRGOH (Activation GOH reset mode of the tunnel card) | 1 |  | 1
MCAL | TstAllowedManCtrl | TEST_ALLOWED_MAN_CTRL | OVERVIEW |  |  | 21 | bool | :1 | bool blecsMCAL:1; | <field name="blecsMCAL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : MANUAL CONTROL ALLOWED</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsMCAL"/> | pDev->blecsMCAL.set(pAcq->blecsMCAL); |  |  | Active test : MANUAL CONTROL ALLOWED | 1 |  | 1
MCAC | TstActiveManCtrl | TEST_ACTIVE_MAN_CTRL | Overview + Tests |  |  | 20 | bool | :1 | bool blecsMCAC:1; | <field name="blecsMCAC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : MANUAL CONTROL ACTIVE</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsMCAC"/> | pDev->blecsMCAC.set(pAcq->blecsMCAC); |  |  | Active test : MANUAL CONTROL ACTIVE | 1 |  | 1
|  |  | - |  |  | 15..19 | unsigned | :5 | unsigned blecsSPARE53:5; | [6 empty] | SPARE
SW1 | SwitchAllowWriteReg |  | OVERVIEW |  |  | 14 | bool | :1 | bool blecsSW1:1; | <field name="blecsSW1" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>SWITCH1: if '1' allow to write into registers (701D04 - 701D4F)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsSW1"/> |  | pDev->blecsSW1.set(pAcq->blecsSW1); | if (data.itemAvailable("blecsSW1") == true)blecsAcq.blecsSW1 = data.blecsSW1.get(); |  | SWITCH1: if '1' allow to write into registers (701D04 - 701D4F) |  |  | 1
SW2 | Switch2 |  | OVERVIEW |  |  | 13 | bool | :1 | bool blecsSW2:1; | <field name="blecsSW2" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>SWITCH2</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsSW2"/> |  | pDev->blecsSW2.set(pAcq->blecsSW2); | if (data.itemAvailable("blecsSW2") == true)blecsAcq.blecsSW2 = data.blecsSW2.get(); |  | SWITCH2
|  |  | - |  |  | 12..8 | unsigned | :5 | unsigned blecsSPARE56:5; | [6 empty] | SPARE
BLECFALLTON | BlecfAllStatusTstCfcOkOn | BLECF_ALL_TST_CFC_ON | BLECF TEST |  |  | 7 | bool | :1 | bool blecsBLECFALLTON:1; | <field name="blecsBLECFALLTON" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>All BLECF in Test CFC (100pA test) Active</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFALLTON"/> |  | pDev->blecsBLECFALLTON.set(pAcq->blecsBLECFALLTON); | if (data.itemAvailable("blecsBLECFALLTON") == true)blecsAcq.blecsBLECFALLTON = data.blecsBLECFALLTON.get(); |  | All BLECF in Test CFC (100pA test) Active |  | 1 | 1
BLECFALLRDON | BlecfAllStatusRstDacOkOn | BLECF_ALL_RST_DAC_ON | BLECF TEST |  |  | 6 | bool | :1 | bool blecsBLECFALLRDON:1; | <field name="blecsBLECFALLRDON" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>All BLECF in Reset DAC Active</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFALLRDON"/> |  | pDev->blecsBLECFALLRDON.set(pAcq->blecsBLECFALLRDON); | if (data.itemAvailable("blecsBLECFALLRDON") == true)blecsAcq.blecsBLECFALLRDON = data.blecsBLECFALLRDON.get(); |  | All BLECF in Reset DAC Active |  | 1 | 1
BLECFALLRGON | BlecfAllStatusRstGohOkOn | BLECF_ALL_RST_GOH_ON | BLECF TEST |  |  | 5 | bool | :1 | bool blecsBLECFALLRGON:1; | <field name="blecsBLECFALLRGON" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>All BLECF in Reset GOH Active</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFALLRGON"/> |  | pDev->blecsBLECFALLRGON.set(pAcq->blecsBLECFALLRGON); | if (data.itemAvailable("blecsBLECFALLRGON") == true)blecsAcq.blecsBLECFALLRGON = data.blecsBLECFALLRGON.get(); |  | All BLECF in Reset GOH Active |  | 1 | 1
BLECFALLHV | BlecfAllStatusHvOk | BLECF_ALL_HV_OK | BLECF TEST |  |  | 4 | bool | :1 | bool blecsBLECFALLHV:1; | <field name="blecsBLECFALLHV" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>All BLECF HV ok</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFALLHV"/> |  | pDev->blecsBLECFALLHV.set(pAcq->blecsBLECFALLHV); | if (data.itemAvailable("blecsBLECFALLHV") == true)blecsAcq.blecsBLECFALLHV = data.blecsBLECFALLHV.get(); |  | All BLECF HV ok |  | 1 | 1
BLECFALLT | BlecfAllStatusTstCfcOk | BLECF_ALL_TST_CFC_OK | BLECF TEST |  |  | 3 | bool | :1 | bool blecsBLECFALLT:1; | <field name="blecsBLECFALLT" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>All BLECF in Test CFC (100pA test)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFALLT"/> |  | pDev->blecsBLECFALLT.set(pAcq->blecsBLECFALLT); | if (data.itemAvailable("blecsBLECFALLT") == true)blecsAcq.blecsBLECFALLT = data.blecsBLECFALLT.get(); |  | All BLECF in Test CFC (100pA test) |  | 1 | 1
BLECFALLRD | BlecfAllStatusRstDacOk | BLECF_ALL_RST_DAC_OK | BLECF TEST |  |  | 2 | bool | :1 | bool blecsBLECFALLRD:1; | <field name="blecsBLECFALLRD" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>All BLECF in Reset DAC</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFALLRD"/> |  | pDev->blecsBLECFALLRD.set(pAcq->blecsBLECFALLRD); | if (data.itemAvailable("blecsBLECFALLRD") == true)blecsAcq.blecsBLECFALLRD = data.blecsBLECFALLRD.get(); |  | All BLECF in Reset DAC |  | 1 | 1
BLECFALLRG | BlecfAllStatusRstGohOk | BLECF_ALL_RST_GOH_OK | BLECF TEST |  |  | 1 | bool | :1 | bool blecsBLECFALLRG:1; | <field name="blecsBLECFALLRG" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>All BLECF in Reset GOH</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFALLRG"/> |  | pDev->blecsBLECFALLRG.set(pAcq->blecsBLECFALLRG); | if (data.itemAvailable("blecsBLECFALLRG") == true)blecsAcq.blecsBLECFALLRG = data.blecsBLECFALLRG.get(); |  | All BLECF in Reset GOH |  | 1 | 1
BLECFER | BlecfNotAllCorrectTst |  | BLECF TEST |  |  | 0 | bool | :1 | bool blecsBLECFER:1; | <field name="blecsBLECFER" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>One or more BLECF not in the correct test</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFER"/> |  | pDev->blecsBLECFER.set(pAcq->blecsBLECFER); | if (data.itemAvailable("blecsBLECFER") == true)blecsAcq.blecsBLECFER = data.blecsBLECFER.get(); |  | One or more BLECF not in the correct test
TSLST | TimeSinceLstSysTst |  | TESTS | 700010 | 700014 |  | unsigned | :32 | unsigned blecsTSLST:32; | <field name="blecsTSLST" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Time since the last system test [s]</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTSLST"/> |  | pDev->blecsTSLST.set(pAcq->blecsTSLST); | if (data.itemAvailable("blecsTSLST") == true)blecsAcq.blecsTSLST = data.blecsTSLST.get(); | Time since last System test | Time since the last system test [s] | [6 empty] | 8 |  |  | if (vme_addr[22..0] >= h"700010") and (vme_addr[22..0] < h"700014") then mux_sel[] = 8; end if; --Time since last System test / comments: Time since the last system test [s]
RSTP | TstPassedSysTst | TEST_PASSED_SYS_TST | TESTS | 700014 | 700018 | 31 | bool | :1 | bool blecsRSTP:1; | <field name="blecsRSTP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : "SYSTEM TEST"</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsRSTP"/> | pDev->blecsRSTP.set(pAcq->blecsRSTP); |  | Result of last tests | Test Passed : "SYSTEM TEST" | 1 |  | 1 |  |  |  | 9 |  |  | if (vme_addr[22..0] >= h"700014") and (vme_addr[22..0] < h"700018") then mux_sel[] = 9; end if; --Result of last tests / comments: Test Passed : "SYSTEM TEST"
RCTP | TstPassedConsist | TEST_PASSED_CONSIST | TESTS |  |  | 30 | bool | :1 | bool blecsRCTP:1; | <field name="blecsRCTP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : CONSISTENCY (threshold table control)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsRCTP"/> | pDev->blecsRCTP.set(pAcq->blecsRCTP); |  |  | Test Passed : CONSISTENCY (threshold table control) | 1 |  | 1
RTBPLP | TstPassedThresBpl | TEST_PASSED_THRES_BPL | TESTS |  |  | 29 | bool | :1 | bool blecsRTBPLP:1; | <field name="blecsRTBPLP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : THRESHOLD to BPL</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsRTBPLP"/> | pDev->blecsRTBPLP.set(pAcq->blecsRTBPLP); |  |  | Test Passed : THRESHOLD to BPL | 1 |  | 1
RETP | TstPassedEngy | TEST_PASSED_ENGY | TESTS |  |  | 28 | bool | :1 | bool blecsRETP:1; | <field name="blecsRETP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : ENERGY TEST</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsRETP"/> | pDev->blecsRETP.set(pAcq->blecsRETP); |  |  | Test Passed : ENERGY TEST | 1 |  | 1
RBPBISP | TstPassedBmPermBis | TEST_PASSED_BP_BIS | TESTS |  |  | 27 | bool | :1 | bool blecsRBPBISP:1; | <field name="blecsRBPBISP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsRBPBISP"/> | pDev->blecsRBPBISP.set(pAcq->blecsRBPBISP); |  |  | Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 1 |  | 1
RBPTCP | TstPassedBmPermTc | TEST_PASSED_BP_TC | TESTS |  |  | 26 | bool | :1 | bool blecsRBPTCP:1; | <field name="blecsRBPTCP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsRBPTCP"/> | pDev->blecsRBPTCP.set(pAcq->blecsRBPTCP); |  |  | Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 1 |  | 1 |  | second
RHVLFP | TstPassedLowFreqModul | TEST_PASSED_LF_MOD | TESTS |  |  | 25 | bool | :1 | bool blecsRHVLFP:1; | <field name="blecsRHVLFP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : HVLF (Modulation of the high voltage and analysis)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsRHVLFP"/> | pDev->blecsRHVLFP.set(pAcq->blecsRHVLFP); |  |  | Test Passed : HVLF (Modulation of the high voltage and analysis) | 1 |  | 1 |  | every
RHVTCFCP | TstPassedCfcTest | TEST_PASSED_CFC_TEST | TESTS |  |  | 24 | bool | :1 | bool blecsRHVTCFCP:1; | <field name="blecsRHVTCFCP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsRHVTCFCP"/> | pDev->blecsRHVTCFCP.set(pAcq->blecsRHVTCFCP); |  |  | Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card) | 1 |  | 1 |  | Read it
|  |  | - |  |  | 23..1 | unsigned | :23 | unsigned blecsSPARE74:23; | [6 empty] | SPARE
TRR | TstResRdy | TEST_RESULT_RDY | TESTS |  |  | 0 | bool | :1 | bool blecsTRR:1; | <field name="blecsTRR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>New test results are available (Check active test to find out which test)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsTRR"/> | pDev->blecsTRR.set(pAcq->blecsTRR); |  |  | New test results are available (Check active test to find out which test) | 1 |  | 1
TSLR | TimeSinceLstRst |  | OVERVIEW | 700018 | 70001C |  | unsigned | :32 | unsigned blecsTSLR:32; | <field name="blecsTSLR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Time since the last reset of the board [s]</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTSLR"/> |  | pDev->blecsTSLR.set(pAcq->blecsTSLR); | if (data.itemAvailable("blecsTSLR") == true)blecsAcq.blecsTSLR = data.blecsTSLR.get(); | Time | Time since the last reset of the board [s] | [6 empty] | 10 |  |  | if (vme_addr[22..0] >= h"700018") and (vme_addr[22..0] < h"70001C") then mux_sel[] = 10; end if; --Time / comments: Time since the last reset of the board [s]
UDL | DumpLineUmsk | BEAM_PERM_TC_LINE_U | Beam PERMIT | 70001C | 700020 | 31 | bool | :1 | bool blecsUDL:1; | <field name="blecsUDL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold Comparators Unmaskable beam permit line arriving to the BLECS</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsUDL"/> | pDev->blecsUDL.set(pAcq->blecsUDL); |  | User Beam Permit | Threshold Comparators Unmaskable beam permit line arriving to the BLECS | 1 |  | 1 |  |  |  | 11 |  |  | if (vme_addr[22..0] >= h"70001C") and (vme_addr[22..0] < h"700020") then mux_sel[] = 11; end if; --User Beam Permit / comments: Threshold Comparators Unmaskable beam permit line arriving to the BLECS
MDL | DumpLineMsk | BEAM_PERM_TC_LINE_M | Beam PERMIT |  |  | 30 | bool | :1 | bool blecsMDL:1; | <field name="blecsMDL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold Comparators Maskable beam permit line arriving to the BLECS</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsMDL"/> | pDev->blecsMDL.set(pAcq->blecsMDL); |  |  | Threshold Comparators Maskable beam permit line arriving to the BLECS | 1 |  | 1
UAI | BmPermUmskInA | BEAM_PERM_U_IN_A | Beam PERMIT |  |  | 29 | bool | :1 | bool blecsUAI:1; | <field name="blecsUAI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Unmaskable A Input from the upper Combiner</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsUAI"/> | pDev->blecsUAI.set(pAcq->blecsUAI); |  |  | Beam Permit Unmaskable A Input from the upper Combiner | 1 |  | 1
UBI | BmPermUmskInB | BEAM_PERM_U_IN_B | Beam PERMIT |  |  | 28 | bool | :1 | bool blecsUBI:1; | <field name="blecsUBI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Unmaskable B Input from the upper Combiner</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsUBI"/> | pDev->blecsUBI.set(pAcq->blecsUBI); |  |  | Beam Permit Unmaskable B Input from the upper Combiner | 1 |  | 1
MAI | BmPermMskInA | BEAM_PERM_M_IN_A | Beam PERMIT |  |  | 27 | bool | :1 | bool blecsMAI:1; | <field name="blecsMAI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Maskable A Input from the upper Combiner</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsMAI"/> | pDev->blecsMAI.set(pAcq->blecsMAI); |  |  | Beam Permit Maskable A Input from the upper Combiner | 1 |  | 1
MBI | BmPermMskInB | BEAM_PERM_M_IN_B | Beam PERMIT |  |  | 26 | bool | :1 | bool blecsMBI:1; | <field name="blecsMBI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Maskable B Input from the upper Combiner</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsMBI"/> | pDev->blecsMBI.set(pAcq->blecsMBI); |  |  | Beam Permit Maskable B Input from the upper Combiner | 1 |  | 1
UAO | BmPermUmskOutA | BEAM_PERM_U_OUT_A | Beam PERMIT |  |  | 25 | bool | :1 | bool blecsUAO:1; | <field name="blecsUAO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Unmaskable A Output to lower Combiner (or CIBUS)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsUAO"/> | pDev->blecsUAO.set(pAcq->blecsUAO); |  |  | Beam Permit Unmaskable A Output to lower Combiner (or CIBUS) | 1 |  | 1
UBO | BmPermUmskOutB | BEAM_PERM_U_OUT_B | Beam PERMIT |  |  | 24 | bool | :1 | bool blecsUBO:1; | <field name="blecsUBO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Unmaskable B Output to lower Combiner (or CIBUS)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsUBO"/> | pDev->blecsUBO.set(pAcq->blecsUBO); |  |  | Beam Permit Unmaskable B Output to lower Combiner (or CIBUS) | 1 |  | 1
MAO | BmPermMskOutA | BEAM_PERM_M_OUT_A | Beam PERMIT |  |  | 23 | bool | :1 | bool blecsMAO:1; | <field name="blecsMAO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Maskable A Output to lower Combiner (or CIBUS)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsMAO"/> | pDev->blecsMAO.set(pAcq->blecsMAO); |  |  | Beam Permit Maskable A Output to lower Combiner (or CIBUS) | 1 |  | 1
MBO | BmPermMskOutB | BEAM_PERM_M_OUT_B | Beam PERMIT |  |  | 22 | bool | :1 | bool blecsMBO:1; | <field name="blecsMBO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Maskable B Output to lower Combiner (or CIBUS)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsMBO"/> | pDev->blecsMBO.set(pAcq->blecsMBO); |  |  | Beam Permit Maskable B Output to lower Combiner (or CIBUS) | 1 |  | 1
UBS | BmInfoUmsk | BEAM_PERM_INFO_U | Beam PERMIT |  |  | 21 | bool | :1 | bool blecsUBS:1; | <field name="blecsUBS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Unmaskable Beam Info (Beam status)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsUBS"/> | pDev->blecsUBS.set(pAcq->blecsUBS); |  |  | Unmaskable Beam Info (Beam status) | 1 |  | 1
MBS | BmInfoMsk | BEAM_PERM_INFO_M | Beam PERMIT |  |  | 20 | bool | :1 | bool blecsMBS:1; | <field name="blecsMBS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Maskable Beam Info (Beam status)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsMBS"/> | pDev->blecsMBS.set(pAcq->blecsMBS); |  |  | Maskable Beam Info (Beam status) | 1 |  | 1
UBSR | BmInfoHadRespToDumpUmsk | BEAM_PERM_CIBUS_U_FAULT | Beam PERMIT |  |  | 19 | bool | :1 | bool blecsUBSR:1; | <field name="blecsUBSR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsUBSR"/> | pDev->blecsUBSR.set(pAcq->blecsUBSR); |  |  | Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP | 1 |  | 1
MBSR | BmInfoHadRespToDumpMsk | BEAM_PERM_CIBUS_M_FAULT | Beam PERMIT |  |  | 18 | bool | :1 | bool blecsMBSR:1; | <field name="blecsMBSR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Maskable Beam Info (beam status) had not respond to last BLECS-DUMP</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsMBSR"/> | pDev->blecsMBSR.set(pAcq->blecsMBSR); |  |  | Maskable Beam Info (beam status) had not respond to last BLECS-DUMP | 1 |  | 1
|  |  | - |  |  | 16..17 | unsigned | :2 | unsigned blecsSPARE91:2; | [6 empty] | SPARE
UDC | CntDumpUmsk |  | Beam PERMIT |  |  | 8..15 | unsigned | :8 | unsigned blecsUDC:8; | <field name="blecsUDC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Unmaskable Dump counter since last reset</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsUDC"/> |  | pDev->blecsUDC.set(pAcq->blecsUDC); | if (data.itemAvailable("blecsUDC") == true)blecsAcq.blecsUDC = data.blecsUDC.get(); |  | Unmaskable Dump counter since last reset
MDC | CntDumpMsk |  | Beam PERMIT |  |  | 0..7 | unsigned | :8 | unsigned blecsMDC:8; | <field name="blecsMDC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Maskable Dump counter since last reset</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsMDC"/> |  | pDev->blecsMDC.set(pAcq->blecsMDC); | if (data.itemAvailable("blecsMDC") == true)blecsAcq.blecsMDC = data.blecsMDC.get(); |  | Maskable Dump counter since last reset
UTTR | TimeBmInfoToRespDumpUmsk |  | Beam PERMIT | 700020 | 700024 | 16..31 | unsigned | :16 | unsigned blecsUTTR:16; | <field name="blecsUTTR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Time delay from unmaskable dump to unmaskable beam status goes off (microsecond resolution)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsUTTR"/> |  | pDev->blecsUTTR.set(pAcq->blecsUTTR); | if (data.itemAvailable("blecsUTTR") == true)blecsAcq.blecsUTTR = data.blecsUTTR.get(); | Time response of the BIS | Time delay from unmaskable dump to unmaskable beam status goes off (microsecond resolution) | [6 empty] | 12 |  |  | if (vme_addr[22..0] >= h"700020") and (vme_addr[22..0] < h"700024") then mux_sel[] = 12; end if; --Time response of the BIS / comments: Time delay from unmaskable dump to unmaskable beam status goes off (microsecond resolution)
MTTR | TimeBmInfoToRespDumpMsk |  | Beam PERMIT |  |  | 0..15 | unsigned | :16 | unsigned blecsMTTR:16; | <field name="blecsMTTR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Time to delay from maskable dump to maskable beam status goes off (microsecond resolution)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsMTTR"/> |  | pDev->blecsMTTR.set(pAcq->blecsMTTR); | if (data.itemAvailable("blecsMTTR") == true)blecsAcq.blecsMTTR = data.blecsMTTR.get(); |  | Time to delay from maskable dump to maskable beam status goes off (microsecond resolution)
TCLD | TurnCntWhenDumpReq | TURN_CNT_WN_DUMP_REQ | Beam PERMIT | 700024 | 700028 | 0..31 | unsigned | :32 | unsigned blecsTCLD:32; | <field name="blecsTCLD" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Turn counter value when the last dump was requested (32 bits )</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsTCLD"/> | pDev->blecsTCLD.set(pAcq->blecsTCLD); |  | Turn counter when last dump request | Turn counter value when the last dump was requested (32 bits ) | 1 |  | 32 |  |  |  | 13 |  |  | if (vme_addr[22..0] >= h"700024") and (vme_addr[22..0] < h"700028") then mux_sel[] = 13; end if; --Turn counter when last dump request / comments: Turn counter value when the last dump was requested (32 bits )
|  |  |  | 700028 | 70002C | 12..31 | unsigned | :20 | unsigned blecsSPARE97:20; | [5 empty] | Bunch counter when last dump request | SPARE | [6 empty] | 14 |  |  | if (vme_addr[22..0] >= h"700028") and (vme_addr[22..0] < h"70002C") then mux_sel[] = 14; end if; --Bunch counter when last dump request / comments: SPARE
BCLD | BunchCntWhenDumpReq | BUNCH_CNT_WN_DUMP_REQ | Beam PERMIT |  |  | 0..11 | unsigned | :12 | unsigned blecsBCLD:12; | <field name="blecsBCLD" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Bunch counter value when the last dump was requested (12 bits)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBCLD"/> | pDev->blecsBCLD.set(pAcq->blecsBCLD); |  |  | Bunch counter value when the last dump was requested (12 bits) | 1 |  | 12
TDPM | TimeDumpToPmfreezeMsk | BEAM_PERM_TIME_TO_PM_FRZE_M | Beam PERMIT | 70002C | 700030 | 31..16 | unsigned | :16 | unsigned blecsTDPM:16; | <field name="blecsTDPM" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Time from the dump request maskable to PM_freeze (microsecond resolution)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsTDPM"/> | pDev->blecsTDPM.set(pAcq->blecsTDPM); |  | Time from dump to PM freeze | Time from the dump request maskable to PM_freeze (microsecond resolution) | 1 |  | 16 |  |  |  | 15 |  |  | if (vme_addr[22..0] >= h"70002C") and (vme_addr[22..0] < h"700030") then mux_sel[] = 15; end if; --Time from dump to PM freeze / comments: Time from the dump request maskable to PM_freeze (microsecond resolution)
TDPU | TimeDumpToPmfreezeUmsk | BEAM_PERM_TIME_TO_PM_FRZE_U | Beam PERMIT |  |  | 15..0 | unsigned | :16 | unsigned blecsTDPU:16; | <field name="blecsTDPU" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Time from the dump request unmaskable to PM_freeze (microsecond resolution)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsTDPU"/> | pDev->blecsTDPU.set(pAcq->blecsTDPU); |  |  | Time from the dump request unmaskable to PM_freeze (microsecond resolution) | 1 |  | 16
BEI | BmEngyIn | BEAM_ENGY_IN | Beam ENERGY | 700030 | 700034 | 16..31 | unsigned | :16 | unsigned blecsBEI:16; | <field name="blecsBEI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Energy Input coming from timing distribution (CISV card AB/CO)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBEI"/> | pDev->blecsBEI.set(pAcq->blecsBEI); |  | Beam Energy 1 | Beam Energy Input coming from timing distribution (CISV card AB/CO) | 1 |  | 16 |  |  |  | 16 |  |  | if (vme_addr[22..0] >= h"700030") and (vme_addr[22..0] < h"700034") then mux_sel[] = 16; end if; --Beam Energy 1 / comments: Beam Energy Input coming from timing distribution (CISV card AB/CO)
BEO | BmEngyOut | BEAM_ENGY_LEVEL_OUT | Beam ENERGY |  |  | 11..15 | unsigned | :5 | unsigned blecsBEO:5; | <field name="blecsBEO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Energy levels output to the Threshold Comparators (send to BLETC cards)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBEO"/> | pDev->blecsBEO.set(pAcq->blecsBEO); |  |  | Beam Energy levels output to the Threshold Comparators (send to BLETC cards) | 1 |  | 5 |  | second
BEE | BmEngyErr | BEAM_ENGY_ERR | Beam ENERGY |  |  | 10 | bool | :1 | bool blecsBEE:1; | <field name="blecsBEE" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Energy error (send to BLETC cards)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsBEE"/> | pDev->blecsBEE.set(pAcq->blecsBEE); |  |  | Beam Energy error (send to BLETC cards) | 1 |  | 1 |  | every
SRTC | BmEngySoftRstTC | BEAM_ENGY_SOFT_RST_TC | Beam ENERGY |  |  | 9 | bool | :1 | bool blecsSRTC:1; | <field name="blecsSRTC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>SoftResetTC (send to BLETC cards)</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsSRTC"/> | pDev->blecsSRTC.set(pAcq->blecsSRTC); |  |  | SoftResetTC (send to BLETC cards) | 1 |  | 1 |  | Read it
SUT | BmEngySysUnderTst |  | Beam ENERGY |  |  | 8 | bool | :1 | bool blecsSUT:1; | <field name="blecsSUT" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>SystemUnderTest (send to BLETC cards)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsSUT"/> |  | pDev->blecsSUT.set(pAcq->blecsSUT); | if (data.itemAvailable("blecsSUT") == true)blecsAcq.blecsSUT = data.blecsSUT.get(); |  | SystemUnderTest (send to BLETC cards)
BEUBI | BmEngyBmInfoUmsk |  | Beam ENERGY |  |  | 7 | bool | :1 | bool blecsBEUBI:1; | <field name="blecsBEUBI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Unmaskable Beam Info (send to BLETC cards)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBEUBI"/> |  | pDev->blecsBEUBI.set(pAcq->blecsBEUBI); | if (data.itemAvailable("blecsBEUBI") == true)blecsAcq.blecsBEUBI = data.blecsBEUBI.get(); |  | Unmaskable Beam Info (send to BLETC cards)
BEMBI | BmEngyBmInfoMsk |  | Beam ENERGY |  |  | 6 | bool | :1 | bool blecsBEMBI:1; | <field name="blecsBEMBI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Maskable Beam Info (send to BLETC cards)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBEMBI"/> |  | pDev->blecsBEMBI.set(pAcq->blecsBEMBI); | if (data.itemAvailable("blecsBEMBI") == true)blecsAcq.blecsBEMBI = data.blecsBEMBI.get(); |  | Maskable Beam Info (send to BLETC cards)
BPULTA | BmPermTcActTstUmsk |  | Beam PERMIT |  |  | 5 | bool | :1 | bool blecsBPULTA:1; | <field name="blecsBPULTA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam permit line unmaskable test activation (send to BLETC cards)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBPULTA"/> |  | pDev->blecsBPULTA.set(pAcq->blecsBPULTA); | if (data.itemAvailable("blecsBPULTA") == true)blecsAcq.blecsBPULTA = data.blecsBPULTA.get(); |  | Beam permit line unmaskable test activation (send to BLETC cards)
BPMLTA | BmPermTcActTstMsk |  | Beam PERMIT |  |  | 4 | bool | :1 | bool blecsBPMLTA:1; | <field name="blecsBPMLTA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam permit line maskable test activation (send to BLETC cards)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBPMLTA"/> |  | pDev->blecsBPMLTA.set(pAcq->blecsBPMLTA); | if (data.itemAvailable("blecsBPMLTA") == true)blecsAcq.blecsBPMLTA = data.blecsBPMLTA.get(); |  | Beam permit line maskable test activation (send to BLETC cards)
BPLTCN | BmPermTcActCrdNbr |  | Beam PERMIT |  |  | 0..3 | unsigned | :4 | unsigned blecsBPLTCN:4; | <field name="blecsBPLTCN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam permit line test TC card number (send to BLETC cards)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBPLTCN"/> |  | pDev->blecsBPLTCN.set(pAcq->blecsBPLTCN); | if (data.itemAvailable("blecsBPLTCN") == true)blecsAcq.blecsBPLTCN = data.blecsBPLTCN.get(); |  | Beam permit line test TC card number (send to BLETC cards)
BEFCA | BmEngyFrmCntA |  | Beam ENERGY | 700034 | 700038 | 16..31 | unsigned | :16 | unsigned blecsBEFCA:16; | <field name="blecsBEFCA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam energy reception Frame counter channel A</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBEFCA"/> |  | pDev->blecsBEFCA.set(pAcq->blecsBEFCA); | if (data.itemAvailable("blecsBEFCA") == true)blecsAcq.blecsBEFCA = data.blecsBEFCA.get(); | Beam Energy 2 | Beam energy reception Frame counter channel A | [6 empty] | 17 |  |  | if (vme_addr[22..0] >= h"700034") and (vme_addr[22..0] < h"700038") then mux_sel[] = 17; end if; --Beam Energy 2 / comments: Beam energy reception Frame counter channel A
BEFCB | BmEngyFrmCntB |  | Beam ENERGY |  |  | 0..15 | unsigned | :16 | unsigned blecsBEFCB:16; | <field name="blecsBEFCB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam energy reception Frame counter channel B</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBEFCB"/> |  | pDev->blecsBEFCB.set(pAcq->blecsBEFCB); | if (data.itemAvailable("blecsBEFCB") == true)blecsAcq.blecsBEFCB = data.blecsBEFCB.get(); |  | Beam energy reception Frame counter channel B
BECRCA | BmEngyCrcCntA | BEAM_ENGY_CRC_CNT_A | Beam ENERGY | 700038 | 70003C | 16..31 | unsigned | :16 | unsigned blecsBECRCA:16; | <field name="blecsBECRCA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBECRCA"/> | pDev->blecsBECRCA.set(pAcq->blecsBECRCA); |  | Beam Energy 3 | Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 1 |  | 16 |  |  |  | 18 |  |  | if (vme_addr[22..0] >= h"700038") and (vme_addr[22..0] < h"70003C") then mux_sel[] = 18; end if; --Beam Energy 3 / comments: Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO)
BECRCB | BmEngyCrcCntB | BEAM_ENGY_CRC_CNT_B | Beam ENERGY |  |  | 0..15 | unsigned | :16 | unsigned blecsBECRCB:16; | <field name="blecsBECRCB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Number of CRC errors at the reception of the beam energy (channel B) from the CISV card (AB/CO)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBECRCB"/> | pDev->blecsBECRCB.set(pAcq->blecsBECRCB); |  |  | Number of CRC errors at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 1 |  | 16
BELFA | BmEngyLstFrmCntA | BEAM_ENGY_LOSTF_CNT_A | Beam ENERGY | 70003C | 700040 | 16..31 | unsigned | :16 | unsigned blecsBELFA:16; | <field name="blecsBELFA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBELFA"/> | pDev->blecsBELFA.set(pAcq->blecsBELFA); |  | Beam Energy 4 | Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 1 |  | 16 |  |  |  | 19 |  |  | if (vme_addr[22..0] >= h"70003C") and (vme_addr[22..0] < h"700040") then mux_sel[] = 19; end if; --Beam Energy 4 / comments: Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO)
BELFB | BmEngyLstFrmCntB | BEAM_ENGY_LOSTF_CNT_B | Beam ENERGY |  |  | 0..15 | unsigned | :16 | unsigned blecsBELFB:16; | <field name="blecsBELFB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Number of lost frames (packets) at the reception of the beam energy (channel B) from the CISV card (AB/CO)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBELFB"/> | pDev->blecsBELFB.set(pAcq->blecsBELFB); |  |  | Number of lost frames (packets) at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 1 |  | 16
BETOV | BmEngyTimeOut | BEAM_ENGY_TIME_OUT | Beam ENERGY | 700040 | 700044 | 28..31 | unsigned | :4 | unsigned blecsBETOV:4; | <field name="blecsBETOV" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBETOV"/> | pDev->blecsBETOV.set(pAcq->blecsBETOV); |  | Beam Energy 5 | Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO) | 1 |  | 16 |  |  |  | 20 |  |  | if (vme_addr[22..0] >= h"700040") and (vme_addr[22..0] < h"700044") then mux_sel[] = 20; end if; --Beam Energy 5 / comments: Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO)
|  |  | - |  |  | 0..27 | unsigned | :28 | unsigned blecsSPARE118:28; | [6 empty] | SPARE
V1H | HighV1VHigh | HV_1_VOLT_HIGH |  | 700044 | 700048 | 31 | bool | :1 | bool blecsV1H:1; | <field name="blecsV1H" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / voltage higher</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsV1H"/> | pDev->blecsV1H.set(pAcq->blecsV1H); |  | High Voltage 1 survey | High Voltage power supply 1 / voltage higher | 1 |  | 1 |  |  |  | 21 |  |  | if (vme_addr[22..0] >= h"700044") and (vme_addr[22..0] < h"700048") then mux_sel[] = 21; end if; --High Voltage 1 survey / comments: High Voltage power supply 1 / voltage higher
V1L | HighV1VLow | HV_1_VOLT_LOW |  |  |  | 30 | bool | :1 | bool blecsV1L:1; | <field name="blecsV1L" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / voltage lower</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsV1L"/> | pDev->blecsV1L.set(pAcq->blecsV1L); |  |  | High Voltage power supply 1 / voltage lower | 1 |  | 1
I1H | HighV1IHigh | HV_1_CURR_HIGH |  |  |  | 29 | bool | :1 | bool blecsI1H:1; | <field name="blecsI1H" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / current higher</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsI1H"/> | pDev->blecsI1H.set(pAcq->blecsI1H); |  |  | High Voltage power supply 1 / current higher | 1 |  | 1
I1L | HighV1ILow | HV_1_CURR_LOW |  |  |  | 28 | bool | :1 | bool blecsI1L:1; | <field name="blecsI1L" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / current lower</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsI1L"/> | pDev->blecsI1L.set(pAcq->blecsI1L); |  |  | High Voltage power supply 1 / current lower | 1 |  | 1
[6 empty] | 24..27 | unsigned | :4 | unsigned blecsSPARE123:4; | [6 empty] | SPARE
HV1V | HighV1VValue | HV_1_VOLT_VALUE |  |  |  | 0..23 | unsigned | :24 | unsigned blecsHV1V:24; | <field name="blecsHV1V" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / voltage value in [V] from 24 bits [HV = 3000/2**24*code]</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsHV1V"/> | pDev->blecsHV1V.set(pAcq->blecsHV1V); |  | High Voltage 2 survey | High Voltage power supply 1 / voltage value in [V] from 24 bits [HV = 3000/2**24*code] | 1 |  | 24 | use filter |  | 3000/2**24*code
V2H | HighV2VHigh | HV_2_VOLT_HIGH | Survey | 700048 | 70004C | 31 | bool | :1 | bool blecsV2H:1; | <field name="blecsV2H" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / voltage higher</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsV2H"/> | pDev->blecsV2H.set(pAcq->blecsV2H); |  |  | High Voltage power supply 2 / voltage higher | 1 |  | 1 |  |  |  | 22 |  |  | if (vme_addr[22..0] >= h"700048") and (vme_addr[22..0] < h"70004C") then mux_sel[] = 22; end if; -- / comments: High Voltage power supply 2 / voltage higher
V2L | HighV2VLow | HV_2_VOLT_LOW |  |  |  | 30 | bool | :1 | bool blecsV2L:1; | <field name="blecsV2L" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / voltage lower</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsV2L"/> | pDev->blecsV2L.set(pAcq->blecsV2L); |  |  | High Voltage power supply 2 / voltage lower | 1 |  | 1
I2H | HighV2IHigh | HV_2_CURR_HIGH |  |  |  | 29 | bool | :1 | bool blecsI2H:1; | <field name="blecsI2H" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / current higher</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsI2H"/> | pDev->blecsI2H.set(pAcq->blecsI2H); |  |  | High Voltage power supply 2 / current higher | 1 |  | 1
I2L | HighV2ILow | HV_2_CURR_LOW |  |  |  | 28 | bool | :1 | bool blecsI2L:1; | <field name="blecsI2L" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / current lower</description> <scalar type="bool"/> </field> |  | <data-field-ref-item field-name-ref="blecsI2L"/> | pDev->blecsI2L.set(pAcq->blecsI2L); |  |  | High Voltage power supply 2 / current lower | 1 |  | 1
[6 empty] | 24..27 | unsigned | :4 | unsigned blecsSPARE129:4; | [6 empty] | SPARE
HV2V | HighV2VValue | HV_2_VOLT_VALUE |  |  |  | 0..23 | unsigned | :24 | unsigned blecsHV2V:24; | <field name="blecsHV2V" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsHV2V"/> | pDev->blecsHV2V.set(pAcq->blecsHV2V); |  |  | High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 1 |  | 24 | use filter |  | 3000/2**24*code
HV1MAX | HighV1VMax |  | HV & LV | 70004C | 700050 | 16..31 | unsigned | :16 | unsigned blecsHV1MAX:16; | <field name="blecsHV1MAX" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Max Value of HV1 since last read of the CPU (16bits) [Hvmax = 3000/2**16*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV1MAX"/> |  | pDev->blecsHV1MAX.set(pAcq->blecsHV1MAX); | if (data.itemAvailable("blecsHV1MAX") == true)blecsAcq.blecsHV1MAX = data.blecsHV1MAX.get(); | High Voltage 1 MinMax | Max Value of HV1 since last read of the CPU (16bits) [Hvmax = 3000/2**16*code] in V | [5 empty] | 3000/2**16*code | 23 |  |  | if (vme_addr[22..0] >= h"70004C") and (vme_addr[22..0] < h"700050") then mux_sel[] = 23; end if; --High Voltage 1 MinMax / comments: Max Value of HV1 since last read of the CPU (16bits) [Hvmax = 3000/2**16*code] in V
HV1MIN | HighV1VMin |  | Control and |  |  | 0..15 | unsigned | :16 | unsigned blecsHV1MIN:16; | <field name="blecsHV1MIN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Min Value of HV1 since last read of the CPU (16bits) [Hvmax = 3000/2**16*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV1MIN"/> |  | pDev->blecsHV1MIN.set(pAcq->blecsHV1MIN); | if (data.itemAvailable("blecsHV1MIN") == true)blecsAcq.blecsHV1MIN = data.blecsHV1MIN.get(); |  | Min Value of HV1 since last read of the CPU (16bits) [Hvmax = 3000/2**16*code] in V | [5 empty] | 3000/2**16*code
HV2MAX | HighV2VMax |  |  | 700050 | 700054 | 16..31 | unsigned | :16 | unsigned blecsHV2MAX:16; | <field name="blecsHV2MAX" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Max Value of HV2 since last read of the CPU (16bits) [Hvmax = 3000/2**16*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV2MAX"/> |  | pDev->blecsHV2MAX.set(pAcq->blecsHV2MAX); | if (data.itemAvailable("blecsHV2MAX") == true)blecsAcq.blecsHV2MAX = data.blecsHV2MAX.get(); | High Voltage 2 MinMax | Max Value of HV2 since last read of the CPU (16bits) [Hvmax = 3000/2**16*code] in V | [5 empty] | 3000/2**16*code | 24 |  |  | if (vme_addr[22..0] >= h"700050") and (vme_addr[22..0] < h"700054") then mux_sel[] = 24; end if; --High Voltage 2 MinMax / comments: Max Value of HV2 since last read of the CPU (16bits) [Hvmax = 3000/2**16*code] in V
HV2MIN | HighV2VMin |  |  |  |  | 0..15 | unsigned | :16 | unsigned blecsHV2MIN:16; | <field name="blecsHV2MIN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Min Value of HV2 since last read of the CPU (16bits) [Hvmax = 3000/2**16*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV2MIN"/> |  | pDev->blecsHV2MIN.set(pAcq->blecsHV2MIN); | if (data.itemAvailable("blecsHV2MIN") == true)blecsAcq.blecsHV2MIN = data.blecsHV2MIN.get(); |  | Min Value of HV2 since last read of the CPU (16bits) [Hvmax = 3000/2**16*code] in V | [5 empty] | 3000/2**16*code
HV1I | HighV1IValue | HV_1_CURR_VALUE |  | 700054 | 700058 | 16..31 | unsigned | :16 | unsigned blecsHV1I:16; | <field name="blecsHV1I" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsHV1I"/> | pDev->blecsHV1I.set(pAcq->blecsHV1I); |  | High Voltage Currents | High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA | 1 |  | 16 | use filter |  | 20/2**16*code | 25 |  |  | if (vme_addr[22..0] >= h"700054") and (vme_addr[22..0] < h"700058") then mux_sel[] = 25; end if; --High Voltage Currents / comments: High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA
HV2I | HighV2IValue | HV_2_CURR_VALUE |  |  |  | 0..15 | unsigned | :16 | unsigned blecsHV2I:16; | <field name="blecsHV2I" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsHV2I"/> | pDev->blecsHV2I.set(pAcq->blecsHV2I); |  |  | High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA | 1 |  | 16 | use filter |  | 20/2**16*code
HV1IMAX | HighV1IMax |  |  | 700058 | 70005C | 16..31 | unsigned | :16 | unsigned blecsHV1IMAX:16; | <field name="blecsHV1IMAX" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Max Value of the HV1 current since last read of the CPU (16 bits) [I = 20/2**16*code] in mA</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV1IMAX"/> |  | pDev->blecsHV1IMAX.set(pAcq->blecsHV1IMAX); | if (data.itemAvailable("blecsHV1IMAX") == true)blecsAcq.blecsHV1IMAX = data.blecsHV1IMAX.get(); | High Voltage 1 Current MinMax | Max Value of the HV1 current since last read of the CPU (16 bits) [I = 20/2**16*code] in mA | [5 empty] | 20/2**16*code | 26 |  |  | if (vme_addr[22..0] >= h"700058") and (vme_addr[22..0] < h"70005C") then mux_sel[] = 26; end if; --High Voltage 1 Current MinMax / comments: Max Value of the HV1 current since last read of the CPU (16 bits) [I = 20/2**16*code] in mA
HV1IMIN | HighV1IMin |  |  |  |  | 0..15 | unsigned | :16 | unsigned blecsHV1IMIN:16; | <field name="blecsHV1IMIN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Min Value of the HV1 current since last read of the CPU (16 bits) [I = 20/2**16*code] in mA</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV1IMIN"/> |  | pDev->blecsHV1IMIN.set(pAcq->blecsHV1IMIN); | if (data.itemAvailable("blecsHV1IMIN") == true)blecsAcq.blecsHV1IMIN = data.blecsHV1IMIN.get(); |  | Min Value of the HV1 current since last read of the CPU (16 bits) [I = 20/2**16*code] in mA | [5 empty] | 20/2**16*code
HV2IMAX | HighV2IMax |  |  | 70005C | 700060 | 16..31 | unsigned | :16 | unsigned blecsHV2IMAX:16; | <field name="blecsHV2IMAX" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Max Value of the HV2 current since last read of the CPU (16 bits) [I = 20/2**16*code] in mA</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV2IMAX"/> |  | pDev->blecsHV2IMAX.set(pAcq->blecsHV2IMAX); | if (data.itemAvailable("blecsHV2IMAX") == true)blecsAcq.blecsHV2IMAX = data.blecsHV2IMAX.get(); | High Voltage 2 Current MinMax | Max Value of the HV2 current since last read of the CPU (16 bits) [I = 20/2**16*code] in mA | [5 empty] | 20/2**16*code | 27 |  |  | if (vme_addr[22..0] >= h"70005C") and (vme_addr[22..0] < h"700060") then mux_sel[] = 27; end if; --High Voltage 2 Current MinMax / comments: Max Value of the HV2 current since last read of the CPU (16 bits) [I = 20/2**16*code] in mA
HV2IMIN | HighV2IMin |  |  |  |  | 0..15 | unsigned | :16 | unsigned blecsHV2IMIN:16; | <field name="blecsHV2IMIN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Min Value of the HV2 current since last read of the CPU (16 bits) [I = 20/2**16*code] in mA</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV2IMIN"/> |  | pDev->blecsHV2IMIN.set(pAcq->blecsHV2IMIN); | if (data.itemAvailable("blecsHV2IMIN") == true)blecsAcq.blecsHV2IMIN = data.blecsHV2IMIN.get(); |  | Min Value of the HV2 current since last read of the CPU (16 bits) [I = 20/2**16*code] in mA | [5 empty] | 20/2**16*code
|  |  |  | 700060 | 700064 | 29..31 | unsigned | :3 | unsigned blecsSPARE141:3; | [5 empty] | VME Voltages 1 | SPARE |  |  |  |  | second |  | 28 |  |  | if (vme_addr[22..0] >= h"700060") and (vme_addr[22..0] < h"700064") then mux_sel[] = 28; end if; --VME Voltages 1 / comments: SPARE
VME3V3 | Vme3Volt3Value |  |  |  |  | 16..28 | unsigned | :13 | unsigned blecsVME3V3:13; | <field name="blecsVME3V3" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>VME 3V3 (13 bits) with 10V full scale</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsVME3V3"/> |  | pDev->blecsVME3V3.set(pAcq->blecsVME3V3); | if (data.itemAvailable("blecsVME3V3") == true)blecsAcq.blecsVME3V3 = data.blecsVME3V3.get(); |  | VME 3V3 (13 bits) with 10V full scale |  |  |  |  | every | 10/2**13*code
[6 empty] | 13..15 | unsigned | :3 | unsigned blecsSPARE143:3; | [6 empty] | SPARE |  |  |  |  | Read it
VME3V3DEL | Vme3Volt3Delta |  |  |  |  | 0..12 | unsigned | :13 | unsigned blecsVME3V3DEL:13; | <field name="blecsVME3V3DEL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Delta value of the VME3V3 => Vmax-Vmin [V=10/2**13*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsVME3V3DEL"/> |  | pDev->blecsVME3V3DEL.set(pAcq->blecsVME3V3DEL); | if (data.itemAvailable("blecsVME3V3DEL") == true)blecsAcq.blecsVME3V3DEL = data.blecsVME3V3DEL.get(); |  | Delta value of the VME3V3 => Vmax-Vmin [V=10/2**13*code] in V | [5 empty] | 10/2**13*code
|  |  |  | 700064 | 700068 | 29..31 | unsigned | :3 | unsigned blecsSPARE145:3; | [5 empty] | VME Voltages 2 | [7 empty] | 29 |  |  | if (vme_addr[22..0] >= h"700064") and (vme_addr[22..0] < h"700068") then mux_sel[] = 29; end if; --VME Voltages 2 / comments:
VME5V | Vme5VoltValue |  |  |  |  | 16..28 | unsigned | :13 | unsigned blecsVME5V:13; | <field name="blecsVME5V" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>VME 5V (13 bits) with 10V full scale</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsVME5V"/> |  | pDev->blecsVME5V.set(pAcq->blecsVME5V); | if (data.itemAvailable("blecsVME5V") == true)blecsAcq.blecsVME5V = data.blecsVME5V.get(); |  | VME 5V (13 bits) with 10V full scale | [5 empty] | 10/2**13*code
[6 empty] | 13..15 | unsigned | :3 | unsigned blecsSPARE147:3;
VME5VDEL | Vme5VoltDelta |  |  |  |  | 0..12 | unsigned | :13 | unsigned blecsVME5VDEL:13; | <field name="blecsVME5VDEL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Delta value of the VME5V => Vmax-Vmin [V=10/2**13*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsVME5VDEL"/> |  | pDev->blecsVME5VDEL.set(pAcq->blecsVME5VDEL); | if (data.itemAvailable("blecsVME5VDEL") == true)blecsAcq.blecsVME5VDEL = data.blecsVME5VDEL.get(); |  | Delta value of the VME5V => Vmax-Vmin [V=10/2**13*code] in V | [5 empty] | 10/2**13*code
|  |  |  | 700068 | 70006C | 29..31 | unsigned | :3 | unsigned blecsSPARE149:3; | [5 empty] | Analog Voltages 1 | SPARE | [6 empty] | 30 |  |  | if (vme_addr[22..0] >= h"700068") and (vme_addr[22..0] < h"70006C") then mux_sel[] = 30; end if; --Analog Voltages 1 / comments: SPARE
AV5V | Analog5VoltValue |  | Survey |  |  | 16..28 | unsigned | :13 | unsigned blecsAV5V:13; | <field name="blecsAV5V" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Analog 5V (13 bits) with 10V full scale</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsAV5V"/> |  | pDev->blecsAV5V.set(pAcq->blecsAV5V); | if (data.itemAvailable("blecsAV5V") == true)blecsAcq.blecsAV5V = data.blecsAV5V.get(); |  | Analog 5V (13 bits) with 10V full scale | [5 empty] | 10/2**13*code
|  |  | HV & LV |  |  | 13..15 | unsigned | :3 | unsigned blecsSPARE151:3; | [6 empty] | SPARE
AV5VDEL | Analog5VoltDelta |  | Control and |  |  | 0..12 | unsigned | :13 | unsigned blecsAV5VDEL:13; | <field name="blecsAV5VDEL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Delta value of the AV5V => Vmax-Vmin [V=10/2**13*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsAV5VDEL"/> |  | pDev->blecsAV5VDEL.set(pAcq->blecsAV5VDEL); | if (data.itemAvailable("blecsAV5VDEL") == true)blecsAcq.blecsAV5VDEL = data.blecsAV5VDEL.get(); |  | Delta value of the AV5V => Vmax-Vmin [V=10/2**13*code] in V | [5 empty] | 10/2**13*code
|  |  |  | 70006C | 700070 | 29..31 | unsigned | :3 | unsigned blecsSPARE153:3; | [5 empty] | Analog Voltages 2 | SPARE | [6 empty] | 31 |  |  | if (vme_addr[22..0] >= h"70006C") and (vme_addr[22..0] < h"700070") then mux_sel[] = 31; end if; --Analog Voltages 2 / comments: SPARE
AVP15V | AnalogP15VoltValue |  |  |  |  | 16..28 | unsigned | :13 | unsigned blecsAVP15V:13; | <field name="blecsAVP15V" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Analog P15V reference (13 bits) [V = 10/2**13*code*3] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsAVP15V"/> |  | pDev->blecsAVP15V.set(pAcq->blecsAVP15V); | if (data.itemAvailable("blecsAVP15V") == true)blecsAcq.blecsAVP15V = data.blecsAVP15V.get(); |  | Analog P15V reference (13 bits) [V = 10/2**13*code*3] in V | [5 empty] | 30/2**13*code
[6 empty] | 13..15 | unsigned | :3 | unsigned blecsSPARE155:3; | [6 empty] | SPARE
AVP15VDEL | AnalogP15VoltDelta |  |  |  |  | 0..12 | unsigned | :13 | unsigned blecsAVP15VDEL:13; | <field name="blecsAVP15VDEL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Delta value of the AP15V => Vmax-Vmin [V=30/2**13*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsAVP15VDEL"/> |  | pDev->blecsAVP15VDEL.set(pAcq->blecsAVP15VDEL); | if (data.itemAvailable("blecsAVP15VDEL") == true)blecsAcq.blecsAVP15VDEL = data.blecsAVP15VDEL.get(); |  | Delta value of the AP15V => Vmax-Vmin [V=30/2**13*code] in V | [5 empty] | 30/2**13*code
|  |  |  | 700070 | 700074 | 29..31 | unsigned | :3 | unsigned blecsSPARE157:3; | [5 empty] | Analog Voltages 3 | SPARE | [6 empty] | 32 |  |  | if (vme_addr[22..0] >= h"700070") and (vme_addr[22..0] < h"700074") then mux_sel[] = 32; end if; --Analog Voltages 3 / comments: SPARE
AVM15V | AnalogM15VoltValue |  |  |  |  | 16..28 | unsigned | :13 | unsigned blecsAVM15V:13; | <field name="blecsAVM15V" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Analog M15V reference (13 bits) [V = (20/2**13*code-10)*3] in V (result in signed val)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsAVM15V"/> |  | pDev->blecsAVM15V.set(pAcq->blecsAVM15V); | if (data.itemAvailable("blecsAVM15V") == true)blecsAcq.blecsAVM15V = data.blecsAVM15V.get(); |  | Analog M15V reference (13 bits) [V = (20/2**13*code-10)*3] in V (result in signed val) | [5 empty] | (20/2**13*code-10)*3
[6 empty] | 13..15 | unsigned | :3 | unsigned blecsSPARE159:3; | [6 empty] | SPARE
AVM15VDEL | AnalogM15VoltDelta |  |  |  |  | 0..12 | unsigned | :13 | unsigned blecsAVM15VDEL:13; | <field name="blecsAVM15VDEL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Delta value of the M15V => Vmin-Vmax [V=(20/2**13*code-10)*3] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsAVM15VDEL"/> |  | pDev->blecsAVM15VDEL.set(pAcq->blecsAVM15VDEL); | if (data.itemAvailable("blecsAVM15VDEL") == true)blecsAcq.blecsAVM15VDEL = data.blecsAVM15VDEL.get(); |  | Delta value of the M15V => Vmin-Vmax [V=(20/2**13*code-10)*3] in V | [5 empty] | (20/2**13*code-10)*3
|  |  |  | 700074 | 700078 | 30..31 | unsigned | :2 | unsigned blecsSPARE161:2; | [5 empty] | Analog Voltages 4 | SPARE | [6 empty] | 33 |  |  | if (vme_addr[22..0] >= h"700074") and (vme_addr[22..0] < h"700078") then mux_sel[] = 33; end if; --Analog Voltages 4 / comments: SPARE
AV5VREF | AnalogRef5VoltValue |  |  |  |  | 20..29 | unsigned | :10 | unsigned blecsAV5VREF:10; | <field name="blecsAV5VREF" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Analog 5V reference (10 bits) with [V=10/2**10*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsAV5VREF"/> |  | pDev->blecsAV5VREF.set(pAcq->blecsAV5VREF); | if (data.itemAvailable("blecsAV5VREF") == true)blecsAcq.blecsAV5VREF = data.blecsAV5VREF.get(); |  | Analog 5V reference (10 bits) with [V=10/2**10*code] in V | [5 empty] | 10/2**10*code
AV10VREFA | AnalogRefA10VoltValue |  |  |  |  | 10..19 | unsigned | :10 | unsigned blecsAV10VREFA:10; | <field name="blecsAV10VREFA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Analog 10V_A reference (10 bits) with [V=20/2**10*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsAV10VREFA"/> |  | pDev->blecsAV10VREFA.set(pAcq->blecsAV10VREFA); | if (data.itemAvailable("blecsAV10VREFA") == true)blecsAcq.blecsAV10VREFA = data.blecsAV10VREFA.get(); |  | Analog 10V_A reference (10 bits) with [V=20/2**10*code] in V | [5 empty] | 20/2**10*code
AV10VREFB | AnalogRefB10VoltValue |  |  |  |  | 0..9 | unsigned | :10 | unsigned blecsAV10VREFB:10; | <field name="blecsAV10VREFB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Analog 10V_B reference (10 bits) with [V=20/2**10*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsAV10VREFB"/> |  | pDev->blecsAV10VREFB.set(pAcq->blecsAV10VREFB); | if (data.itemAvailable("blecsAV10VREFB") == true)blecsAcq.blecsAV10VREFB = data.blecsAV10VREFB.get(); |  | Analog 10V_B reference (10 bits) with [V=20/2**10*code] in V | [5 empty] | 20/2**10*code
LV5A | Analog5VoltOk |  |  | 700078 | 70007C | 31 | bool | :1 | bool blecsLV5A:1; | <field name="blecsLV5A" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>LV Analog 5V ok</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsLV5A"/> |  | pDev->blecsLV5A.set(pAcq->blecsLV5A); | if (data.itemAvailable("blecsLV5A") == true)blecsAcq.blecsLV5A = data.blecsLV5A.get(); | Low Voltage survey 1 | LV Analog 5V ok | [6 empty] | 34 |  |  | if (vme_addr[22..0] >= h"700078") and (vme_addr[22..0] < h"70007C") then mux_sel[] = 34; end if; --Low Voltage survey 1 / comments: LV Analog 5V ok
LV5AC | Analog5VoltCnt |  |  |  |  | 0..30 | unsigned | :31 | unsigned blecsLV5AC:31; | <field name="blecsLV5AC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>LV Analog 5V Counter. When below level, count with 25ns period</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsLV5AC"/> |  | pDev->blecsLV5AC.set(pAcq->blecsLV5AC); | if (data.itemAvailable("blecsLV5AC") == true)blecsAcq.blecsLV5AC = data.blecsLV5AC.get(); |  | LV Analog 5V Counter. When below level, count with 25ns period
LV15A | Analog15VoltOk |  |  | 70007C | 700080 | 31 | bool | :1 | bool blecsLV15A:1; | <field name="blecsLV15A" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>LV Analog +-15V ok</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsLV15A"/> |  | pDev->blecsLV15A.set(pAcq->blecsLV15A); | if (data.itemAvailable("blecsLV15A") == true)blecsAcq.blecsLV15A = data.blecsLV15A.get(); | Low Voltage survey 2 | LV Analog +-15V ok | [6 empty] | 35 |  |  | if (vme_addr[22..0] >= h"70007C") and (vme_addr[22..0] < h"700080") then mux_sel[] = 35; end if; --Low Voltage survey 2 / comments: LV Analog +-15V ok
LV15AC | Analog15VoltCnt |  |  |  |  | 0..30 | unsigned | :31 | unsigned blecsLV15AC:31; | <field name="blecsLV15AC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>LV Analog +-15V Counter. When below or higher level, count with 25ns period</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsLV15AC"/> |  | pDev->blecsLV15AC.set(pAcq->blecsLV15AC); | if (data.itemAvailable("blecsLV15AC") == true)blecsAcq.blecsLV15AC = data.blecsLV15AC.get(); |  | LV Analog +-15V Counter. When below or higher level, count with 25ns period
LV3D3 | Vme3Volt3Ok |  |  | 700080 | 700084 | 31 | bool | :1 | bool blecsLV3D3:1; | <field name="blecsLV3D3" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>LV Digital 3.3V ok</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsLV3D3"/> |  | pDev->blecsLV3D3.set(pAcq->blecsLV3D3); | if (data.itemAvailable("blecsLV3D3") == true)blecsAcq.blecsLV3D3 = data.blecsLV3D3.get(); | Low Voltage survey 3 | LV Digital 3.3V ok | [6 empty] | 36 |  |  | if (vme_addr[22..0] >= h"700080") and (vme_addr[22..0] < h"700084") then mux_sel[] = 36; end if; --Low Voltage survey 3 / comments: LV Digital 3.3V ok
LV3D3C | Vme3Volt3Cnt |  |  |  |  | 0..30 | unsigned | :31 | unsigned blecsLV3D3C:31; | <field name="blecsLV3D3C" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>LV Digital 3.3V Counter. When below level, count with 25ns period</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsLV3D3C"/> |  | pDev->blecsLV3D3C.set(pAcq->blecsLV3D3C); | if (data.itemAvailable("blecsLV3D3C") == true)blecsAcq.blecsLV3D3C = data.blecsLV3D3C.get(); |  | LV Digital 3.3V Counter. When below level, count with 25ns period
LV5D | Vme5VoltOk |  |  | 700084 | 700088 | 31 | bool | :1 | bool blecsLV5D:1; | <field name="blecsLV5D" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>LV Digital 5V ok</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsLV5D"/> |  | pDev->blecsLV5D.set(pAcq->blecsLV5D); | if (data.itemAvailable("blecsLV5D") == true)blecsAcq.blecsLV5D = data.blecsLV5D.get(); | Low Voltage survey 4 | LV Digital 5V ok | [6 empty] | 37 |  |  | if (vme_addr[22..0] >= h"700084") and (vme_addr[22..0] < h"700088") then mux_sel[] = 37; end if; --Low Voltage survey 4 / comments: LV Digital 5V ok
LV5DC | Vme5VoltCnt |  |  |  |  | 0..30 | unsigned | :31 | unsigned blecsLV5DC:31; | <field name="blecsLV5DC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>LV Digital 5V Counter. When below level, count with 25ns period</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsLV5DC"/> |  | pDev->blecsLV5DC.set(pAcq->blecsLV5DC); | if (data.itemAvailable("blecsLV5DC") == true)blecsAcq.blecsLV5DC = data.blecsLV5DC.get(); |  | LV Digital 5V Counter. When below level, count with 25ns period
LV12D | Vme12VoltOk |  |  | 700088 | 70008C | 31 | bool | :1 | bool blecsLV12D:1; | <field name="blecsLV12D" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>LV Digital +-12V ok</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsLV12D"/> |  | pDev->blecsLV12D.set(pAcq->blecsLV12D); | if (data.itemAvailable("blecsLV12D") == true)blecsAcq.blecsLV12D = data.blecsLV12D.get(); | Low Voltage survey 5 | LV Digital +-12V ok | [6 empty] | 38 |  |  | if (vme_addr[22..0] >= h"700088") and (vme_addr[22..0] < h"70008C") then mux_sel[] = 38; end if; --Low Voltage survey 5 / comments: LV Digital +-12V ok
LV12DC | Vmr12VoltCnt |  |  |  |  | 0..30 | unsigned | :31 | unsigned blecsLV12DC:31; | <field name="blecsLV12DC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>LV Digital +-12V Counter. When below or higher level, count with 25ns period</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsLV12DC"/> |  | pDev->blecsLV12DC.set(pAcq->blecsLV12DC); | if (data.itemAvailable("blecsLV12DC") == true)blecsAcq.blecsLV12DC = data.blecsLV12DC.get(); |  | LV Digital +-12V Counter. When below or higher level, count with 25ns period
TCFUAB | TurnCntWhenfallUAB | TURN_CNT_WN_FALL_U_AORB | Beam PERMIT | 70008C | 700090 | 0..31 | unsigned | :32 | unsigned blecsTCFUAB:32; | <field name="blecsTCFUAB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Turn counter when last fall of UA or UB (32 bits)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsTCFUAB"/> | pDev->blecsTCFUAB.set(pAcq->blecsTCFUAB); |  | Turn counter when last fall of UA or UB | Turn counter when last fall of UA or UB (32 bits) | 1 |  | 32 |  |  |  | 39 |  |  | if (vme_addr[22..0] >= h"70008C") and (vme_addr[22..0] < h"700090") then mux_sel[] = 39; end if; --Turn counter when last fall of UA or UB / comments: Turn counter when last fall of UA or UB (32 bits)
TCFMAB | TurnCntWhenfallMAB | TURN_CNT_WN_FALL_M_AORB | Beam PERMIT | 700090 | 700094 | 0..31 | unsigned | :32 | unsigned blecsTCFMAB:32; | <field name="blecsTCFMAB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Turn counter when last fall of MA or MB (32 bits)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsTCFMAB"/> | pDev->blecsTCFMAB.set(pAcq->blecsTCFMAB); |  | Turn counter when last fall of MA or MB | Turn counter when last fall of MA or MB (32 bits) | 1 |  | 32 |  |  |  | 40 |  |  | if (vme_addr[22..0] >= h"700090") and (vme_addr[22..0] < h"700094") then mux_sel[] = 40; end if; --Turn counter when last fall of MA or MB / comments: Turn counter when last fall of MA or MB (32 bits)
|  |  |  | 700094 | 700098 | 24..31 | unsigned | :8 | unsigned blecsSPARE177:8; | [5 empty] | Bunch counters last fall of A or B | SPARE | [6 empty] | 41 |  |  | if (vme_addr[22..0] >= h"700094") and (vme_addr[22..0] < h"700098") then mux_sel[] = 41; end if; --Bunch counters last fall of A or B / comments: SPARE
BCFMAB | BunchCntWhenfallMAB | BUNCH_CNT_WN_FALL_M_AORB | Beam PERMIT |  |  | 12..23 | unsigned | :12 | unsigned blecsBCFMAB:12; | <field name="blecsBCFMAB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Bunch counter when last fall of MA or MB 12 bits counter</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBCFMAB"/> | pDev->blecsBCFMAB.set(pAcq->blecsBCFMAB); |  |  | Bunch counter when last fall of MA or MB 12 bits counter | 1 |  | 16
BCFUAB | BunchCntWhenfallUAB | BUNCH_CNT_WN_FALL_U_AORB | Beam PERMIT |  |  | 0..11 | unsigned | :12 | unsigned blecsBCFUAB:12; | <field name="blecsBCFUAB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Bunch counter when last fall of UA or UB 12 bits counter</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBCFUAB"/> | pDev->blecsBCFUAB.set(pAcq->blecsBCFUAB); |  |  | Bunch counter when last fall of UA or UB 12 bits counter | 1 |  | 16
FIRM_REV |  |  | OVERVIEW | 700098 | 70009C | 31..0 | unsigned | :32 | unsigned blecsFIRM_REV:32; | <field name="blecsFIRM_REV" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Firmware version of the FPGA</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFIRM_REV"/> |  | pDev->blecsFIRM_REV.set(pAcq->blecsFIRM_REV); | if (data.itemAvailable("blecsFIRM_REV") == true)blecsAcq.blecsFIRM_REV = data.blecsFIRM_REV.get(); |  | Firmware version of the FPGA | [6 empty] | 85 |  |  | if (vme_addr[22..0] >= h"700098") and (vme_addr[22..0] < h"70009C") then mux_sel[] = 85; end if; -- / comments: Firmware version of the FPGA
SERIAL_UPPER |  |  | OVERVIEW | 70009C | 7000A0 | 31..0 | unsigned | :32 | unsigned blecsSERIAL_UPPER:32; | <field name="blecsSERIAL_UPPER" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Upper part of the serial number of the board</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsSERIAL_UPPER"/> |  | pDev->blecsSERIAL_UPPER.set(pAcq->blecsSERIAL_UPPER); | if (data.itemAvailable("blecsSERIAL_UPPER") == true)blecsAcq.blecsSERIAL_UPPER = data.blecsSERIAL_UPPER.get(); |  | Upper part of the serial number of the board | [6 empty] | 87 |  |  | if (vme_addr[22..0] >= h"70009C") and (vme_addr[22..0] < h"7000A0") then mux_sel[] = 87; end if; -- / comments: Upper part of the serial number of the board
SERIAL_LOWER |  |  | OVERVIEW | 7000A0 | 7000A4 | 31..0 | unsigned | :32 | unsigned blecsSERIAL_LOWER:32; | <field name="blecsSERIAL_LOWER" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Lower part of serial number of the board</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsSERIAL_LOWER"/> |  | pDev->blecsSERIAL_LOWER.set(pAcq->blecsSERIAL_LOWER); | if (data.itemAvailable("blecsSERIAL_LOWER") == true)blecsAcq.blecsSERIAL_LOWER = data.blecsSERIAL_LOWER.get(); |  | Lower part of serial number of the board | [6 empty] | 88 |  |  | if (vme_addr[22..0] >= h"7000A0") and (vme_addr[22..0] < h"7000A4") then mux_sel[] = 88; end if; -- / comments: Lower part of serial number of the board
|  |  | - | 7000A4 | 7000FB |  | unsigned | [22] | unsigned blecsSPARE183[22]; | [6 empty] | SPARE
NRCPUD | NewCpuReadDoneCnt |  | OVERVIEW | 7000FC | 700100 | 0..30 | unsigned | :32 | unsigned blecsNRCPUD:32; | <field name="blecsNRCPUD" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Counter from 0 to 2**32-1. At startup the counter is "0", increase by 1 after each read of this address</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsNRCPUD"/> |  | pDev->blecsNRCPUD.set(pAcq->blecsNRCPUD); | if (data.itemAvailable("blecsNRCPUD") == true)blecsAcq.blecsNRCPUD = data.blecsNRCPUD.get(); | Counter for New Read by CPU Done | Counter from 0 to 2**32-1. At startup the counter is "0", increase by 1 after each read of this address | [6 empty] | 42 |  |  | if (vme_addr[22..0] >= h"7000FC") and (vme_addr[22..0] < h"700100") then mux_sel[] = 42; end if; --Counter for New Read by CPU Done / comments: Counter from 0 to 2**32-1. At startup the counter is "0", increase by 1 after each read of this address
|  |  |  | 700100 | 700200 |  | unsigned | [64] | unsigned blecsSPARE185[64];
TEST RESULTS
|  |  | - | 700200 | 700204 | 16..31 | unsigned | :16 | unsigned blecsSPARE187:16; | [5 empty] | General tests info TC Board present | SPARE | [6 empty] | 43 |  |  | if (vme_addr[22..0] >= h"700200") and (vme_addr[22..0] < h"700204") then mux_sel[] = 43; end if; --General tests info TC Board present / comments: SPARE
BLETCPRES | - | TEST_RES_BLETC_PRES | TESTS |  |  | 0..15 | unsigned | :16 | unsigned blecsBLETCPRES:16; | <field name="blecsBLETCPRES" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Vector of BLETC Board present in the crate (MSBit TC1 LSBit TC16)</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBLETCPRES"/> | pDev->blecsBLETCPRES.set(pAcq->blecsBLETCPRES); |  |  | Vector of BLETC Board present in the crate (MSBit TC1 LSBit TC16) | 1 |  | 16
BLECFPRES | - | TEST_RES_BLECF_PRES | BLECF TEST | 700204 | 700208 | 31..0 | unsigned | :32 | unsigned blecsBLECFPRES:32; | <field name="blecsBLECFPRES" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Vector of BLECF Board present in the crate (MSBit BLECF1(TC1) LSBit BLECF32 (TC16))</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsBLECFPRES"/> | pDev->blecsBLECFPRES.set(pAcq->blecsBLECFPRES); |  | General tests info BLECF Board present | Vector of BLECF Board present in the crate (MSBit BLECF1(TC1) LSBit BLECF32 (TC16)) | 1 |  | 32 |  |  |  | 44 |  |  | if (vme_addr[22..0] >= h"700204") and (vme_addr[22..0] < h"700208") then mux_sel[] = 44; end if; --General tests info BLECF Board present / comments: Vector of BLECF Board present in the crate (MSBit BLECF1(TC1) LSBit BLECF32 (TC16))
|  |  |  | 700208 | 700240 |  | unsigned | [14] | unsigned blecsSPARE190[14]; | [5 empty] | SPARE
TRTBLETC | - |  | TESTS | 700240 | 700244 | 16..31 | unsigned | :16 | unsigned blecsTRTBLETC:16; | <field name="blecsTRTBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the Treshold to BPL test per BLETC (MSBit TC1) 1=passed</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTRTBLETC"/> |  | pDev->blecsTRTBLETC.set(pAcq->blecsTRTBLETC); | if (data.itemAvailable("blecsTRTBLETC") == true)blecsAcq.blecsTRTBLETC = data.blecsTRTBLETC.get(); | Tests result per BLETC 1 | Result of the Treshold to BPL test per BLETC (MSBit TC1) 1=passed |  |  |  |  | the concentrator |  | 45 |  |  | if (vme_addr[22..0] >= h"700240") and (vme_addr[22..0] < h"700244") then mux_sel[] = 45; end if; --Tests result per BLETC 1 / comments: Result of the Treshold to BPL test per BLETC (MSBit TC1) 1=passed
TRCBLETC | - | TEST_RES_CONSIST_PER_BLETC | TESTS |  |  | 0..15 | unsigned | :16 | unsigned blecsTRCBLETC:16; | <field name="blecsTRCBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the Consistency test per BLETC (MSBit TC1) 1=passed</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsTRCBLETC"/> | pDev->blecsTRCBLETC.set(pAcq->blecsTRCBLETC); |  |  | Result of the Consistency test per BLETC (MSBit TC1) 1=passed | 1 |  | 16 |  | when TRR=1
TRBPUBLETC | - | TEST_RES_BP_TC_PER_BLETC_U | TESTS | 700244 | 700248 | 16..31 | unsigned | :16 | unsigned blecsTRBPUBLETC:16; | <field name="blecsTRBPUBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsTRBPUBLETC"/> | pDev->blecsTRBPUBLETC.set(pAcq->blecsTRBPUBLETC); |  | Tests result per BLETC 2 | Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed | 1 |  | 16 |  | specific location |  | 46 |  |  | if (vme_addr[22..0] >= h"700244") and (vme_addr[22..0] < h"700248") then mux_sel[] = 46; end if; --Tests result per BLETC 2 / comments: Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed
TRBPMBLETC | - | TEST_RES_BP_TC_PER_BLETC_M | TESTS |  |  | 0..15 | unsigned | :16 | unsigned blecsTRBPMBLETC:16; | <field name="blecsTRBPMBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the BPLTC test maskable line (MSBit TC1) 1=passed</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsTRBPMBLETC"/> | pDev->blecsTRBPMBLETC.set(pAcq->blecsTRBPMBLETC); |  |  | Result of the BPLTC test maskable line (MSBit TC1) 1=passed | 1 |  | 16 |  | write it into
|  |  | - | 700248 | 70024C | 16..31 | unsigned | :16 | unsigned blecsSPARE195:16; | [5 empty] | Tests result per BLETC 3 | SPARE |  |  |  |  | don’t save it |  | 47 |  |  | if (vme_addr[22..0] >= h"700248") and (vme_addr[22..0] < h"70024C") then mux_sel[] = 47; end if; --Tests result per BLETC 3 / comments: SPARE
TRHVBLETC | - | TEST_RES_LF_MOD_PER_BLETC | TESTS |  |  | 0..15 | unsigned | :16 | unsigned blecsTRHVBLETC:16; | <field name="blecsTRHVBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsTRHVBLETC"/> | pDev->blecsTRHVBLETC.set(pAcq->blecsTRHVBLETC); |  |  | Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed | 1 |  | 16 |  | otherwise
HVCFCRES | - | TEST_RES_CFCTEST_PER_BLECF | BLECF TEST | 70024C | 700250 | 31..0 | unsigned | :32 | unsigned blecsHVCFCRES:32; | <field name="blecsHVCFCRES" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed</description> <scalar type="long"/> </field> |  | <data-field-ref-item field-name-ref="blecsHVCFCRES"/> | pDev->blecsHVCFCRES.set(pAcq->blecsHVCFCRES); |  | HVCFC result per BLECF | Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed | 1 |  | 32 |  |  |  | 48 |  |  | if (vme_addr[22..0] >= h"70024C") and (vme_addr[22..0] < h"700250") then mux_sel[] = 48; end if; --HVCFC result per BLECF / comments: Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed
|  |  | - | 700250 | 700300 |  | unsigned | [44] | unsigned blecsSPARE198[44]; | [5 empty] | SPARE
HVLFRESSETUP1 | - | TEST_RES_LF_MOD_STP_1 | HVLF TEST | 700300 | 700700 |  | unsigned | [256] | unsigned blecsHVLFRESSETUP1[256]; | <field name="blecsHVLFRESSETUP1" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the HVLF modulation test per channel, setup 1 (one DWORD per Channel)</description> <array type="long"><dim>256</dim></array> </field> | <data-field-ref-item field-name-ref="blecsHVLFRESSETUP1"/> |  | pDev->blecsHVLFRESSETUP1.set((const long int*)pAcq->blecsHVLFRESSETUP1,256); | if (data.itemAvailable("blecsHVLFRESSETUP1") == true)memcpy(blecsAcq.blecsHVLFRESSETUP1,data.blecsHVLFRESSETUP1.get(sz1),sizeof(blecsAcq.blecsHVLFRESSETUP1)); | HvlfResultSetup1[Nbr of Channels] | Result of the HVLF modulation test per channel, setup 1 (one DWORD per Channel) |  |  | 8192 |  |  |  | 49 | struct HvlfResult { |  | if (vme_addr[22..0] >= h"700300") and (vme_addr[22..0] < h"700700") then mux_sel[] = 49; end if; --HvlfResultSetup1[Nbr of Channels] / comments: Result of the HVLF modulation test per channel, setup 1 (one DWORD per Channel)
HVLFRESSETUP2 | - | TEST_RES_LF_MOD_STP_2 | HVLF TEST | 700700 | 700B00 |  | unsigned | [256] | unsigned blecsHVLFRESSETUP2[256]; | <field name="blecsHVLFRESSETUP2" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the HVLF modulation test per channel, setup 2 (one DWORD per Channel)</description> <array type="long"><dim>256</dim></array> </field> | <data-field-ref-item field-name-ref="blecsHVLFRESSETUP2"/> |  | pDev->blecsHVLFRESSETUP2.set((const long int*)pAcq->blecsHVLFRESSETUP2,256); | if (data.itemAvailable("blecsHVLFRESSETUP2") == true)memcpy(blecsAcq.blecsHVLFRESSETUP2,data.blecsHVLFRESSETUP2.get(sz1),sizeof(blecsAcq.blecsHVLFRESSETUP2)); | HvlfResultSetup2[Nbr of Channels] | Result of the HVLF modulation test per channel, setup 2 (one DWORD per Channel) |  |  | 8192 |  | second |  | 50 | bool HVLFPRES : 1; // Channel is connected ? | Monitor Status 1=present 0=not present | if (vme_addr[22..0] >= h"700700") and (vme_addr[22..0] < h"700B00") then mux_sel[] = 50; end if; --HvlfResultSetup2[Nbr of Channels] / comments: Result of the HVLF modulation test per channel, setup 2 (one DWORD per Channel)
[21 empty] | every |  |  | bool HVLFRES : 1; // Result | test Result parameter 1=Passed 0=Failed
[21 empty] | Read it |  |  | unsigned HVLFGAIN : 15; // Gain | Amplitude parameter (15 bits)
[24 empty] | unsigned HVLFPH : 15; // phase | Phase parameter (15 bits)
};
|  |  |  | 700B04 | 700C00 |  | unsigned | [64] | unsigned blecsSPARE205[64];
FLASH IMAGE
|  |  | FLASH | 700C00 | 700C04 | 31..16 | unsigned | :16 | unsigned blecsSPARE207:16; | [5 empty] | Present TC board table | SPARE | [6 empty] | 43 |  |  | if (vme_addr[22..0] >= h"700C00") and (vme_addr[22..0] < h"700C04") then mux_sel[] = 43; end if; --Present TC board table / comments: SPARE
FBLETCPRES | - |  | FLASH |  |  | 16..0 | unsigned | :16 | unsigned blecsFBLETCPRES:16; | <field name="blecsFBLETCPRES" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Present BLETC, MSBit=BLETC-1, LSBit=BLETC-16, 1=present, 0=no board </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFBLETCPRES"/> |  | pDev->blecsFBLETCPRES.set(pAcq->blecsFBLETCPRES); | if (data.itemAvailable("blecsFBLETCPRES") == true)blecsAcq.blecsFBLETCPRES = data.blecsFBLETCPRES.get(); |  | Present BLETC, MSBit=BLETC-1, LSBit=BLETC-16, 1=present, 0=no board
FBLECFPRES | - |  | FLASH | 700C04 | 700C08 | 31..0 | unsigned | :32 | unsigned blecsFBLECFPRES:32; | <field name="blecsFBLECFPRES" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Present BLECF, MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16), 1=Present 0=no board</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFBLECFPRES"/> |  | pDev->blecsFBLECFPRES.set(pAcq->blecsFBLECFPRES); | if (data.itemAvailable("blecsFBLECFPRES") == true)blecsAcq.blecsFBLECFPRES = data.blecsFBLECFPRES.get(); | Present BLECF board table | Present BLECF, MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16), 1=Present 0=no board | [6 empty] | 44 |  |  | if (vme_addr[22..0] >= h"700C04") and (vme_addr[22..0] < h"700C08") then mux_sel[] = 44; end if; --Present BLECF board table / comments: Present BLECF, MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16), 1=Present 0=no board
FCHPRES | - |  | FLASH | 700C08 | 700C0C | 31..0 | unsigned | [8] | unsigned blecsFCHPRES[8]; | <field name="blecsFCHPRES" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>TC 1 and TC 2 Present Channels (MSBit=CH1, LSBit=CH32)</description> <array type="long"><dim>8</dim></array> </field> | <data-field-ref-item field-name-ref="blecsFCHPRES"/> |  | pDev->blecsFCHPRES.set((const long int*)pAcq->blecsFCHPRES,8); | if (data.itemAvailable("blecsFCHPRES") == true)memcpy(blecsAcq.blecsFCHPRES,data.blecsFCHPRES.get(sz1),sizeof(blecsAcq.blecsFCHPRES)); | Present channels table[0] | TC 1 and TC 2 Present Channels (MSBit=CH1, LSBit=CH32) | [6 empty] | 52 |  |  | if (vme_addr[22..0] >= h"700C08") and (vme_addr[22..0] < h"700C0C") then mux_sel[] = 52; end if; --Present channels table[0] / comments: TC 1 and TC 2 Present Channels (MSBit=CH1, LSBit=CH32)
|  |  | FLASH | 700C0C | 700C10 | 31..0 | [8 empty] | Present channels table[1] | TC 3 and TC 4 (MSBit=CH33, LSBit=CH65)
|  |  | FLASH | 700C10 | 700C14 | 31..0 | [8 empty] | Present channels table[2] | TC 5 and TC 6 ….
|  |  | FLASH | 700C14 | 700C18 | 31..0 | [8 empty] | Present channels table[3] | TC 7 and TC 8 '1'= Channel connected to a monitor, '0'=not connected |  |  |  |  | From
|  |  | FLASH | 700C18 | 700C1C | 31..0 | [8 empty] | Present channels table[4] | TC 9 and TC 10 …. |  |  |  |  | LSA
|  |  | FLASH | 700C1C | 700C20 | 31..0 | [8 empty] | Present channels table[5] | TC 11 and TC 12 …. |  |  |  |  | DB
|  |  | FLASH | 700C20 | 700C24 | 31..0 | [8 empty] | Present channels table[6] | TC 13 and TC 14 ….
|  |  | FLASH | 700C24 | 700C28 | 31..0 | [8 empty] | Present channels table[7] | TC 15 and TC 16 (MSBit=CH224, LSBit=CH256)
FSHVNO | HV normal operation |  | FLASH | 700C28 | 700C2C | 31..16 | unsigned | :16 | unsigned blecsFSHVNO:16; | <field name="blecsFSHVNO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage Normal Operation in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFSHVNO"/> |  | pDev->blecsFSHVNO.set(pAcq->blecsFSHVNO); | if (data.itemAvailable("blecsFSHVNO") == true)blecsAcq.blecsFSHVNO = data.blecsFSHVNO.get(); | HVCommunSetup 1 | Set High Voltage Normal Operation in dac value | [6 empty] | 53 |  |  | if (vme_addr[22..0] >= h"700C28") and (vme_addr[22..0] < h"700C2C") then mux_sel[] = 53; end if; --HVCommunSetup 1 / comments: Set High Voltage Normal Operation in dac value
FSHVDIF | HV1 - HV2 normal operation |  | FLASH |  |  | 15..0 | unsigned | :16 | unsigned blecsFSHVDIF:16; | <field name="blecsFSHVDIF" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set Difference between HV1 and HV2 (HV2 lower than HV1) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFSHVDIF"/> |  | pDev->blecsFSHVDIF.set(pAcq->blecsFSHVDIF); | if (data.itemAvailable("blecsFSHVDIF") == true)blecsAcq.blecsFSHVDIF = data.blecsFSHVDIF.get(); |  | Set Difference between HV1 and HV2 (HV2 lower than HV1) in dac value
FSHVCFC | HV for HVCFC test |  | FLASH | 700C2C | 700C30 | 31..16 | unsigned | :16 | unsigned blecsFSHVCFC:16; | <field name="blecsFSHVCFC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage HVCFC (100 pA test) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFSHVCFC"/> |  | pDev->blecsFSHVCFC.set(pAcq->blecsFSHVCFC); | if (data.itemAvailable("blecsFSHVCFC") == true)blecsAcq.blecsFSHVCFC = data.blecsFSHVCFC.get(); | HVCommunSetup 2 | Set High Voltage HVCFC (100 pA test) in dac value |  |  |  |  | dabinit into the |  | 54 |  |  | if (vme_addr[22..0] >= h"700C2C") and (vme_addr[22..0] < h"700C30") then mux_sel[] = 54; end if; --HVCommunSetup 2 / comments: Set High Voltage HVCFC (100 pA test) in dac value
FHVLFSHV | HV for Modulation test |  | FLASH |  |  | 15..0 | unsigned | :16 | unsigned blecsFHVLFSHV:16; | <field name="blecsFHVLFSHV" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage for modulation test</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFHVLFSHV"/> |  | pDev->blecsFHVLFSHV.set(pAcq->blecsFHVLFSHV); | if (data.itemAvailable("blecsFHVLFSHV") == true)blecsAcq.blecsFHVLFSHV = data.blecsFHVLFSHV.get(); |  | Set High Voltage for modulation test |  |  |  |  | Write it with
FSHVGOH | HV for RGOH reset |  | FLASH | 700C30 | 700C34 | 31..16 | unsigned | :16 | unsigned blecsFSHVGOH:16; | <field name="blecsFSHVGOH" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage HVGOH (GOH Reset) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFSHVGOH"/> |  | pDev->blecsFSHVGOH.set(pAcq->blecsFSHVGOH); | if (data.itemAvailable("blecsFSHVGOH") == true)blecsAcq.blecsFSHVGOH = data.blecsFSHVGOH.get(); | HVCommunSetup 3 | Set High Voltage HVGOH (GOH Reset) in dac value | [6 empty] | 55 |  |  | if (vme_addr[22..0] >= h"700C30") and (vme_addr[22..0] < h"700C34") then mux_sel[] = 55; end if; --HVCommunSetup 3 / comments: Set High Voltage HVGOH (GOH Reset) in dac value
FSHVRDAC | HV for RDAC reset |  | FLASH |  |  | 15..0 | unsigned | :16 | unsigned blecsFSHVRDAC:16; | <field name="blecsFSHVRDAC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage HVRDAC (Dac Reset) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFSHVRDAC"/> |  | pDev->blecsFSHVRDAC.set(pAcq->blecsFSHVRDAC); | if (data.itemAvailable("blecsFSHVRDAC") == true)blecsAcq.blecsFSHVRDAC = data.blecsFSHVRDAC.get(); |  | Set High Voltage HVRDAC (Dac Reset) in dac value |  |  |  |  | FLASH
FSHVRDACP | HV peak for RDAC reset |  | FLASH | 700C34 | 700C38 | 31..16 | unsigned | :16 | unsigned blecsFSHVRDACP:16; | <field name="blecsFSHVRDACP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage (Peek Value) HVRDAC (Dac Reset) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFSHVRDACP"/> |  | pDev->blecsFSHVRDACP.set(pAcq->blecsFSHVRDACP); | if (data.itemAvailable("blecsFSHVRDACP") == true)blecsAcq.blecsFSHVRDACP = data.blecsFSHVRDACP.get(); | HVCommunSetup 4 | Set High Voltage (Peek Value) HVRDAC (Dac Reset) in dac value |  |  |  |  | addresses |  | 56 |  |  | if (vme_addr[22..0] >= h"700C34") and (vme_addr[22..0] < h"700C38") then mux_sel[] = 56; end if; --HVCommunSetup 4 / comments: Set High Voltage (Peek Value) HVRDAC (Dac Reset) in dac value
FSHVCFCP | HV peak for HVCFC test |  | FLASH |  |  | 15..0 | unsigned | :16 | unsigned blecsFSHVCFCP:16; | <field name="blecsFSHVCFCP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage (Peek Value) HVCFC (100 pA test) in dac value </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFSHVCFCP"/> |  | pDev->blecsFSHVCFCP.set(pAcq->blecsFSHVCFCP); | if (data.itemAvailable("blecsFSHVCFCP") == true)blecsAcq.blecsFSHVCFCP = data.blecsFSHVCFCP.get(); |  | Set High Voltage (Peek Value) HVCFC (100 pA test) in dac value |  |  |  |  | Only read at these
|  |  | FLASH | 700C38 | 700C3C | 31..16 | unsigned | :16 | unsigned blecsSPARE226:16; | [5 empty] | HVCommunSetup 5 | SPARE | [6 empty] | 57 |  |  | if (vme_addr[22..0] >= h"700C38") and (vme_addr[22..0] < h"700C3C") then mux_sel[] = 57; end if; --HVCommunSetup 5 / comments: SPARE
FSHVRGOHP | HV peak for RGOH reset |  | FLASH |  |  | 15..0 | unsigned | :16 | unsigned blecsFSHVRGOHP:16; | <field name="blecsFSHVRGOHP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage (Peek Value) HVGOH (GOH Reset) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFSHVRGOHP"/> |  | pDev->blecsFSHVRGOHP.set(pAcq->blecsFSHVRGOHP); | if (data.itemAvailable("blecsFSHVRGOHP") == true)blecsAcq.blecsFSHVRGOHP = data.blecsFSHVRGOHP.get(); |  | Set High Voltage (Peek Value) HVGOH (GOH Reset) in dac value
FHVLFP1 | Modulation 1 Period |  | FLASH | 700C3C | 700C40 | 31..16 | unsigned | :16 | unsigned blecsFHVLFP1:16; | <field name="blecsFHVLFP1" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Modulation Period (Period[s] = 2048/10**6)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFHVLFP1"/> |  | pDev->blecsFHVLFP1.set(pAcq->blecsFHVLFP1); | if (data.itemAvailable("blecsFHVLFP1") == true)blecsAcq.blecsFHVLFP1 = data.blecsFHVLFP1.get(); | ModulationCommunSetup 1 | Modulation Period (Period[s] = 2048/10**6) | [6 empty] | 58 |  |  | if (vme_addr[22..0] >= h"700C3C") and (vme_addr[22..0] < h"700C40") then mux_sel[] = 58; end if; --ModulationCommunSetup 1 / comments: Modulation Period (Period[s] = 2048/10**6)
FHVLFV1 | Modulation 1 Voltage |  | FLASH |  |  | 15..0 | unsigned | :16 | unsigned blecsFHVLFV1:16; | <field name="blecsFHVLFV1" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set modulation peak-peak voltage (Upp[V] = 11.72/300*Value)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFHVLFV1"/> |  | pDev->blecsFHVLFV1.set(pAcq->blecsFHVLFV1); | if (data.itemAvailable("blecsFHVLFV1") == true)blecsAcq.blecsFHVLFV1 = data.blecsFHVLFV1.get(); |  | Set modulation peak-peak voltage (Upp[V] = 11.72/300*Value)
FHVLFP2 | Modulation 2 Voltage |  | FLASH | 700C40 | 700C44 | 31..16 | unsigned | :16 | unsigned blecsFHVLFP2:16; | <field name="blecsFHVLFP2" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Modulation Period [Period[s] = 2048/10**6]</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFHVLFP2"/> |  | pDev->blecsFHVLFP2.set(pAcq->blecsFHVLFP2); | if (data.itemAvailable("blecsFHVLFP2") == true)blecsAcq.blecsFHVLFP2 = data.blecsFHVLFP2.get(); | ModulationCommunSetup 2 | Modulation Period [Period[s] = 2048/10**6] | [6 empty] | 59 |  |  | if (vme_addr[22..0] >= h"700C40") and (vme_addr[22..0] < h"700C44") then mux_sel[] = 59; end if; --ModulationCommunSetup 2 / comments: Modulation Period [Period[s] = 2048/10**6]
FHVLFV2 | Modulation 2 Period |  | FLASH |  |  | 15..0 | unsigned | :16 | unsigned blecsFHVLFV2:16; | <field name="blecsFHVLFV2" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set modulation peak-peak voltage [Upp = 11.72/300*Value] </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFHVLFV2"/> |  | pDev->blecsFHVLFV2.set(pAcq->blecsFHVLFV2); | if (data.itemAvailable("blecsFHVLFV2") == true)blecsAcq.blecsFHVLFV2 = data.blecsFHVLFV2.get(); |  | Set modulation peak-peak voltage [Upp = 11.72/300*Value]
FFIRM_REV |  |  | FLASH | 700C44 | 700C48 | 31..0 | unsigned | :32 | unsigned blecsFFIRM_REV:32; | <field name="blecsFFIRM_REV" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Firmware revison</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFFIRM_REV"/> |  | pDev->blecsFFIRM_REV.set(pAcq->blecsFFIRM_REV); | if (data.itemAvailable("blecsFFIRM_REV") == true)blecsAcq.blecsFFIRM_REV = data.blecsFFIRM_REV.get(); |  | Firmware revison | [6 empty] | 60 | 60 |  | if (vme_addr[22..0] >= h"700C44") and (vme_addr[22..0] < h"700C48") then mux_sel[] = 60; end if; -- / comments: Firmware revison
FSERIAL_UPPER |  |  | FLASH | 700C48 | 700C4C | 31..0 | unsigned | :32 | unsigned blecsFSERIAL_UPPER:32; | <field name="blecsFSERIAL_UPPER" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Serial number upper DWORD</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFSERIAL_UPPER"/> |  | pDev->blecsFSERIAL_UPPER.set(pAcq->blecsFSERIAL_UPPER); | if (data.itemAvailable("blecsFSERIAL_UPPER") == true)blecsAcq.blecsFSERIAL_UPPER = data.blecsFSERIAL_UPPER.get(); |  | Serial number upper DWORD | [6 empty] | 61 | 61 |  | if (vme_addr[22..0] >= h"700C48") and (vme_addr[22..0] < h"700C4C") then mux_sel[] = 61; end if; -- / comments: Serial number upper DWORD
FSERIAL_LOWER |  |  | FLASH | 700C4C | 700C50 | 31..0 | unsigned | :32 | unsigned blecsFSERIAL_LOWER:32; | <field name="blecsFSERIAL_LOWER" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Serial number lower DWORD</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFSERIAL_LOWER"/> |  | pDev->blecsFSERIAL_LOWER.set(pAcq->blecsFSERIAL_LOWER); | if (data.itemAvailable("blecsFSERIAL_LOWER") == true)blecsAcq.blecsFSERIAL_LOWER = data.blecsFSERIAL_LOWER.get(); |  | Serial number lower DWORD | [6 empty] | 62 | 62 |  | if (vme_addr[22..0] >= h"700C4C") and (vme_addr[22..0] < h"700C50") then mux_sel[] = 62; end if; -- / comments: Serial number lower DWORD
MOD_1_GMAX |  |  | FLASH | 700C50 | 700E50 | 31..0 | unsigned | [128] | unsigned blecsMOD_1_GMAX[128]; | <field name="blecsMOD_1_GMAX" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold value for the modulation test, maximum gain for the setup parameters 1</description> <array type="long"><dim>128</dim></array> </field> | <data-field-ref-item field-name-ref="blecsMOD_1_GMAX"/> |  | pDev->blecsMOD_1_GMAX.set((const long int*)pAcq->blecsMOD_1_GMAX,128); | if (data.itemAvailable("blecsMOD_1_GMAX") == true)memcpy(blecsAcq.blecsMOD_1_GMAX,data.blecsMOD_1_GMAX.get(sz1),sizeof(blecsAcq.blecsMOD_1_GMAX)); | Modulation setup 1 maximum gain | Threshold value for the modulation test, maximum gain for the setup parameters 1 | [6 empty] | 63 | 63 |  | if (vme_addr[22..0] >= h"700C50") and (vme_addr[22..0] < h"700E50") then mux_sel[] = 63; end if; --Modulation setup 1 maximum gain / comments: Threshold value for the modulation test, maximum gain for the setup parameters 1
MOD_1_GMIN |  |  | FLASH | 700E50 | 701050 | 31..0 | unsigned | [128] | unsigned blecsMOD_1_GMIN[128]; | <field name="blecsMOD_1_GMIN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold value for the modulation test, minimum gain for the setup parameters 1</description> <array type="long"><dim>128</dim></array> </field> | <data-field-ref-item field-name-ref="blecsMOD_1_GMIN"/> |  | pDev->blecsMOD_1_GMIN.set((const long int*)pAcq->blecsMOD_1_GMIN,128); | if (data.itemAvailable("blecsMOD_1_GMIN") == true)memcpy(blecsAcq.blecsMOD_1_GMIN,data.blecsMOD_1_GMIN.get(sz1),sizeof(blecsAcq.blecsMOD_1_GMIN)); | Modulation setup 1 minimum gain | Threshold value for the modulation test, minimum gain for the setup parameters 1 | [6 empty] | 64 | 64 |  | if (vme_addr[22..0] >= h"700E50") and (vme_addr[22..0] < h"701050") then mux_sel[] = 64; end if; --Modulation setup 1 minimum gain / comments: Threshold value for the modulation test, minimum gain for the setup parameters 1
MOD_1_PMAX |  |  | FLASH | 701050 | 701250 | 31..0 | unsigned | [128] | unsigned blecsMOD_1_PMAX[128]; | <field name="blecsMOD_1_PMAX" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold value for the modulation test, maximum phase for the setup parameters 1</description> <array type="long"><dim>128</dim></array> </field> | <data-field-ref-item field-name-ref="blecsMOD_1_PMAX"/> |  | pDev->blecsMOD_1_PMAX.set((const long int*)pAcq->blecsMOD_1_PMAX,128); | if (data.itemAvailable("blecsMOD_1_PMAX") == true)memcpy(blecsAcq.blecsMOD_1_PMAX,data.blecsMOD_1_PMAX.get(sz1),sizeof(blecsAcq.blecsMOD_1_PMAX)); | modulation setup 1 maximum phase | Threshold value for the modulation test, maximum phase for the setup parameters 1 | [6 empty] | 65 | 65 |  | if (vme_addr[22..0] >= h"701050") and (vme_addr[22..0] < h"701250") then mux_sel[] = 65; end if; --modulation setup 1 maximum phase / comments: Threshold value for the modulation test, maximum phase for the setup parameters 1
MOD_1_PMIN |  |  | FLASH | 701250 | 701450 | 31..0 | unsigned | [128] | unsigned blecsMOD_1_PMIN[128]; | <field name="blecsMOD_1_PMIN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold value for the modulation test, minimum phase for the setup parameters 1</description> <array type="long"><dim>128</dim></array> </field> | <data-field-ref-item field-name-ref="blecsMOD_1_PMIN"/> |  | pDev->blecsMOD_1_PMIN.set((const long int*)pAcq->blecsMOD_1_PMIN,128); | if (data.itemAvailable("blecsMOD_1_PMIN") == true)memcpy(blecsAcq.blecsMOD_1_PMIN,data.blecsMOD_1_PMIN.get(sz1),sizeof(blecsAcq.blecsMOD_1_PMIN)); | Modulation setup 1 minimum phase | Threshold value for the modulation test, minimum phase for the setup parameters 1 | [6 empty] | 66 | 66 |  | if (vme_addr[22..0] >= h"701250") and (vme_addr[22..0] < h"701450") then mux_sel[] = 66; end if; --Modulation setup 1 minimum phase / comments: Threshold value for the modulation test, minimum phase for the setup parameters 1
MOD_2_GMAX |  |  | FLASH | 701450 | 701650 | 31..0 | unsigned | [128] | unsigned blecsMOD_2_GMAX[128]; | <field name="blecsMOD_2_GMAX" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold value for the modulation test, maximum gain for the setup parameters 2</description> <array type="long"><dim>128</dim></array> </field> | <data-field-ref-item field-name-ref="blecsMOD_2_GMAX"/> |  | pDev->blecsMOD_2_GMAX.set((const long int*)pAcq->blecsMOD_2_GMAX,128); | if (data.itemAvailable("blecsMOD_2_GMAX") == true)memcpy(blecsAcq.blecsMOD_2_GMAX,data.blecsMOD_2_GMAX.get(sz1),sizeof(blecsAcq.blecsMOD_2_GMAX)); | Modulation setup 2 maximum gain | Threshold value for the modulation test, maximum gain for the setup parameters 2 | [6 empty] | 67 | 67 |  | if (vme_addr[22..0] >= h"701450") and (vme_addr[22..0] < h"701650") then mux_sel[] = 67; end if; --Modulation setup 2 maximum gain / comments: Threshold value for the modulation test, maximum gain for the setup parameters 2
MOD_2_GMIN |  |  | FLASH | 701650 | 701850 | 31..0 | unsigned | [128] | unsigned blecsMOD_2_GMIN[128]; | <field name="blecsMOD_2_GMIN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold value for the modulation test, minimum gain for the setup parameters 2</description> <array type="long"><dim>128</dim></array> </field> | <data-field-ref-item field-name-ref="blecsMOD_2_GMIN"/> |  | pDev->blecsMOD_2_GMIN.set((const long int*)pAcq->blecsMOD_2_GMIN,128); | if (data.itemAvailable("blecsMOD_2_GMIN") == true)memcpy(blecsAcq.blecsMOD_2_GMIN,data.blecsMOD_2_GMIN.get(sz1),sizeof(blecsAcq.blecsMOD_2_GMIN)); | Modulation setup 2 minimum gain | Threshold value for the modulation test, minimum gain for the setup parameters 2 | [6 empty] | 68 | 68 |  | if (vme_addr[22..0] >= h"701650") and (vme_addr[22..0] < h"701850") then mux_sel[] = 68; end if; --Modulation setup 2 minimum gain / comments: Threshold value for the modulation test, minimum gain for the setup parameters 2
MOD_2_PMAX |  |  | FLASH | 701850 | 701A50 | 31..0 | unsigned | [128] | unsigned blecsMOD_2_PMAX[128]; | <field name="blecsMOD_2_PMAX" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold value for the modulation test, maximum phase for the setup parameters 2</description> <array type="long"><dim>128</dim></array> </field> | <data-field-ref-item field-name-ref="blecsMOD_2_PMAX"/> |  | pDev->blecsMOD_2_PMAX.set((const long int*)pAcq->blecsMOD_2_PMAX,128); | if (data.itemAvailable("blecsMOD_2_PMAX") == true)memcpy(blecsAcq.blecsMOD_2_PMAX,data.blecsMOD_2_PMAX.get(sz1),sizeof(blecsAcq.blecsMOD_2_PMAX)); | modulation setup 2 maximum phase | Threshold value for the modulation test, maximum phase for the setup parameters 2 | [6 empty] | 69 | 69 |  | if (vme_addr[22..0] >= h"701850") and (vme_addr[22..0] < h"701A50") then mux_sel[] = 69; end if; --modulation setup 2 maximum phase / comments: Threshold value for the modulation test, maximum phase for the setup parameters 2
MOD_2_PMIN |  |  | FLASH | 701A50 | 701C50 | 31..0 | unsigned | [128] | unsigned blecsMOD_2_PMIN[128]; | <field name="blecsMOD_2_PMIN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold value for the modulation test, minimum phase for the setup parameters 2</description> <array type="long"><dim>128</dim></array> </field> | <data-field-ref-item field-name-ref="blecsMOD_2_PMIN"/> |  | pDev->blecsMOD_2_PMIN.set((const long int*)pAcq->blecsMOD_2_PMIN,128); | if (data.itemAvailable("blecsMOD_2_PMIN") == true)memcpy(blecsAcq.blecsMOD_2_PMIN,data.blecsMOD_2_PMIN.get(sz1),sizeof(blecsAcq.blecsMOD_2_PMIN)); | Modulation setup 2 minimum phase | Threshold value for the modulation test, minimum phase for the setup parameters 2 | [6 empty] | 70 | 70 |  | if (vme_addr[22..0] >= h"701A50") and (vme_addr[22..0] < h"701C50") then mux_sel[] = 70; end if; --Modulation setup 2 minimum phase / comments: Threshold value for the modulation test, minimum phase for the setup parameters 2
|  |  |  | 701C50 | 701D00 |  | unsigned | [44] | unsigned blecsSPARE243[44]; | [5 empty] | SPARE
REGISTERS
USTR | TstRequUsrSysTest |  | TESTS | 701D00 | 701D04 | 31 | bool | :1 | bool blecsUSTR:1; | <field name="blecsUSTR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User System test</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUSTR"/> |  | pDev->blecsUSTR.set(pAcq->blecsUSTR); | if (data.itemAvailable("blecsUSTR") == true)blecsAcq.blecsUSTR = data.blecsUSTR.get(); | TestsRequest | User System test | [6 empty] | 71 | 71 |  | if (vme_addr[22..0] >= h"701D00") and (vme_addr[22..0] < h"701D04") then mux_sel[] = 71; TestsRequest = vme_acc; end if; --TestsRequest / comments: User System test
UCTR | TstRequUsrConsist |  | TESTS |  |  | 30 | bool | :1 | bool blecsUCTR:1; | <field name="blecsUCTR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User Consistency test</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUCTR"/> |  | pDev->blecsUCTR.set(pAcq->blecsUCTR); | if (data.itemAvailable("blecsUCTR") == true)blecsAcq.blecsUCTR = data.blecsUCTR.get(); |  | User Consistency test
UTBPLR | TstRequUsrThresBpl |  | TESTS |  |  | 29 | bool | :1 | bool blecsUTBPLR:1; | <field name="blecsUTBPLR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User Threshold to BPL test</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUTBPLR"/> |  | pDev->blecsUTBPLR.set(pAcq->blecsUTBPLR); | if (data.itemAvailable("blecsUTBPLR") == true)blecsAcq.blecsUTBPLR = data.blecsUTBPLR.get(); |  | User Threshold to BPL test
UETR | TstRequUsrEngy |  | TESTS |  |  | 28 | bool | :1 | bool blecsUETR:1; | <field name="blecsUETR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User Energy Test</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUETR"/> |  | pDev->blecsUETR.set(pAcq->blecsUETR); | if (data.itemAvailable("blecsUETR") == true)blecsAcq.blecsUETR = data.blecsUETR.get(); |  | User Energy Test
UBPBISTR | TstRequUsrBmPermBis |  | TESTS |  |  | 27 | bool | :1 | bool blecsUBPBISTR:1; | <field name="blecsUBPBISTR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User BPBIS (BLECS to Beam interlock)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUBPBISTR"/> |  | pDev->blecsUBPBISTR.set(pAcq->blecsUBPBISTR); | if (data.itemAvailable("blecsUBPBISTR") == true)blecsAcq.blecsUBPBISTR = data.blecsUBPBISTR.get(); |  | User BPBIS (BLECS to Beam interlock)
ESTR | TstRequExpSysTest |  | TESTS |  |  | 26 | bool | :1 | bool blecsESTR:1; | <field name="blecsESTR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert System test</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsESTR"/> |  | pDev->blecsESTR.set(pAcq->blecsESTR); | if (data.itemAvailable("blecsESTR") == true)blecsAcq.blecsESTR = data.blecsESTR.get(); |  | Expert System test |  |  |  |  | or External
ECTR | TstRequExpConsist |  | TESTS |  |  | 25 | bool | :1 | bool blecsECTR:1; | <field name="blecsECTR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Consistency test</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsECTR"/> |  | pDev->blecsECTR.set(pAcq->blecsECTR); | if (data.itemAvailable("blecsECTR") == true)blecsAcq.blecsECTR = data.blecsECTR.get(); |  | Expert Consistency test |  |  |  |  | from GUI
ETBPLR | TstRequExpThresBpl |  | TESTS |  |  | 24 | bool | :1 | bool blecsETBPLR:1; | <field name="blecsETBPLR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Threshold to BPL test</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsETBPLR"/> |  | pDev->blecsETBPLR.set(pAcq->blecsETBPLR); | if (data.itemAvailable("blecsETBPLR") == true)blecsAcq.blecsETBPLR = data.blecsETBPLR.get(); |  | Expert Threshold to BPL test |  |  |  |  | demande
EETR | TstRequExpEngy |  | TESTS |  |  | 23 | bool | :1 | bool blecsEETR:1; | <field name="blecsEETR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Energy Test</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEETR"/> |  | pDev->blecsEETR.set(pAcq->blecsEETR); | if (data.itemAvailable("blecsEETR") == true)blecsAcq.blecsEETR = data.blecsEETR.get(); |  | Expert Energy Test |  |  |  |  | R/W on
EBPBISR | TstRequExpBmPermBis |  | TESTS |  |  | 22 | bool | :1 | bool blecsEBPBISR:1; | <field name="blecsEBPBISR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert BPBIS (BLECS to Beam interlock)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEBPBISR"/> |  | pDev->blecsEBPBISR.set(pAcq->blecsEBPBISR); | if (data.itemAvailable("blecsEBPBISR") == true)blecsAcq.blecsEBPBISR = data.blecsEBPBISR.get(); |  | Expert BPBIS (BLECS to Beam interlock)
EBPTCR | TstRequExpBmPermTc |  | TESTS |  |  | 21 | bool | :1 | bool blecsEBPTCR:1; | <field name="blecsEBPTCR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert BPTC</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEBPTCR"/> |  | pDev->blecsEBPTCR.set(pAcq->blecsEBPTCR); | if (data.itemAvailable("blecsEBPTCR") == true)blecsAcq.blecsEBPTCR = data.blecsEBPTCR.get(); |  | Expert BPTC
EHVLFR | TstRequExpLowFreqModul |  | TESTS |  |  | 20 | bool | :1 | bool blecsEHVLFR:1; | <field name="blecsEHVLFR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVLF test</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEHVLFR"/> |  | pDev->blecsEHVLFR.set(pAcq->blecsEHVLFR); | if (data.itemAvailable("blecsEHVLFR") == true)blecsAcq.blecsEHVLFR = data.blecsEHVLFR.get(); |  | Expert HVLF test
EHVCFCR | TstRequExpCfcTest |  | TESTS |  |  | 19 | bool | :1 | bool blecsEHVCFCR:1; | <field name="blecsEHVCFCR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVCFC (100 pA test)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEHVCFCR"/> |  | pDev->blecsEHVCFCR.set(pAcq->blecsEHVCFCR); | if (data.itemAvailable("blecsEHVCFCR") == true)blecsAcq.blecsEHVCFCR = data.blecsEHVCFCR.get(); |  | Expert HVCFC (100 pA test)
EHVRDACR | TstRequExpRstDac |  | TESTS |  |  | 18 | bool | :1 | bool blecsEHVRDACR:1; | <field name="blecsEHVRDACR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVRDAC (Dac Reset)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEHVRDACR"/> |  | pDev->blecsEHVRDACR.set(pAcq->blecsEHVRDACR); | if (data.itemAvailable("blecsEHVRDACR") == true)blecsAcq.blecsEHVRDACR = data.blecsEHVRDACR.get(); |  | Expert HVRDAC (Dac Reset)
EHVRGOHR | TstRequExpRstGoh |  | TESTS |  |  | 17 | bool | :1 | bool blecsEHVRGOHR:1; | <field name="blecsEHVRGOHR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVGOH (GOH Reset)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEHVRGOHR"/> |  | pDev->blecsEHVRGOHR.set(pAcq->blecsEHVRGOHR); | if (data.itemAvailable("blecsEHVRGOHR") == true)blecsAcq.blecsEHVRGOHR = data.blecsEHVRGOHR.get(); |  | Expert HVGOH (GOH Reset)
EMCR | TstRequExpManCtrl |  | OVERVIEW |  |  | 16 | bool | :1 | bool blecsEMCR:1; | <field name="blecsEMCR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Manual control Request</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEMCR"/> |  | pDev->blecsEMCR.set(pAcq->blecsEMCR); | if (data.itemAvailable("blecsEMCR") == true)blecsAcq.blecsEMCR = data.blecsEMCR.get(); |  | Expert Manual control Request
[6 empty] | 0..15 | unsigned | :16 | unsigned blecsSPARE261:16; | [6 empty] | SPARE
FDUTC | FoceDumpTcUmsk |  | Beam Permit | 701D04 | 701D08 | 31 | bool | :1 | bool blecsFDUTC:1; | <field name="blecsFDUTC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Send command to TC: Force dump event Unmaskable on specific TC </description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFDUTC"/> |  | pDev->blecsFDUTC.set(pAcq->blecsFDUTC); | if (data.itemAvailable("blecsFDUTC") == true)blecsAcq.blecsFDUTC = data.blecsFDUTC.get(); | ManualControl | Send command to TC: Force dump event Unmaskable on specific TC |  |  |  |  | write from |  | 72 | 72 |  | if (vme_addr[22..0] >= h"701D04") and (vme_addr[22..0] < h"701D08") then mux_sel[] = 72; ManualControl = vme_acc; end if; --ManualControl / comments: Send command to TC: Force dump event Unmaskable on specific TC
FDMTC | ForceDumpTcMsk |  | Beam Permit |  |  | 30 | bool | :1 | bool blecsFDMTC:1; | <field name="blecsFDMTC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Send command to TC: Force dump event Maskable on specific TC </description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFDMTC"/> |  | pDev->blecsFDMTC.set(pAcq->blecsFDMTC); | if (data.itemAvailable("blecsFDMTC") == true)blecsAcq.blecsFDMTC = data.blecsFDMTC.get(); |  | Send command to TC: Force dump event Maskable on specific TC |  |  |  |  | Allow
FDCN | ForceDumpCdrNbr |  | Beam Permit |  |  | 26..29 | unsigned | :4 | unsigned blecsFDCN:4; | <field name="blecsFDCN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Force dump on TC card number</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsFDCN"/> |  | pDev->blecsFDCN.set(pAcq->blecsFDCN); | if (data.itemAvailable("blecsFDCN") == true)blecsAcq.blecsFDCN = data.blecsFDCN.get(); |  | Force dump on TC card number |  |  |  |  | SW1=1
BEESE | EnableSetEngy |  | Beam Energy |  |  | 25 | bool | :1 | bool blecsBEESE:1; | <field name="blecsBEESE" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Enargy Enable Set Energy to send to TC</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBEESE"/> |  | pDev->blecsBEESE.set(pAcq->blecsBEESE); | if (data.itemAvailable("blecsBEESE") == true)blecsAcq.blecsBEESE = data.blecsBEESE.get(); |  | Beam Enargy Enable Set Energy to send to TC |  |  |  |  | when
BEV | SetEngyValue |  | Beam Energy |  |  | 20..24 | unsigned | :5 | unsigned blecsBEV:5; | <field name="blecsBEV" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Energy value to be send</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBEV"/> |  | pDev->blecsBEV.set(pAcq->blecsBEV); | if (data.itemAvailable("blecsBEV") == true)blecsAcq.blecsBEV = data.blecsBEV.get(); |  | Beam Energy value to be send
FDEU | ForceDumpUmsk |  | Beam Permit |  |  | 19 | bool | :1 | bool blecsFDEU:1; | <field name="blecsFDEU" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Force Dump event Unmascable</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFDEU"/> |  | pDev->blecsFDEU.set(pAcq->blecsFDEU); | if (data.itemAvailable("blecsFDEU") == true)blecsAcq.blecsFDEU = data.blecsFDEU.get(); |  | Force Dump event Unmascable
FDEM | ForceDumpMsk |  | Beam Permit |  |  | 18 | bool | :1 | bool blecsFDEM:1; | <field name="blecsFDEM" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Force Dump event Mascable</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFDEM"/> |  | pDev->blecsFDEM.set(pAcq->blecsFDEM); | if (data.itemAvailable("blecsFDEM") == true)blecsAcq.blecsFDEM = data.blecsFDEM.get(); |  | Force Dump event Mascable
FBPLI | ForceBmPermEnable |  | Beam Permit |  |  | 17 | bool | :1 | bool blecsFBPLI:1; | <field name="blecsFBPLI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Force Beam Permit (Mascable and Unmascable) Lines individually</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFBPLI"/> |  | pDev->blecsFBPLI.set(pAcq->blecsFBPLI); | if (data.itemAvailable("blecsFBPLI") == true)blecsAcq.blecsFBPLI = data.blecsFBPLI.get(); |  | Force Beam Permit (Mascable and Unmascable) Lines individually
FBPUA | ForceBmPermUmskA |  | Beam Permit |  |  | 16 | bool | :1 | bool blecsFBPUA:1; | <field name="blecsFBPUA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Force Beam Permit UA</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFBPUA"/> |  | pDev->blecsFBPUA.set(pAcq->blecsFBPUA); | if (data.itemAvailable("blecsFBPUA") == true)blecsAcq.blecsFBPUA = data.blecsFBPUA.get(); |  | Force Beam Permit UA
FBPUB | ForceBmPermUmskB |  | Beam Permit |  |  | 15 | bool | :1 | bool blecsFBPUB:1; | <field name="blecsFBPUB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Force Beam Permit UB</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFBPUB"/> |  | pDev->blecsFBPUB.set(pAcq->blecsFBPUB); | if (data.itemAvailable("blecsFBPUB") == true)blecsAcq.blecsFBPUB = data.blecsFBPUB.get(); |  | Force Beam Permit UB
FBPMA | ForceBmPermMskA |  | Beam Permit |  |  | 14 | bool | :1 | bool blecsFBPMA:1; | <field name="blecsFBPMA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Force Beam Permit MA</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFBPMA"/> |  | pDev->blecsFBPMA.set(pAcq->blecsFBPMA); | if (data.itemAvailable("blecsFBPMA") == true)blecsAcq.blecsFBPMA = data.blecsFBPMA.get(); |  | Force Beam Permit MA
FBPMB | ForceBmPermMskB |  | Beam Permit |  |  | 13 | bool | :1 | bool blecsFBPMB:1; | <field name="blecsFBPMB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Force Beam Permit MB</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFBPMB"/> |  | pDev->blecsFBPMB.set(pAcq->blecsFBPMB); | if (data.itemAvailable("blecsFBPMB") == true)blecsAcq.blecsFBPMB = data.blecsFBPMB.get(); |  | Force Beam Permit MB | [8 empty] | when FBPLI=1 A or B can be 'true' while the other stay to 'false'
FSRTC | ForceSoftRstTC |  | Beam Energy |  |  | 12 | bool | :1 | bool blecsFSRTC:1; | <field name="blecsFSRTC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Force SoftResetTC</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFSRTC"/> |  | pDev->blecsFSRTC.set(pAcq->blecsFSRTC); | if (data.itemAvailable("blecsFSRTC") == true)blecsAcq.blecsFSRTC = data.blecsFSRTC.get(); |  | Force SoftResetTC
FSUT | ForceSysUnderTst |  | Beam Energy |  |  | 11 | bool | :1 | bool blecsFSUT:1; | <field name="blecsFSUT" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Force SystemUnderTest</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsFSUT"/> |  | pDev->blecsFSUT.set(pAcq->blecsFSUT); | if (data.itemAvailable("blecsFSUT") == true)blecsAcq.blecsFSUT = data.blecsFSUT.get(); |  | Force SystemUnderTest
[6 empty] | 0..10 | unsigned | :11 | unsigned blecsSPARE276:11; | [6 empty] | SPARE
|  |  |  | 701D08 | 701D20 |  | unsigned | [6] | unsigned blecsSPARE277[6]; | [5 empty] | SPARE | [5 empty] | the GUI
HVLFMP | LowFrequModulSetPeriod |  | HVLF test | 701D20 | 701D24 | 31..16 | unsigned | :16 | unsigned blecsHVLFMP:16; | <field name="blecsHVLFMP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Modulation Period [Period[s] = 2048/10**6]</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHVLFMP"/> |  | pDev->blecsHVLFMP.set(pAcq->blecsHVLFMP); | if (data.itemAvailable("blecsHVLFMP") == true)blecsAcq.blecsHVLFMP = data.blecsHVLFMP.get(); | DefaultHVLFControl_1 | Modulation Period [Period[s] = 2048/10**6] | [6 empty] | 58 |  |  | if (vme_addr[22..0] >= h"701D20") and (vme_addr[22..0] < h"701D24") then mux_sel[] = 58; DefaultHVLFControl_1 = vme_acc; end if; --DefaultHVLFControl_1 / comments: Modulation Period [Period[s] = 2048/10**6]
HVLFMAA | LowFrequModulSetAttenu |  | HVLF test |  |  | 15..8 | unsigned | :8 | unsigned blecsHVLFMAA:8; | <field name="blecsHVLFMAA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Modulation analog attenuator [Total = 100+Att/256*400]</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHVLFMAA"/> |  | pDev->blecsHVLFMAA.set(pAcq->blecsHVLFMAA); | if (data.itemAvailable("blecsHVLFMAA") == true)blecsAcq.blecsHVLFMAA = data.blecsHVLFMAA.get(); |  | Modulation analog attenuator [Total = 100+Att/256*400]
HVLFMDM | LowFrequModulSetMultipl |  | HVLF test |  |  | 7..0 | unsigned | :8 | unsigned blecsHVLFMDM:8; | <field name="blecsHVLFMDM" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Modulation digital multiplicator [V = 11.72*DigitalMultiplicator] </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHVLFMDM"/> |  | pDev->blecsHVLFMDM.set(pAcq->blecsHVLFMDM); | if (data.itemAvailable("blecsHVLFMDM") == true)blecsAcq.blecsHVLFMDM = data.blecsHVLFMDM.get(); |  | Modulation digital multiplicator [V = 11.72*DigitalMultiplicator]
HVLFMP2 | LowFrequModulSetPeriod |  | HVLF test | 701D24 | 701D28 | 31..16 | unsigned | :16 | unsigned blecsHVLFMP2:16; | <field name="blecsHVLFMP2" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Modulation Period [Period[s] = 2048/10**6]</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHVLFMP2"/> |  | pDev->blecsHVLFMP2.set(pAcq->blecsHVLFMP2); | if (data.itemAvailable("blecsHVLFMP2") == true)blecsAcq.blecsHVLFMP2 = data.blecsHVLFMP2.get(); | DefaultHVLFControl_2 | Modulation Period [Period[s] = 2048/10**6] | [6 empty] | 59 |  |  | if (vme_addr[22..0] >= h"701D24") and (vme_addr[22..0] < h"701D28") then mux_sel[] = 59; DefaultHVLFControl_2 = vme_acc; end if; --DefaultHVLFControl_2 / comments: Modulation Period [Period[s] = 2048/10**6]
HVLFMAA2 | LowFrequModulSetAttenu |  | HVLF test |  |  | 15..8 | unsigned | :8 | unsigned blecsHVLFMAA2:8; | <field name="blecsHVLFMAA2" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Modulation analog attenuator [Total = 100+Att/256*400]</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHVLFMAA2"/> |  | pDev->blecsHVLFMAA2.set(pAcq->blecsHVLFMAA2); | if (data.itemAvailable("blecsHVLFMAA2") == true)blecsAcq.blecsHVLFMAA2 = data.blecsHVLFMAA2.get(); |  | Modulation analog attenuator [Total = 100+Att/256*400]
HVLFMDM2 | LowFrequModulSetMultipl |  | HVLF test |  |  | 7..0 | unsigned | :8 | unsigned blecsHVLFMDM2:8; | <field name="blecsHVLFMDM2" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Modulation digital multiplicator [V = 11.72*DigitalMultiplicator] </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHVLFMDM2"/> |  | pDev->blecsHVLFMDM2.set(pAcq->blecsHVLFMDM2); | if (data.itemAvailable("blecsHVLFMDM2") == true)blecsAcq.blecsHVLFMDM2 = data.blecsHVLFMDM2.get(); |  | Modulation digital multiplicator [V = 11.72*DigitalMultiplicator]
SHVNO | HighVSetNormalVolt |  |  | 701D28 | 701D2C | 31..16 | unsigned | :16 | unsigned blecsSHVNO:16; | <field name="blecsSHVNO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage Normal Operation in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsSHVNO"/> |  | pDev->blecsSHVNO.set(pAcq->blecsSHVNO); | if (data.itemAvailable("blecsSHVNO") == true)blecsAcq.blecsSHVNO = data.blecsSHVNO.get(); | DefaultHVvalues_1 | Set High Voltage Normal Operation in dac value | [6 empty] | 53 |  |  | if (vme_addr[22..0] >= h"701D28") and (vme_addr[22..0] < h"701D2C") then mux_sel[] = 53; DefaultHVvalues_1 = vme_acc; end if; -- DefaultHVvalues_1 / comments: Set High Voltage Normal Operation in dac value
SHVDIF | HighVSetNormalVoltDiff |  |  |  |  | 15..0 | unsigned | :16 | unsigned blecsSHVDIF:16; | <field name="blecsSHVDIF" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set Difference between HV1 and HV2 (HV2 lower than HV1) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsSHVDIF"/> |  | pDev->blecsSHVDIF.set(pAcq->blecsSHVDIF); | if (data.itemAvailable("blecsSHVDIF") == true)blecsAcq.blecsSHVDIF = data.blecsSHVDIF.get(); |  | Set Difference between HV1 and HV2 (HV2 lower than HV1) in dac value
SHVCFC | HighVSetCfcTstVolt |  | HV & LV | 701D2C | 701D30 | 31..16 | unsigned | :16 | unsigned blecsSHVCFC:16; | <field name="blecsSHVCFC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage HVCFC (100 pA test) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsSHVCFC"/> |  | pDev->blecsSHVCFC.set(pAcq->blecsSHVCFC); | if (data.itemAvailable("blecsSHVCFC") == true)blecsAcq.blecsSHVCFC = data.blecsSHVCFC.get(); | DefaultHVvalues_2 | Set High Voltage HVCFC (100 pA test) in dac value | [6 empty] | 54 |  |  | if (vme_addr[22..0] >= h"701D2C") and (vme_addr[22..0] < h"701D30") then mux_sel[] = 54; DefaultHVvalues_2 = vme_acc; end if; --DefaultHVvalues_2 / comments: Set High Voltage HVCFC (100 pA test) in dac value
HVLFSHV | LowFrequModulSetVoltValue |  | Control and |  |  | 15..0 | unsigned | :16 | unsigned blecsHVLFSHV:16; | <field name="blecsHVLFSHV" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage for modulation in dac values (U[V] = 3000/2**16*dacValue)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHVLFSHV"/> |  | pDev->blecsHVLFSHV.set(pAcq->blecsHVLFSHV); | if (data.itemAvailable("blecsHVLFSHV") == true)blecsAcq.blecsHVLFSHV = data.blecsHVLFSHV.get(); |  | Set High Voltage for modulation in dac values (U[V] = 3000/2**16*dacValue) | [8 empty] | In DAC values
SHVGOH | HighVSetRstGohVolt |  |  | 701D30 | 701D34 | 31..16 | unsigned | :16 | unsigned blecsSHVGOH:16; | <field name="blecsSHVGOH" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage HVGOH (GOH Reset) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsSHVGOH"/> |  | pDev->blecsSHVGOH.set(pAcq->blecsSHVGOH); | if (data.itemAvailable("blecsSHVGOH") == true)blecsAcq.blecsSHVGOH = data.blecsSHVGOH.get(); | DefaultHVvalues_3 | Set High Voltage HVGOH (GOH Reset) in dac value | [6 empty] | 55 |  | 3000/2**16*dacValue | if (vme_addr[22..0] >= h"701D30") and (vme_addr[22..0] < h"701D34") then mux_sel[] = 55; DefaultHVvalues_3 = vme_acc; end if; --DefaultHVvalues_3 / comments: Set High Voltage HVGOH (GOH Reset) in dac value
SHVRDAC | HighVSetRstDacVolt |  | Survey |  |  | 15..0 | unsigned | :16 | unsigned blecsSHVRDAC:16; | <field name="blecsSHVRDAC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage HVRDAC (Dac Reset) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsSHVRDAC"/> |  | pDev->blecsSHVRDAC.set(pAcq->blecsSHVRDAC); | if (data.itemAvailable("blecsSHVRDAC") == true)blecsAcq.blecsSHVRDAC = data.blecsSHVRDAC.get(); |  | Set High Voltage HVRDAC (Dac Reset) in dac value | [8 empty] | U(V) =
SHVRDACP | HighVSetRstDacPeekVolt |  |  | 701D34 | 701D38 | 31..16 | unsigned | :16 | unsigned blecsSHVRDACP:16; | <field name="blecsSHVRDACP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage (Peek Value) HVRDAC (Dac Reset) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsSHVRDACP"/> |  | pDev->blecsSHVRDACP.set(pAcq->blecsSHVRDACP); | if (data.itemAvailable("blecsSHVRDACP") == true)blecsAcq.blecsSHVRDACP = data.blecsSHVRDACP.get(); | DefaultHVvalues_4 | Set High Voltage (Peek Value) HVRDAC (Dac Reset) in dac value | [6 empty] | 56 |  |  | if (vme_addr[22..0] >= h"701D34") and (vme_addr[22..0] < h"701D38") then mux_sel[] = 56; DefaultHVvalues_4 = vme_acc; end if; --DefaultHVvalues_4 / comments: Set High Voltage (Peek Value) HVRDAC (Dac Reset) in dac value
SHVCFCP | HighVSetCfcTstPeekVolt |  |  |  |  | 15..0 | unsigned | :16 | unsigned blecsSHVCFCP:16; | <field name="blecsSHVCFCP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage (Peek Value) HVCFC (100 pA test) in dac value </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsSHVCFCP"/> |  | pDev->blecsSHVCFCP.set(pAcq->blecsSHVCFCP); | if (data.itemAvailable("blecsSHVCFCP") == true)blecsAcq.blecsSHVCFCP = data.blecsSHVCFCP.get(); |  | Set High Voltage (Peek Value) HVCFC (100 pA test) in dac value
|  |  |  | 701D38 | 701D3C | 31..16 | unsigned | :16 | unsigned blecsSPARE292:16; | [5 empty] | DefaultHVvalues_5 | SPARE | [6 empty] | 57 |  |  | if (vme_addr[22..0] >= h"701D38") and (vme_addr[22..0] < h"701D3C") then mux_sel[] = 57; DefaultHVvalues_5 = vme_acc; end if; --DefaultHVvalues_5 / comments: SPARE
SHVRGOHP | HighVSetRstGohPeekVolt |  |  |  |  | 15..0 | unsigned | :16 | unsigned blecsSHVRGOHP:16; | <field name="blecsSHVRGOHP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Set High Voltage (Peek Value) HVGOH (GOH Reset) in dac value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsSHVRGOHP"/> |  | pDev->blecsSHVRGOHP.set(pAcq->blecsSHVRGOHP); | if (data.itemAvailable("blecsSHVRGOHP") == true)blecsAcq.blecsSHVRGOHP = data.blecsSHVRGOHP.get(); |  | Set High Voltage (Peek Value) HVGOH (GOH Reset) in dac value
|  |  |  | 701D3C | 701D50 |  | unsigned | [5] | unsigned blecsSPARE294[5]; | [5 empty] | SPARE
MAXRUNTABLE | RunMaxTbl |  | HVLF test | 701D50 | 702150 | 31..0 | unsigned | [256] | unsigned blecsMAXRUNTABLE[256]; | <field name="blecsMAXRUNTABLE" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Channel 1 (TC 1) Running Maximum Nº9 </description> <array type="long"><dim>256</dim></array> </field> | <data-field-ref-item field-name-ref="blecsMAXRUNTABLE"/> |  | pDev->blecsMAXRUNTABLE.set((const long int*)pAcq->blecsMAXRUNTABLE,256); | if (data.itemAvailable("blecsMAXRUNTABLE") == true)memcpy(blecsAcq.blecsMAXRUNTABLE,data.blecsMAXRUNTABLE.get(sz1),sizeof(blecsAcq.blecsMAXRUNTABLE)); | RunningMaximumN9table | Channel 1 (TC 1) Running Maximum Nº9 | [6 empty] | 73 | 73 |  | if (vme_addr[22..0] >= h"701D50") and (vme_addr[22..0] < h"702150") then mux_sel[] = 73; RunningMaximumN9table = vme_acc; end if; --RunningMaximumN9table / comments: Channel 1 (TC 1) Running Maximum Nº9
|  |  | HVLF test |  |  | 31..0 | [9 empty] | Channel 2 (TC 1)
|  |  | HVLF test |  |  | 31..0 | [9 empty] | … | [8 empty] | Modulation test
|  |  | HVLF test |  |  | 31..0 | [9 empty] | Channel 255 (TC 16)
|  |  | HVLF test |  |  | 31..0 | [9 empty] | Channel 256 (TC 16) |  |  |  |  | Write at
BLECFSHV | BlecfHighVStatus |  | BLECF TEST | 702150 | 702154 | 31..0 | unsigned | :32 | unsigned blecsBLECFSHV:32; | <field name="blecsBLECFSHV" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Status HV. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFSHV"/> |  | pDev->blecsBLECFSHV.set(pAcq->blecsBLECFSHV); | if (data.itemAvailable("blecsBLECFSHV") == true)blecsAcq.blecsBLECFSHV = data.blecsBLECFSHV.get(); | StatusHV_BLECF | Status HV. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) |  |  |  |  | every |  | 74 | 74 | Info compiled by the CPU from all BLETC | if (vme_addr[22..0] >= h"702150") and (vme_addr[22..0] < h"702154") then mux_sel[] = 74; StatusHV_BLECF = vme_acc; end if; --StatusHV_BLECF / comments: Status HV. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16)
BLECFSTC | BlecfTstCfcStatus |  | BLECF TEST | 702154 | 702158 | 31..0 | unsigned | :32 | unsigned blecsBLECFSTC:32; | <field name="blecsBLECFSTC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Status TESTCFC. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFSTC"/> |  | pDev->blecsBLECFSTC.set(pAcq->blecsBLECFSTC); | if (data.itemAvailable("blecsBLECFSTC") == true)blecsAcq.blecsBLECFSTC = data.blecsBLECFSTC.get(); | StatusTESTCFC_BLECF | Status TESTCFC. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) |  |  |  |  | logging |  | 75 | 75 | Update every login read when "system under test". | if (vme_addr[22..0] >= h"702154") and (vme_addr[22..0] < h"702158") then mux_sel[] = 75; StatusTESTCFC_BLECF = vme_acc; end if; --StatusTESTCFC_BLECF / comments: Status TESTCFC. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16)
BLECFSRD | BlecfRstDacStatus |  | BLECF TEST | 702158 | 70215C | 31..0 | unsigned | :32 | unsigned blecsBLECFSRD:32; | <field name="blecsBLECFSRD" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Status RSTDAC. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFSRD"/> |  | pDev->blecsBLECFSRD.set(pAcq->blecsBLECFSRD); | if (data.itemAvailable("blecsBLECFSRD") == true)blecsAcq.blecsBLECFSRD = data.blecsBLECFSRD.get(); | StatusRSTDAC_BLECF | Status RSTDAC. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) |  |  |  |  | when |  | 76 | 76 |  | if (vme_addr[22..0] >= h"702158") and (vme_addr[22..0] < h"70215C") then mux_sel[] = 76; StatusRSTDAC_BLECF = vme_acc; end if; --StatusRSTDAC_BLECF / comments: Status RSTDAC. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16)
BLECFSRG | BlecfRstGohStatus |  | BLECF TEST | 70215C | 702160 | 31..0 | unsigned | :32 | unsigned blecsBLECFSRG:32; | <field name="blecsBLECFSRG" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Status RSTGOH. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFSRG"/> |  | pDev->blecsBLECFSRG.set(pAcq->blecsBLECFSRG); | if (data.itemAvailable("blecsBLECFSRG") == true)blecsAcq.blecsBLECFSRG = data.blecsBLECFSRG.get(); | StatusRSTGOH_BLECF | Status RSTGOH. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) |  |  |  |  | TEST=1 |  | 77 | 77 |  | if (vme_addr[22..0] >= h"70215C") and (vme_addr[22..0] < h"702160") then mux_sel[] = 77; StatusRSTGOH_BLECF = vme_acc; end if; --StatusRSTGOH_BLECF / comments: Status RSTGOH. MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16)
BLECFSTCON | BlecfTstCfcStatusOn |  | BLECF TEST | 702160 | 702164 | 31..0 | unsigned | :32 | unsigned blecsBLECFSTCON:32; | <field name="blecsBLECFSTCON" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Status TESTCFC Received by the CFC MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFSTCON"/> |  | pDev->blecsBLECFSTCON.set(pAcq->blecsBLECFSTCON); | if (data.itemAvailable("blecsBLECFSTCON") == true)blecsAcq.blecsBLECFSTCON = data.blecsBLECFSTCON.get(); | StatusTESTCFC_BLECF_ON | Status TESTCFC Received by the CFC MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) | [6 empty] | 78 | 78 |  | if (vme_addr[22..0] >= h"702160") and (vme_addr[22..0] < h"702164") then mux_sel[] = 78; StatusTESTCFC_BLECF_ON = vme_acc; end if; --StatusTESTCFC_BLECF_ON / comments: Status TESTCFC Received by the CFC MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16)
BLECFSRDON | BlecfRstDacStatusOn |  | BLECF TEST | 702164 | 702168 | 31..0 | unsigned | :32 | unsigned blecsBLECFSRDON:32; | <field name="blecsBLECFSRDON" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Status RSTDAC Received by the CFC MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFSRDON"/> |  | pDev->blecsBLECFSRDON.set(pAcq->blecsBLECFSRDON); | if (data.itemAvailable("blecsBLECFSRDON") == true)blecsAcq.blecsBLECFSRDON = data.blecsBLECFSRDON.get(); | StatusRSTDAC_BLECF_ON | Status RSTDAC Received by the CFC MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) | [6 empty] | 79 | 79 |  | if (vme_addr[22..0] >= h"702164") and (vme_addr[22..0] < h"702168") then mux_sel[] = 79; StatusRSTDAC_BLECF_ON = vme_acc; end if; --StatusRSTDAC_BLECF_ON / comments: Status RSTDAC Received by the CFC MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16)
BLECFSRGON | BlecfRstGohStatusON |  | BLECF TEST | 702168 | 70216C | 31..0 | unsigned | :32 | unsigned blecsBLECFSRGON:32; | <field name="blecsBLECFSRGON" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Status RSTGOH Received by the CFC MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) </description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFSRGON"/> |  | pDev->blecsBLECFSRGON.set(pAcq->blecsBLECFSRGON); | if (data.itemAvailable("blecsBLECFSRGON") == true)blecsAcq.blecsBLECFSRGON = data.blecsBLECFSRGON.get(); | StatusRSTGOH_BLECF_ON | Status RSTGOH Received by the CFC MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16) | [6 empty] | 80 | 80 |  | if (vme_addr[22..0] >= h"702168") and (vme_addr[22..0] < h"70216C") then mux_sel[] = 80; StatusRSTGOH_BLECF_ON = vme_acc; end if; --StatusRSTGOH_BLECF_ON / comments: Status RSTGOH Received by the CFC MSBit=BLECF-1 (TC1), LSBit=BLECF-32 (TC16)
|  |  |  | 70216C | 702170 |  | unsigned | [1] | unsigned blecsSPARE307[1]; | [5 empty] | SPARE
|  |  | - | 702170 | 702174 | 6..31 | unsigned | :26 | unsigned blecsSPARE308:26; | [5 empty] | ExternalTestResults | SPARE | [6 empty] | 81 | 81 |  | if (vme_addr[22..0] >= h"702170") and (vme_addr[22..0] < h"702174") then mux_sel[] = 81; ExternalTestResults = vme_acc; end if; --ExternalTestResults / comments: SPARE
ETRBPBISR | ExtTstResRdyBmPermBis |  | - |  |  | 5 | bool | :1 | bool blecsETRBPBISR:1; | <field name="blecsETRBPBISR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>External Test result ready: BPBIS</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsETRBPBISR"/> |  | pDev->blecsETRBPBISR.set(pAcq->blecsETRBPBISR); | if (data.itemAvailable("blecsETRBPBISR") == true)blecsAcq.blecsETRBPBISR = data.blecsETRBPBISR.get(); |  | External Test result ready: BPBIS
ETRBPBIS | ExtTstResBmPermBis |  | - |  |  | 4 | bool | :1 | bool blecsETRBPBIS:1; | <field name="blecsETRBPBIS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>External Test result: BPBIS given by the CPU 1=passed, 0=failed</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsETRBPBIS"/> |  | pDev->blecsETRBPBIS.set(pAcq->blecsETRBPBIS); | if (data.itemAvailable("blecsETRBPBIS") == true)blecsAcq.blecsETRBPBIS = data.blecsETRBPBIS.get(); |  | External Test result: BPBIS given by the CPU 1=passed, 0=failed |  |  |  |  | External |  |  |  | will come from the Beam interlock groupe
ETRTBPLR | ExtTstResRdyThresBpl |  | - |  |  | 3 | bool | :1 | bool blecsETRTBPLR:1; | <field name="blecsETRTBPLR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>External Test result ready: Threshold to BPL</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsETRTBPLR"/> |  | pDev->blecsETRTBPLR.set(pAcq->blecsETRTBPLR); | if (data.itemAvailable("blecsETRTBPLR") == true)blecsAcq.blecsETRTBPLR = data.blecsETRTBPLR.get(); |  | External Test result ready: Threshold to BPL |  |  |  |  | from
ETRTBPL | ExtTstResThresBpl |  | - |  |  | 2 | bool | :1 | bool blecsETRTBPL:1; | <field name="blecsETRTBPL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>External Test result: Threshold to BPL given by the CPU 1=passed, 0=failed</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsETRTBPL"/> |  | pDev->blecsETRTBPL.set(pAcq->blecsETRTBPL); | if (data.itemAvailable("blecsETRTBPL") == true)blecsAcq.blecsETRTBPL = data.blecsETRTBPL.get(); |  | External Test result: Threshold to BPL given by the CPU 1=passed, 0=failed |  |  |  |  | demande
ETRCTYR | ExtTstResRdyConsist |  | - |  |  | 1 | bool | :1 | bool blecsETRCTYR:1; | <field name="blecsETRCTYR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>External Test result ready: Consistency</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsETRCTYR"/> |  | pDev->blecsETRCTYR.set(pAcq->blecsETRCTYR); | if (data.itemAvailable("blecsETRCTYR") == true)blecsAcq.blecsETRCTYR = data.blecsETRCTYR.get(); |  | External Test result ready: Consistency |  |  |  |  | R/W on
ETRCTY | ExtTstResConsist |  | - |  |  | 0 | bool | :1 | bool blecsETRCTY:1; | <field name="blecsETRCTY" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>External Test result: Consistency given by the CPU 1=passed, 0=failed</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsETRCTY"/> |  | pDev->blecsETRCTY.set(pAcq->blecsETRCTY); | if (data.itemAvailable("blecsETRCTY") == true)blecsAcq.blecsETRCTY = data.blecsETRCTY.get(); |  | External Test result: Consistency given by the CPU 1=passed, 0=failed | [8 empty] | Will come from the comparaison between LSA data and Hadware data
ETRTPBLETC | - |  | - | 702174 | 702178 | 16..31 | unsigned | :16 | unsigned blecsETRTPBLETC:16; | <field name="blecsETRTPBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Treshold to BPL result per BLETC, MSBit=BLETC_1, LSBit=BLETC_16, 1=passed, 0=failed</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsETRTPBLETC"/> |  | pDev->blecsETRTPBLETC.set(pAcq->blecsETRTPBLETC); | if (data.itemAvailable("blecsETRTPBLETC") == true)blecsAcq.blecsETRTPBLETC = data.blecsETRTPBLETC.get(); | ExternalTestResultPerBLETC | Treshold to BPL result per BLETC, MSBit=BLETC_1, LSBit=BLETC_16, 1=passed, 0=failed |  |  |  |  | include these data |  | 82 | 82 |  | if (vme_addr[22..0] >= h"702174") and (vme_addr[22..0] < h"702178") then mux_sel[] = 82; ExternalTestResultPerBLETC = vme_acc; end if; --ExternalTestResultPerBLETC / comments: Treshold to BPL result per BLETC, MSBit=BLETC_1, LSBit=BLETC_16, 1=passed, 0=failed
ETRCPBLETC | - |  | - |  |  | 0..15 | unsigned | :16 | unsigned blecsETRCPBLETC:16; | <field name="blecsETRCPBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Consistency result per BLETC, MSBit=BLETC_1, LSBit=BLETC_16, 1=passed, 0=failed</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsETRCPBLETC"/> |  | pDev->blecsETRCPBLETC.set(pAcq->blecsETRCPBLETC); | if (data.itemAvailable("blecsETRCPBLETC") == true)blecsAcq.blecsETRCPBLETC = data.blecsETRCPBLETC.get(); |  | Consistency result per BLETC, MSBit=BLETC_1, LSBit=BLETC_16, 1=passed, 0=failed |  |  |  |  | in the logging DB
|  |  | - | 702178 | 70217C | 5..31 | unsigned | :27 | unsigned blecsSPARE317:27; | [5 empty] | ExternalControl_BPlines | SPARE | [6 empty] | 83 | 83 |  | if (vme_addr[22..0] >= h"702178") and (vme_addr[22..0] < h"70217C") then mux_sel[] = 83; ExternalControl_BPlines = vme_acc; end if; --ExternalControl_BPlines / comments: SPARE
BISFMB | - |  | - |  |  | 4 | bool | :1 | bool blecsBISFMB:1; | <field name="blecsBISFMB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>BIS force beam permit maskable B</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBISFMB"/> |  | pDev->blecsBISFMB.set(pAcq->blecsBISFMB); | if (data.itemAvailable("blecsBISFMB") == true)blecsAcq.blecsBISFMB = data.blecsBISFMB.get(); |  | BIS force beam permit maskable B
BISFMA | - |  | - |  |  | 3 | bool | :1 | bool blecsBISFMA:1; | <field name="blecsBISFMA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>BIS force beam permit maskable A</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBISFMA"/> |  | pDev->blecsBISFMA.set(pAcq->blecsBISFMA); | if (data.itemAvailable("blecsBISFMA") == true)blecsAcq.blecsBISFMA = data.blecsBISFMA.get(); |  | BIS force beam permit maskable A | [8 empty] | to the CIBUS
BISFUB | - |  | - |  |  | 2 | bool | :1 | bool blecsBISFUB:1; | <field name="blecsBISFUB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>BIS force beam permit Unmaskable B</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBISFUB"/> |  | pDev->blecsBISFUB.set(pAcq->blecsBISFUB); | if (data.itemAvailable("blecsBISFUB") == true)blecsAcq.blecsBISFUB = data.blecsBISFUB.get(); |  | BIS force beam permit Unmaskable B | [8 empty] | from the BLM system (with the combiner)
BISFUA | - |  | - |  |  | 1 | bool | :1 | bool blecsBISFUA:1; | <field name="blecsBISFUA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>BIS force beam permit Unmaskable A</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBISFUA"/> |  | pDev->blecsBISFUA.set(pAcq->blecsBISFUA); | if (data.itemAvailable("blecsBISFUA") == true)blecsAcq.blecsBISFUA = data.blecsBISFUA.get(); |  | BIS force beam permit Unmaskable A | [8 empty] | Used by the BIS to test the Beam permit lines
| - |  | - |  |  | 0 | bool | :1 | bool blecsSPARE322:1; | <field name="blecs" multiplexing-criterion="NONE" persistency="VOLATILE"> <description></description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecs"/> |  | pDev->blecs.set(pAcq->blecs); | if (data.itemAvailable("blecs") == true)blecsAcq.blecs = data.blecs.get();
|  |  | - | 70217C | 702200 |  | unsigned | [33] | unsigned blecsSPARE323[33]; | [6 empty] | SPARE
[17 empty] | STATUS logged | to the concentrator | total in bits
| 2176 DWORD to read on the combiner | [14 empty] | Variables to the logging database (bool and unsigned) | 99
[16 empty] | Variables not in the DB but goes to the concentrator |  | 9
[16 empty] | Total of usefull informations in bits | 16955 | 71 | 17026

## DAB64x Control Status Registers
MEMORY MAPPING BLECS
DAB64x Control Status Registers | 7F0000 | 7FFFFF
Description | Start | End | data (x32bits) | Comments
Flash ROM Control/Status | 7F0000 | 7F0004 | #N/A
Revision ADC 0 | 7F0004 | 7F0005 | #N/A | byte0 - 0x7F0004 - Revision ADC High byte
Revision ADC 1 | 7F0005 | 7F0006 | #N/A | byte1 - 0x7F0005 - Revision ADC Low byte
Revision ADC 3 4 | 7F0006 | 7F0008 | #N/A | reserved
SPARE | 7F0008 | 7F0014 | #N/A
Revision Letter/Date 0 | 7F0014 | 7F0015 | #N/A | byte0 - 0x7F0014 - Revision Letter
Revision Letter/Date 1 | 7F0015 | 7F0016 | #N/A | byte1 - 0x7F0015 - Revision Date MM
Revision Letter/Date 2 | 7F0016 | 7F0017 | #N/A | byte2 - 0x7F0016 - Revision Date DD
Revision Letter/Date 3 | 7F0017 | 7F0018 | #N/A | byte3 - 0x7F0017 - Revision Date yy
Monitor Temp/Power 0 | 7F0018 | 7F0019 | #N/A | byte0 - 0x7F0018 - STRATIX Temp High byte
Monitor Temp/Power 1 | 7F0019 | 7F001A | #N/A | byte1 - 0x7F0019 - STRATIX Temp Low byte
Monitor Temp/Power 2 | 7F001A | 7F001B | #N/A | byte2 - 0x7F001A - HWBTN Power Monitor Status
Monitor Temp/Power 3 | 7F001B | 7F001C | #N/A | byte3 - 0x7F001B - VWBTN Power Monitor Status
SPARE | 7F001C | 7F003C | #N/A
Serial Number MSB 0 | 7F003C | 7F003D | #N/A | byte0 - 0x7F003C - Serial Number Upper 4 bytes
Serial Number MSB 1 | 7F003D | 7F003E | #N/A | byte1 - 0x7F003D
Serial Number MSB 2 | 7F003E | 7F003F | #N/A | byte2 - 0x7F003E
Serial Number MSB 4 | 7F003F | 7F0040 | #N/A | byte3 - 0x7F003F
Serial Number LSB 0 | 7F0040 | 7F0041 | #N/A | byte0 - 0x7F0040 - Serial Number Lower 4 bytes
Serial Number LSB 1 | 7F0041 | 7F0042 | #N/A | byte1 - 0x7F0041
Serial Number LSB 2 | 7F0042 | 7F0043 | #N/A | byte2 - 0x7F0042
Serial Number LSB 3 | 7F0043 | 7F0044 | #N/A | byte3 - 0x7F0043
SPARE | 7F0044 | 7F0050 | #N/A
PM SRAM Read Pointer | 7F0050 | 7F0054 | #N/A | MSB 1/0 for full/partial read
BD SRAM Read Pointer | 7F0054 | 7F0058 | #N/A | MSB 1/0 for full/partial read
Xtra SRAM Read Pointer | 7F0058 | 7F005C | #N/A | MSB 1/0 for full/partial read
SPARE | 7F005C | 7F0100 | #N/A
Collimation Data - Ch 01 | 7F0100 | 7F0104 | #N/A | 32 values deep FIFO
Collimation Data - Ch 02 | 7F0104 | 7F0108 | #N/A | 32 values deep FIFO
Collimation Data - Ch 03 | 7F0108 | 7F010C | #N/A | 32 values deep FIFO
Collimation Data - Ch 04 | 7F010C | 7F0110 | #N/A | 32 values deep FIFO
Collimation Data - Ch 05 | 7F0110 | 7F0114 | #N/A | 32 values deep FIFO
Collimation Data - Ch 06 | 7F0114 | 7F0118 | #N/A | 32 values deep FIFO
Collimation Data - Ch 07 | 7F0118 | 7F011C | #N/A | 32 values deep FIFO
Collimation Data - Ch 08 | 7F011C | 7F0120 | #N/A | 32 values deep FIFO
Collimation Data - Ch 09 | 7F0120 | 7F0124 | #N/A | 32 values deep FIFO
Collimation Data - Ch 10 | 7F0124 | 7F0128 | #N/A | 32 values deep FIFO
Collimation Data - Ch 11 | 7F0128 | 7F012C | #N/A | 32 values deep FIFO
Collimation Data - Ch 12 | 7F012C | 7F0130 | #N/A | 32 values deep FIFO
Collimation Data - Ch 13 | 7F0130 | 7F0134 | #N/A | 32 values deep FIFO
Collimation Data - Ch 14 | 7F0134 | 7F0138 | #N/A | 32 values deep FIFO
Collimation Data - Ch 15 | 7F0138 | 7F013C | #N/A | 32 values deep FIFO
Collimation Data - Ch 16 | 7F013C | 7F0140 | #N/A | 32 values deep FIFO
Collimator Data Overflow | 7F0140 | 7F0144 | #N/A | Register
SPARE | 7F0144 | 7FC000 | #N/A
BLM Logging Memory A | 7FC000 | 7FC200 | #N/A | used for the First 8 detectors
BLM Logging Memory B | 7FC200 | 7FC400 | #N/A | used for the Second 8 detectors
BLM ESL Memory A | 7FC400 | 7FC440 | #N/A | used for the First 8 detectors
BLM ESL Memory B | 7FC440 | 7FC480 | #N/A | used for the Second 8 detectors
BLM ESL Memory (ADC RANGE) | 7FC480 | 7FC500 | #N/A | Calculated ADC RANGE for all detectors
BLM Th Memory A | 7FC500 | 7FC700 | #N/A | used for the First 8 detectors
BLM Th Memory B | 7FC700 | 7FC900 | #N/A | used for the Second 8 detectors
SPARE | 7FC900 | 800000 | #N/A

## Search Results
SYS_OP | SysOperational | SYS_OPERATIONAL | OVERVIEW | 700000 | 700004 | 31 | bool | :1 | bool blecsSYS_OP:1; | <field name="blecsSYS_OP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>System is operational (View by the BLECS)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsSYS_OP"/> | pDev->blecsSYS_OP.set(pAcq->blecsSYS_OP); | System level | System is operational (View by the BLECS) | 1 | 1 |  |  |  | 4 | if (vme_addr[22..0] >= h"700000") and (vme_addr[22..0] < h"700004") then mux_sel[] = 4; end if; --System level / comments: System is operational (View by the BLECS)
TEST | SysUnderTest | SYS_UNDER_TEST | OVERVIEW |  |  | 30 | bool | :1 | bool blecsTEST:1; | <field name="blecsTEST" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>System is under test (All BLM crate of the IP are involved but only this crate is tested)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsTEST"/> | pDev->blecsTEST.set(pAcq->blecsTEST); |  | System is under test (All BLM crate of the IP are involved but only this crate is tested) | 1 | 1
T_REQ | TstRequ | TEST_PEND_GLOBAL | OVERVIEW |  |  | 29 | bool | :1 | bool blecsT_REQ:1; | <field name="blecsT_REQ" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test requested (while take place when the conditions are fillfull)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsT_REQ"/> | pDev->blecsT_REQ.set(pAcq->blecsT_REQ); |  | Test requested (while take place when the conditions are fillfull) | 1 | 1
HV1 | FaultHighV1Comp | HV_1_COMP_FAULT | OVERVIEW |  |  | 27 | bool | :1 | bool blecsHV1:1; | <field name="blecsHV1" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High voltage 1 fault (survey by comparators)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsHV1"/> | pDev->blecsHV1.set(pAcq->blecsHV1); |  | High voltage 1 fault (survey by comparators) | 1 | 1
HV2 | FaultHighV2Comp | HV_2_COMP_FAULT | OVERVIEW |  |  | 26 | bool | :1 | bool blecsHV2:1; | <field name="blecsHV2" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High voltage 2 fault (survey by comparators)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsHV2"/> | pDev->blecsHV2.set(pAcq->blecsHV2); |  | High voltage 2 fault (survey by comparators) | 1 | 1
HV12ADC | FaultHighVAdc | HV_ADC_FAULT | OVERVIEW |  |  | 25 | bool | :1 | bool blecsHV12ADC:1; | <field name="blecsHV12ADC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High voltage 1 or 2 fault (survey by ADC)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsHV12ADC"/> | pDev->blecsHV12ADC.set(pAcq->blecsHV12ADC); |  | High voltage 1 or 2 fault (survey by ADC) | 1 | 1
LVCC | FaultLowVComp | LV_COMP_FAULT | OVERVIEW |  |  | 24 | bool | :1 | bool blecsLVCC:1; | <field name="blecsLVCC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Low voltage fault with (survey by comparators)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsLVCC"/> | pDev->blecsLVCC.set(pAcq->blecsLVCC); |  | Low voltage fault with (survey by comparators) | 1 | 1
LVA | FaultLowVAdc | LV_ADC_FAULT | OVERVIEW |  |  | 23 | bool | :1 | bool blecsLVA:1; | <field name="blecsLVA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Low voltage fault (survey by ADC)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsLVA"/> | pDev->blecsLVA.set(pAcq->blecsLVA); |  | Low voltage fault (survey by ADC) | 1 | 1
BE | FaultBmEngy | BEAM_ENGY_FAULT | OVERVIEW |  |  | 22 | bool | :1 | bool blecsBE:1; | <field name="blecsBE" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Energy fault (From CISV card AB/CO)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBE"/> | pDev->blecsBE.set(pAcq->blecsBE); |  | Beam Energy fault (From CISV card AB/CO) | 1 | 1
TCKF | FaultTurnClk | TURN_CLK_FAULT | OVERVIEW |  |  | 21 | bool | :1 | bool blecsTCKF:1; | <field name="blecsTCKF" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Turn Clock fault (From BOBR card)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsTCKF"/> | pDev->blecsTCKF.set(pAcq->blecsTCKF); |  | Turn Clock fault (From BOBR card) | 1 | 1
DUMP | DumpOccured | DUMP_OCCURED | OVERVIEW |  |  | 20 | bool | :1 | bool blecsDUMP:1; | <field name="blecsDUMP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Dump occurred</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsDUMP"/> | pDev->blecsDUMP.set(pAcq->blecsDUMP); |  | Dump occurred | 1 | 1
UST | TstPendUsrSysTest | TEST_PEND_USR_SYS_TST | TESTS |  |  | 19 | bool | :1 | bool blecsUST:1; | <field name="blecsUST" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User System test pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUST"/> | pDev->blecsUST.set(pAcq->blecsUST); |  | User System test pending | 1 | 1
UCT | TstPendUsrConsist | TEST_PEND_USR_CONSIST | TESTS |  |  | 18 | bool | :1 | bool blecsUCT:1; | <field name="blecsUCT" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User Consistency test pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUCT"/> | pDev->blecsUCT.set(pAcq->blecsUCT); |  | User Consistency test pending | 1 | 1
UTBPL | TstPendUsrThresBpl | TEST_PEND_USR_THRES_BPL | TESTS |  |  | 17 | bool | :1 | bool blecsUTBPL:1; | <field name="blecsUTBPL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User Threshold to BPL pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUTBPL"/> | pDev->blecsUTBPL.set(pAcq->blecsUTBPL); |  | User Threshold to BPL pending | 1 | 1
UET | TstPendUsrEngy | TEST_PEND_USR_ENGY | TESTS |  |  | 16 | bool | :1 | bool blecsUET:1; | <field name="blecsUET" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User Energy Test pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUET"/> | pDev->blecsUET.set(pAcq->blecsUET); |  | User Energy Test pending | 1 | 1
UBPBIS | TstPendUsrBmPermBis | TEST_PEND_USR_BP_BIS | TESTS |  |  | 15 | bool | :1 | bool blecsUBPBIS:1; | <field name="blecsUBPBIS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>User BPBIS (BLECS to Beam interlock) pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUBPBIS"/> | pDev->blecsUBPBIS.set(pAcq->blecsUBPBIS); |  | User BPBIS (BLECS to Beam interlock) pending | 1 | 1 |  | second
EST | TstPendExpSysTest | TEST_PEND_EXP_SYS_TST | TESTS |  |  | 14 | bool | :1 | bool blecsEST:1; | <field name="blecsEST" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert System test pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEST"/> | pDev->blecsEST.set(pAcq->blecsEST); |  | Expert System test pending | 1 | 1 |  | every
ECT | TstPendExpConsist | TEST_PEND_EXP_CONSIST | TESTS |  |  | 13 | bool | :1 | bool blecsECT:1; | <field name="blecsECT" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Consistency test pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsECT"/> | pDev->blecsECT.set(pAcq->blecsECT); |  | Expert Consistency test pending | 1 | 1 |  | Read it
ETBPL | TstPendExpThresBpl | TEST_PEND_EXP_THRES_BP | TESTS |  |  | 12 | bool | :1 | bool blecsETBPL:1; | <field name="blecsETBPL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Threshold to BPL pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsETBPL"/> | pDev->blecsETBPL.set(pAcq->blecsETBPL); |  | Expert Threshold to BPL pending | 1 | 1
EET | TstPendExpEngy | TEST_PEND_EXP_ENGY | TESTS |  |  | 11 | bool | :1 | bool blecsEET:1; | <field name="blecsEET" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Energy Test pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEET"/> | pDev->blecsEET.set(pAcq->blecsEET); |  | Expert Energy Test pending | 1 | 1
EBPBIS | TstPendExpBmPermBis | TEST_PEND_EXP_BP_BIS | TESTS |  |  | 10 | bool | :1 | bool blecsEBPBIS:1; | <field name="blecsEBPBIS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert BPBIS (BLECS to Beam interlock) pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEBPBIS"/> | pDev->blecsEBPBIS.set(pAcq->blecsEBPBIS); |  | Expert BPBIS (BLECS to Beam interlock) pending | 1 | 1
EBPTC | TstPendExpBmPermTc | TEST_PEND_EXP_BP_TC | TESTS |  |  | 9 | bool | :1 | bool blecsEBPTC:1; | <field name="blecsEBPTC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert BPTC (Beam Permit BLETC to BLECS) pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEBPTC"/> | pDev->blecsEBPTC.set(pAcq->blecsEBPTC); |  | Expert BPTC (Beam Permit BLETC to BLECS) pending | 1 | 1
EHVLF | TstPendExpLowFreqModul | TEST_PEND_EXP_LF_MOD | TESTS |  |  | 8 | bool | :1 | bool blecsEHVLF:1; | <field name="blecsEHVLF" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVLF test pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEHVLF"/> | pDev->blecsEHVLF.set(pAcq->blecsEHVLF); |  | Expert HVLF test pending | 1 | 1
EHVCFC | TstPendExpCfcTest | TEST_PEND_EXP_CFC_TST | TESTS |  |  | 7 | bool | :1 | bool blecsEHVCFC:1; | <field name="blecsEHVCFC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVCFC (100 pA test) pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEHVCFC"/> | pDev->blecsEHVCFC.set(pAcq->blecsEHVCFC); |  | Expert HVCFC (100 pA test) pending | 1 | 1
EHVRDAC | TstPendExpRstDac | TEST_PEND_EXP_RST_DAC | TESTS |  |  | 6 | bool | :1 | bool blecsEHVRDAC:1; | <field name="blecsEHVRDAC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVRDAC (Dac Reset) pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEHVRDAC"/> | pDev->blecsEHVRDAC.set(pAcq->blecsEHVRDAC); |  | Expert HVRDAC (Dac Reset) pending | 1 | 1
EHVRGOH | TstPendExpRstGoh | TEST_PEND_EXP_RST_GOH | TESTS |  |  | 5 | bool | :1 | bool blecsEHVRGOH:1; | <field name="blecsEHVRGOH" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert HVRGOH (GOH Reset) pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEHVRGOH"/> | pDev->blecsEHVRGOH.set(pAcq->blecsEHVRGOH); |  | Expert HVRGOH (GOH Reset) pending | 1 | 1
EMC | TstPendExpManCtrl | TEST_PEND_EXP_MAN_CTRL | TESTS |  |  | 4 | bool | :1 | bool blecsEMC:1; | <field name="blecsEMC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Expert Manual control Request pending</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsEMC"/> | pDev->blecsEMC.set(pAcq->blecsEMC); |  | Expert Manual control Request pending | 1 | 1
TSTRN | TstPendTimerSysTestNorm | TEST_PEND_TMR_SYS_TST_NOR | TESTS |  |  | 2 | bool | :1 | bool blecsTSTRN:1; | <field name="blecsTSTRN" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Timer System test request normal</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsTSTRN"/> | pDev->blecsTSTRN.set(pAcq->blecsTSTRN); |  | Timer System test request normal | 1 | 1
TSTRP | TstPendTimerSysTestPrio | TEST_PEND_TMR_SYS_TST_PRI | TESTS |  |  | 1 | bool | :1 | bool blecsTSTRP:1; | <field name="blecsTSTRP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Timer System test request Priority (BP away next Dump)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsTSTRP"/> | pDev->blecsTSTRP.set(pAcq->blecsTSTRP); |  | Timer System test request Priority (BP away next Dump) | 1 | 1
AST | TstActiveSysTest | TEST_ACTIVE_SYS_TST | TESTS | 70000C | 700010 | 31 | bool | :1 | bool blecsAST:1; | <field name="blecsAST" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsAST"/> | pDev->blecsAST.set(pAcq->blecsAST); | Active tests overview | Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested) | 1 | 1 |  |  |  | 7 | if (vme_addr[22..0] >= h"70000C") and (vme_addr[22..0] < h"700010") then mux_sel[] = 7; end if; --Active tests overview / comments: Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested)
ACT | TstActiveConsist | TEST_ACTIVE_CONSIST | TESTS |  |  | 30 | bool | :1 | bool blecsACT:1; | <field name="blecsACT" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : CONSISTENCY (threshold table control)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsACT"/> | pDev->blecsACT.set(pAcq->blecsACT); |  | Active test : CONSISTENCY (threshold table control) | 1 | 1
ATBPL | TstActiveThresBpl | TEST_ACTIVE_THRES_BP | TESTS |  |  | 29 | bool | :1 | bool blecsATBPL:1; | <field name="blecsATBPL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : THRESHOLD to BPL</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsATBPL"/> | pDev->blecsATBPL.set(pAcq->blecsATBPL); |  | Active test : THRESHOLD to BPL | 1 | 1
AET | TstActiveEngy | TEST_ACTIVE_ENGY | TESTS |  |  | 28 | bool | :1 | bool blecsAET:1; | <field name="blecsAET" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : ENERGY TEST</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsAET"/> | pDev->blecsAET.set(pAcq->blecsAET); |  | Active test : ENERGY TEST | 1 | 1
ABPBIS | TstActiveBmPermBis | TEST_ACTIVE_BP_BIS | TESTS |  |  | 27 | bool | :1 | bool blecsABPBIS:1; | <field name="blecsABPBIS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsABPBIS"/> | pDev->blecsABPBIS.set(pAcq->blecsABPBIS); |  | Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 1 | 1
ABPTC | TstActiveBmPermTc | TEST_ACTIVE_BP_TC | TESTS |  |  | 26 | bool | :1 | bool blecsABPTC:1; | <field name="blecsABPTC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsABPTC"/> | pDev->blecsABPTC.set(pAcq->blecsABPTC); |  | Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 1 | 1
AHVLF | TstActiveLowFreqModul | TEST_ACTIVE_LF_MOD | TESTS |  |  | 25 | bool | :1 | bool blecsAHVLF:1; | <field name="blecsAHVLF" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : HVLF (Modulation of the high voltage and analysis)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsAHVLF"/> | pDev->blecsAHVLF.set(pAcq->blecsAHVLF); |  | Active test : HVLF (Modulation of the high voltage and analysis) | 1 | 1
AHVCFC | TstActiveCfcTest | TEST_ACTIVE_CFC_TST | TESTS |  |  | 24 | bool | :1 | bool blecsAHVCFC:1; | <field name="blecsAHVCFC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsAHVCFC"/> | pDev->blecsAHVCFC.set(pAcq->blecsAHVCFC); |  | Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card) | 1 | 1
AHVRDAC | TstActiveRstDac | TEST_ACTIVE_RST_DAC | TESTS |  |  | 23 | bool | :1 | bool blecsAHVRDAC:1; | <field name="blecsAHVRDAC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : HVRDAC (Activation DAC reset mode of the tunnel card)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsAHVRDAC"/> | pDev->blecsAHVRDAC.set(pAcq->blecsAHVRDAC); |  | Active test : HVRDAC (Activation DAC reset mode of the tunnel card) | 1 | 1
AHVRGOH | TstActiveRstGoh | TEST_ACTIVE_RST_GOH | TESTS |  |  | 22 | bool | :1 | bool blecsAHVRGOH:1; | <field name="blecsAHVRGOH" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : HVRGOH (Activation GOH reset mode of the tunnel card)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsAHVRGOH"/> | pDev->blecsAHVRGOH.set(pAcq->blecsAHVRGOH); |  | Active test : HVRGOH (Activation GOH reset mode of the tunnel card) | 1 | 1
MCAL | TstAllowedManCtrl | TEST_ALLOWED_MAN_CTRL | OVERVIEW |  |  | 21 | bool | :1 | bool blecsMCAL:1; | <field name="blecsMCAL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : MANUAL CONTROL ALLOWED</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsMCAL"/> | pDev->blecsMCAL.set(pAcq->blecsMCAL); |  | Active test : MANUAL CONTROL ALLOWED | 1 | 1
MCAC | TstActiveManCtrl | TEST_ACTIVE_MAN_CTRL | Overview + Tests |  |  | 20 | bool | :1 | bool blecsMCAC:1; | <field name="blecsMCAC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Active test : MANUAL CONTROL ACTIVE</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsMCAC"/> | pDev->blecsMCAC.set(pAcq->blecsMCAC); |  | Active test : MANUAL CONTROL ACTIVE | 1 | 1
RSTP | TstPassedSysTst | TEST_PASSED_SYS_TST | TESTS | 700014 | 700018 | 31 | bool | :1 | bool blecsRSTP:1; | <field name="blecsRSTP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : "SYSTEM TEST"</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsRSTP"/> | pDev->blecsRSTP.set(pAcq->blecsRSTP); | Result of last tests | Test Passed : "SYSTEM TEST" | 1 | 1 |  |  |  | 9 | if (vme_addr[22..0] >= h"700014") and (vme_addr[22..0] < h"700018") then mux_sel[] = 9; end if; --Result of last tests / comments: Test Passed : "SYSTEM TEST"
RCTP | TstPassedConsist | TEST_PASSED_CONSIST | TESTS |  |  | 30 | bool | :1 | bool blecsRCTP:1; | <field name="blecsRCTP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : CONSISTENCY (threshold table control)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsRCTP"/> | pDev->blecsRCTP.set(pAcq->blecsRCTP); |  | Test Passed : CONSISTENCY (threshold table control) | 1 | 1
RTBPLP | TstPassedThresBpl | TEST_PASSED_THRES_BPL | TESTS |  |  | 29 | bool | :1 | bool blecsRTBPLP:1; | <field name="blecsRTBPLP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : THRESHOLD to BPL</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsRTBPLP"/> | pDev->blecsRTBPLP.set(pAcq->blecsRTBPLP); |  | Test Passed : THRESHOLD to BPL | 1 | 1
RETP | TstPassedEngy | TEST_PASSED_ENGY | TESTS |  |  | 28 | bool | :1 | bool blecsRETP:1; | <field name="blecsRETP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : ENERGY TEST</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsRETP"/> | pDev->blecsRETP.set(pAcq->blecsRETP); |  | Test Passed : ENERGY TEST | 1 | 1
RBPBISP | TstPassedBmPermBis | TEST_PASSED_BP_BIS | TESTS |  |  | 27 | bool | :1 | bool blecsRBPBISP:1; | <field name="blecsRBPBISP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsRBPBISP"/> | pDev->blecsRBPBISP.set(pAcq->blecsRBPBISP); |  | Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 1 | 1
RBPTCP | TstPassedBmPermTc | TEST_PASSED_BP_TC | TESTS |  |  | 26 | bool | :1 | bool blecsRBPTCP:1; | <field name="blecsRBPTCP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsRBPTCP"/> | pDev->blecsRBPTCP.set(pAcq->blecsRBPTCP); |  | Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 1 | 1 |  | second
RHVLFP | TstPassedLowFreqModul | TEST_PASSED_LF_MOD | TESTS |  |  | 25 | bool | :1 | bool blecsRHVLFP:1; | <field name="blecsRHVLFP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : HVLF (Modulation of the high voltage and analysis)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsRHVLFP"/> | pDev->blecsRHVLFP.set(pAcq->blecsRHVLFP); |  | Test Passed : HVLF (Modulation of the high voltage and analysis) | 1 | 1 |  | every
RHVTCFCP | TstPassedCfcTest | TEST_PASSED_CFC_TEST | TESTS |  |  | 24 | bool | :1 | bool blecsRHVTCFCP:1; | <field name="blecsRHVTCFCP" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsRHVTCFCP"/> | pDev->blecsRHVTCFCP.set(pAcq->blecsRHVTCFCP); |  | Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card) | 1 | 1 |  | Read it
TRR | TstResRdy | TEST_RESULT_RDY | TESTS |  |  | 0 | bool | :1 | bool blecsTRR:1; | <field name="blecsTRR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>New test results are available (Check active test to find out which test)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsTRR"/> | pDev->blecsTRR.set(pAcq->blecsTRR); |  | New test results are available (Check active test to find out which test) | 1 | 1
UDL | DumpLineUmsk | BEAM_PERM_TC_LINE_U | Beam PERMIT | 70001C | 700020 | 31 | bool | :1 | bool blecsUDL:1; | <field name="blecsUDL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold Comparators Unmaskable beam permit Line</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUDL"/> | pDev->blecsUDL.set(pAcq->blecsUDL); | User Beam Permit | Threshold Comparators Unmaskable beam permit Line | 1 | 1 |  |  |  | 11 | if (vme_addr[22..0] >= h"70001C") and (vme_addr[22..0] < h"700020") then mux_sel[] = 11; end if; --User Beam Permit / comments: Threshold Comparators Unmaskable beam permit Line
MDL | DumpLineMsk | BEAM_PERM_TC_LINE_M | Beam PERMIT |  |  | 30 | bool | :1 | bool blecsMDL:1; | <field name="blecsMDL" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Threshold Comparators Maskable beam permit Line</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsMDL"/> | pDev->blecsMDL.set(pAcq->blecsMDL); |  | Threshold Comparators Maskable beam permit Line | 1 | 1
UAI | BmPermUmskInA | BEAM_PERM_U_IN_A | Beam PERMIT |  |  | 29 | bool | :1 | bool blecsUAI:1; | <field name="blecsUAI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Unmaskable A Input from the upper Combiner</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUAI"/> | pDev->blecsUAI.set(pAcq->blecsUAI); |  | Beam Permit Unmaskable A Input from the upper Combiner | 1 | 1
UBI | BmPermUmskInB | BEAM_PERM_U_IN_B | Beam PERMIT |  |  | 28 | bool | :1 | bool blecsUBI:1; | <field name="blecsUBI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Unmaskable B Input from the upper Combiner</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUBI"/> | pDev->blecsUBI.set(pAcq->blecsUBI); |  | Beam Permit Unmaskable B Input from the upper Combiner | 1 | 1
MAI | BmPermMskInA | BEAM_PERM_M_IN_A | Beam PERMIT |  |  | 27 | bool | :1 | bool blecsMAI:1; | <field name="blecsMAI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Maskable A Input from the upper Combiner</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsMAI"/> | pDev->blecsMAI.set(pAcq->blecsMAI); |  | Beam Permit Maskable A Input from the upper Combiner | 1 | 1
MBI | BmPermMskInB | BEAM_PERM_M_IN_B | Beam PERMIT |  |  | 26 | bool | :1 | bool blecsMBI:1; | <field name="blecsMBI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Maskable B Input from the upper Combiner</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsMBI"/> | pDev->blecsMBI.set(pAcq->blecsMBI); |  | Beam Permit Maskable B Input from the upper Combiner | 1 | 1
UAO | BmPermUmskOutA | BEAM_PERM_U_OUT_A | Beam PERMIT |  |  | 25 | bool | :1 | bool blecsUAO:1; | <field name="blecsUAO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Unmaskable A Output to lower Combiner (or CIBUS)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUAO"/> | pDev->blecsUAO.set(pAcq->blecsUAO); |  | Beam Permit Unmaskable A Output to lower Combiner (or CIBUS) | 1 | 1
UBO | BmPermUmskOutB | BEAM_PERM_U_OUT_B | Beam PERMIT |  |  | 24 | bool | :1 | bool blecsUBO:1; | <field name="blecsUBO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Unmaskable B Output to lower Combiner (or CIBUS)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUBO"/> | pDev->blecsUBO.set(pAcq->blecsUBO); |  | Beam Permit Unmaskable B Output to lower Combiner (or CIBUS) | 1 | 1
MAO | BmPermMskOutA | BEAM_PERM_M_OUT_A | Beam PERMIT |  |  | 23 | bool | :1 | bool blecsMAO:1; | <field name="blecsMAO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Maskable A Output to lower Combiner (or CIBUS)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsMAO"/> | pDev->blecsMAO.set(pAcq->blecsMAO); |  | Beam Permit Maskable A Output to lower Combiner (or CIBUS) | 1 | 1
MBO | BmPermMskOutB | BEAM_PERM_M_OUT_B | Beam PERMIT |  |  | 22 | bool | :1 | bool blecsMBO:1; | <field name="blecsMBO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Permit Maskable B Output to lower Combiner (or CIBUS)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsMBO"/> | pDev->blecsMBO.set(pAcq->blecsMBO); |  | Beam Permit Maskable B Output to lower Combiner (or CIBUS) | 1 | 1
UBS | BmInfoUmsk | BEAM_PERM_INFO_U | Beam PERMIT |  |  | 21 | bool | :1 | bool blecsUBS:1; | <field name="blecsUBS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Unmaskable Beam Info (Beam status)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUBS"/> | pDev->blecsUBS.set(pAcq->blecsUBS); |  | Unmaskable Beam Info (Beam status) | 1 | 1
MBS | BmInfoMsk | BEAM_PERM_INFO_M | Beam PERMIT |  |  | 20 | bool | :1 | bool blecsMBS:1; | <field name="blecsMBS" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Maskable Beam Info (Beam status)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsMBS"/> | pDev->blecsMBS.set(pAcq->blecsMBS); |  | Maskable Beam Info (Beam status) | 1 | 1
UBSR | BmInfoHadRespToDumpUmsk | BEAM_PERM_CIBUS_U_FAULT | Beam PERMIT |  |  | 19 | bool | :1 | bool blecsUBSR:1; | <field name="blecsUBSR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsUBSR"/> | pDev->blecsUBSR.set(pAcq->blecsUBSR); |  | Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP | 1 | 1
MBSR | BmInfoHadRespToDumpMsk | BEAM_PERM_CIBUS_M_FAULT | Beam PERMIT |  |  | 18 | bool | :1 | bool blecsMBSR:1; | <field name="blecsMBSR" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Maskable Beam Info (beam status) had not respond to last BLECS-DUMP</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsMBSR"/> | pDev->blecsMBSR.set(pAcq->blecsMBSR); |  | Maskable Beam Info (beam status) had not respond to last BLECS-DUMP | 1 | 1
TCLD | TurnCntWhenDumpReq | TURN_CNT_WN_DUMP_REQ | Beam PERMIT | 700024 | 700028 | 0..31 | unsigned | :32 | unsigned blecsTCLD:32; | <field name="blecsTCLD" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Turn counter value when the last dump was requested (32 bits )</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTCLD"/> | pDev->blecsTCLD.set(pAcq->blecsTCLD); | Turn counter when last dump request | Turn counter value when the last dump was requested (32 bits ) | 1 | 32 |  |  |  | 13 | if (vme_addr[22..0] >= h"700024") and (vme_addr[22..0] < h"700028") then mux_sel[] = 13; end if; --Turn counter when last dump request / comments: Turn counter value when the last dump was requested (32 bits )
BCLD | BunchCntWhenDumpReq | BUNCH_CNT_WN_DUMP_REQ | Beam PERMIT |  |  | 0..11 | unsigned | :12 | unsigned blecsBCLD:12; | <field name="blecsBCLD" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Bunch counter value when the last dump was requested (12 bits)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBCLD"/> | pDev->blecsBCLD.set(pAcq->blecsBCLD); |  | Bunch counter value when the last dump was requested (12 bits) | 1 | 12
TDPM | TimeDumpToPmfreezeMsk | BEAM_PERM_TIME_TO_PM_FRZE_M | Beam PERMIT | 70002C | 700030 | 31..16 | unsigned | :16 | unsigned blecsTDPM:16; | <field name="blecsTDPM" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Time from the dump request maskable to PM_freeze (microsecond resolution)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTDPM"/> | pDev->blecsTDPM.set(pAcq->blecsTDPM); | Time from dump to PM freeze | Time from the dump request maskable to PM_freeze (microsecond resolution) | 1 | 16 |  |  |  | 15 | if (vme_addr[22..0] >= h"70002C") and (vme_addr[22..0] < h"700030") then mux_sel[] = 15; end if; --Time from dump to PM freeze / comments: Time from the dump request maskable to PM_freeze (microsecond resolution)
TDPU | TimeDumpToPmfreezeUmsk | BEAM_PERM_TIME_TO_PM_FRZE_U | Beam PERMIT |  |  | 15..0 | unsigned | :16 | unsigned blecsTDPU:16; | <field name="blecsTDPU" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Time from the dump request unmaskable to PM_freeze (microsecond resolution)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTDPU"/> | pDev->blecsTDPU.set(pAcq->blecsTDPU); |  | Time from the dump request unmaskable to PM_freeze (microsecond resolution) | 1 | 16
BEI | BmEngyIn | BEAM_ENGY_IN | Beam ENERGY | 700030 | 700034 | 16..31 | unsigned | :16 | unsigned blecsBEI:16; | <field name="blecsBEI" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Energy Input coming from timing distribution (CISV card AB/CO)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBEI"/> | pDev->blecsBEI.set(pAcq->blecsBEI); | Beam Energy 1 | Beam Energy Input coming from timing distribution (CISV card AB/CO) | 1 | 16 |  |  |  | 16 | if (vme_addr[22..0] >= h"700030") and (vme_addr[22..0] < h"700034") then mux_sel[] = 16; end if; --Beam Energy 1 / comments: Beam Energy Input coming from timing distribution (CISV card AB/CO)
BEO | BmEngyOut | BEAM_ENGY_LEVEL_OUT | Beam ENERGY |  |  | 11..15 | unsigned | :5 | unsigned blecsBEO:5; | <field name="blecsBEO" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Energy levels output to the Threshold Comparators (send to BLETC cards)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBEO"/> | pDev->blecsBEO.set(pAcq->blecsBEO); |  | Beam Energy levels output to the Threshold Comparators (send to BLETC cards) | 1 | 5 |  | second
BEE | BmEngyErr | BEAM_ENGY_ERR | Beam ENERGY |  |  | 10 | bool | :1 | bool blecsBEE:1; | <field name="blecsBEE" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam Energy error (send to BLETC cards)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsBEE"/> | pDev->blecsBEE.set(pAcq->blecsBEE); |  | Beam Energy error (send to BLETC cards) | 1 | 1 |  | every
SRTC | BmEngySoftRstTC | BEAM_ENGY_SOFT_RST_TC | Beam ENERGY |  |  | 9 | bool | :1 | bool blecsSRTC:1; | <field name="blecsSRTC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>SoftResetTC (send to BLETC cards)</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsSRTC"/> | pDev->blecsSRTC.set(pAcq->blecsSRTC); |  | SoftResetTC (send to BLETC cards) | 1 | 1 |  | Read it
BECRCA | BmEngyCrcCntA | BEAM_ENGY_CRC_CNT_A | Beam ENERGY | 700038 | 70003C | 16..31 | unsigned | :16 | unsigned blecsBECRCA:16; | <field name="blecsBECRCA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam energy reception CRC error counter channel A</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBECRCA"/> | pDev->blecsBECRCA.set(pAcq->blecsBECRCA); | Beam Energy 3 | Beam energy reception CRC error counter channel A | 1 | 16 |  |  |  | 18 | if (vme_addr[22..0] >= h"700038") and (vme_addr[22..0] < h"70003C") then mux_sel[] = 18; end if; --Beam Energy 3 / comments: Beam energy reception CRC error counter channel A
BECRCB | BmEngyCrcCntB | BEAM_ENGY_CRC_CNT_B | Beam ENERGY |  |  | 0..15 | unsigned | :16 | unsigned blecsBECRCB:16; | <field name="blecsBECRCB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam energy reception CRC error counter channel B</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBECRCB"/> | pDev->blecsBECRCB.set(pAcq->blecsBECRCB); |  | Beam energy reception CRC error counter channel B | 1 | 16
BELFA | BmEngyLstFrmCntA | BEAM_ENGY_LOSTF_CNT_A | Beam ENERGY | 70003C | 700040 | 16..31 | unsigned | :16 | unsigned blecsBELFA:16; | <field name="blecsBELFA" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam energy reception Lost Frame counter channel A</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBELFA"/> | pDev->blecsBELFA.set(pAcq->blecsBELFA); | Beam Energy 4 | Beam energy reception Lost Frame counter channel A | 1 | 16 |  |  |  | 19 | if (vme_addr[22..0] >= h"70003C") and (vme_addr[22..0] < h"700040") then mux_sel[] = 19; end if; --Beam Energy 4 / comments: Beam energy reception Lost Frame counter channel A
BELFB | BmEngyLstFrmCntB | BEAM_ENGY_LOSTF_CNT_B | Beam ENERGY |  |  | 0..15 | unsigned | :16 | unsigned blecsBELFB:16; | <field name="blecsBELFB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam energy reception Lost Frame counter channel B</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBELFB"/> | pDev->blecsBELFB.set(pAcq->blecsBELFB); |  | Beam energy reception Lost Frame counter channel B | 1 | 16
BETOV | BmEngyTimeOut | BEAM_ENGY_TIME_OUT | Beam ENERGY | 700040 | 700044 | 28..31 | unsigned | :4 | unsigned blecsBETOV:4; | <field name="blecsBETOV" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Beam energy reception Time Out Beam Energy Value</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBETOV"/> | pDev->blecsBETOV.set(pAcq->blecsBETOV); | Beam Energy 5 | Beam energy reception Time Out Beam Energy Value | 1 | 16 |  |  |  | 20 | if (vme_addr[22..0] >= h"700040") and (vme_addr[22..0] < h"700044") then mux_sel[] = 20; end if; --Beam Energy 5 / comments: Beam energy reception Time Out Beam Energy Value
V1H | HighV1VHigh | HV_1_VOLT_HIGH |  | 700044 | 700048 | 31 | bool | :1 | bool blecsV1H:1; | <field name="blecsV1H" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / voltage higher</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsV1H"/> | pDev->blecsV1H.set(pAcq->blecsV1H); | High Voltage 1 survey | High Voltage power supply 1 / voltage higher | 1 | 1 |  |  |  | 21 | if (vme_addr[22..0] >= h"700044") and (vme_addr[22..0] < h"700048") then mux_sel[] = 21; end if; --High Voltage 1 survey / comments: High Voltage power supply 1 / voltage higher
V1L | HighV1VLow | HV_1_VOLT_LOW |  |  |  | 30 | bool | :1 | bool blecsV1L:1; | <field name="blecsV1L" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / voltage lower</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsV1L"/> | pDev->blecsV1L.set(pAcq->blecsV1L); |  | High Voltage power supply 1 / voltage lower | 1 | 1
I1H | HighV1IHigh | HV_1_CURR_HIGH |  |  |  | 29 | bool | :1 | bool blecsI1H:1; | <field name="blecsI1H" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / current higher</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsI1H"/> | pDev->blecsI1H.set(pAcq->blecsI1H); |  | High Voltage power supply 1 / current higher | 1 | 1
I1L | HighV1ILow | HV_1_CURR_LOW |  |  |  | 28 | bool | :1 | bool blecsI1L:1; | <field name="blecsI1L" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / current lower</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsI1L"/> | pDev->blecsI1L.set(pAcq->blecsI1L); |  | High Voltage power supply 1 / current lower | 1 | 1
HV1V | HighV1VValue | HV_1_VOLT_VALUE |  |  |  | 0..23 | unsigned | :24 | unsigned blecsHV1V:24; | <field name="blecsHV1V" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / voltage value (24 bits) [HV = 3000/2**24*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV1V"/> | pDev->blecsHV1V.set(pAcq->blecsHV1V); | High Voltage 2 survey | High Voltage power supply 1 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 1 | 24 | use filter |  | 3000/2**24*code
V2H | HighV2VHigh | HV_2_VOLT_HIGH | Survey | 700048 | 70004C | 31 | bool | :1 | bool blecsV2H:1; | <field name="blecsV2H" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / voltage higher</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsV2H"/> | pDev->blecsV2H.set(pAcq->blecsV2H); |  | High Voltage power supply 2 / voltage higher | 1 | 1 |  |  |  | 22 | if (vme_addr[22..0] >= h"700048") and (vme_addr[22..0] < h"70004C") then mux_sel[] = 22; end if; -- / comments: High Voltage power supply 2 / voltage higher
V2L | HighV2VLow | HV_2_VOLT_LOW |  |  |  | 30 | bool | :1 | bool blecsV2L:1; | <field name="blecsV2L" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / voltage lower</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsV2L"/> | pDev->blecsV2L.set(pAcq->blecsV2L); |  | High Voltage power supply 2 / voltage lower | 1 | 1
I2H | HighV2IHigh | HV_2_CURR_HIGH |  |  |  | 29 | bool | :1 | bool blecsI2H:1; | <field name="blecsI2H" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / current higher</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsI2H"/> | pDev->blecsI2H.set(pAcq->blecsI2H); |  | High Voltage power supply 2 / current higher | 1 | 1
I2L | HighV2ILow | HV_2_CURR_LOW |  |  |  | 28 | bool | :1 | bool blecsI2L:1; | <field name="blecsI2L" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / current lower</description> <scalar type="bool"/> </field> | <data-field-ref-item field-name-ref="blecsI2L"/> | pDev->blecsI2L.set(pAcq->blecsI2L); |  | High Voltage power supply 2 / current lower | 1 | 1
HV2V | HighV2VValue | HV_2_VOLT_VALUE |  |  |  | 0..23 | unsigned | :24 | unsigned blecsHV2V:24; | <field name="blecsHV2V" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV2V"/> | pDev->blecsHV2V.set(pAcq->blecsHV2V); |  | High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 1 | 24 | use filter |  | 3000/2**24*code
HV1I | HighV1IValue | HV_1_CURR_VALUE |  | 700054 | 700058 | 16..31 | unsigned | :16 | unsigned blecsHV1I:16; | <field name="blecsHV1I" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV1I"/> | pDev->blecsHV1I.set(pAcq->blecsHV1I); | High Voltage Currents | High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA | 1 | 16 | use filter |  | 20/2**16*code | 25 | if (vme_addr[22..0] >= h"700054") and (vme_addr[22..0] < h"700058") then mux_sel[] = 25; end if; --High Voltage Currents / comments: High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA
HV2I | HighV2IValue | HV_2_CURR_VALUE |  |  |  | 0..15 | unsigned | :16 | unsigned blecsHV2I:16; | <field name="blecsHV2I" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHV2I"/> | pDev->blecsHV2I.set(pAcq->blecsHV2I); |  | High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA | 1 | 16 | use filter |  | 20/2**16*code
TCFUAB | TurnCntWhenfallUAB | TURN_CNT_WN_FALL_U_AORB | Beam PERMIT | 70008C | 700090 | 0..31 | unsigned | :32 | unsigned blecsTCFUAB:32; | <field name="blecsTCFUAB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Turn counter when last fall of UA or UB (32 bits)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTCFUAB"/> | pDev->blecsTCFUAB.set(pAcq->blecsTCFUAB); | Turn counter when last fall of UA or UB | Turn counter when last fall of UA or UB (32 bits) | 1 | 32 |  |  |  | 39 | if (vme_addr[22..0] >= h"70008C") and (vme_addr[22..0] < h"700090") then mux_sel[] = 39; end if; --Turn counter when last fall of UA or UB / comments: Turn counter when last fall of UA or UB (32 bits)
TCFMAB | TurnCntWhenfallMAB | TURN_CNT_WN_FALL_M_AORB | Beam PERMIT | 700090 | 700094 | 0..31 | unsigned | :32 | unsigned blecsTCFMAB:32; | <field name="blecsTCFMAB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Turn counter when last fall of MA or MB (32 bits)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTCFMAB"/> | pDev->blecsTCFMAB.set(pAcq->blecsTCFMAB); | Turn counter when last fall of MA or MB | Turn counter when last fall of MA or MB (32 bits) | 1 | 32 |  |  |  | 40 | if (vme_addr[22..0] >= h"700090") and (vme_addr[22..0] < h"700094") then mux_sel[] = 40; end if; --Turn counter when last fall of MA or MB / comments: Turn counter when last fall of MA or MB (32 bits)
BCFMAB | BunchCntWhenfallMAB | BUNCH_CNT_WN_FALL_M_AORB | Beam PERMIT |  |  | 12..23 | unsigned | :12 | unsigned blecsBCFMAB:12; | <field name="blecsBCFMAB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Bunch counter when last fall of MA or MB 12 bits counter</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBCFMAB"/> | pDev->blecsBCFMAB.set(pAcq->blecsBCFMAB); |  | Bunch counter when last fall of MA or MB 12 bits counter | 1 | 16
BCFUAB | BunchCntWhenfallUAB | BUNCH_CNT_WN_FALL_U_AORB | Beam PERMIT |  |  | 0..11 | unsigned | :12 | unsigned blecsBCFUAB:12; | <field name="blecsBCFUAB" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Bunch counter when last fall of UA or UB 12 bits counter</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBCFUAB"/> | pDev->blecsBCFUAB.set(pAcq->blecsBCFUAB); |  | Bunch counter when last fall of UA or UB 12 bits counter | 1 | 16
BLETCPRES | - | TEST_RES_BLETC_PRES | TESTS |  |  | 0..15 | unsigned | :16 | unsigned blecsBLETCPRES:16; | <field name="blecsBLETCPRES" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Vector of BLETC Board present in the crate (MSBit TC1 LSBit TC16)</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBLETCPRES"/> | pDev->blecsBLETCPRES.set(pAcq->blecsBLETCPRES); |  | Vector of BLETC Board present in the crate (MSBit TC1 LSBit TC16) | 1 | 16
BLECFPRES | - | TEST_RES_BLECF_PRES | BLECF TEST | 700204 | 700208 | 31..0 | unsigned | :32 | unsigned blecsBLECFPRES:32; | <field name="blecsBLECFPRES" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Vector of BLECF Board present in the crate (MSBit BLECF1(TC1) LSBit BLECF32 (TC16))</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsBLECFPRES"/> | pDev->blecsBLECFPRES.set(pAcq->blecsBLECFPRES); | General tests info BLECF Board present | Vector of BLECF Board present in the crate (MSBit BLECF1(TC1) LSBit BLECF32 (TC16)) | 1 | 32 |  |  |  | 44 | if (vme_addr[22..0] >= h"700204") and (vme_addr[22..0] < h"700208") then mux_sel[] = 44; end if; --General tests info BLECF Board present / comments: Vector of BLECF Board present in the crate (MSBit BLECF1(TC1) LSBit BLECF32 (TC16))
TRCBLETC | - | TEST_RES_CONSIST_PER_BLETC | TESTS |  |  | 0..15 | unsigned | :16 | unsigned blecsTRCBLETC:16; | <field name="blecsTRCBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the Consistency test per BLETC (MSBit TC1) 1=passed</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTRCBLETC"/> | pDev->blecsTRCBLETC.set(pAcq->blecsTRCBLETC); |  | Result of the Consistency test per BLETC (MSBit TC1) 1=passed | 1 | 16 |  | when TRR=1
TRBPUBLETC | - | TEST_RES_BP_TC_PER_BLETC_U | TESTS | 700244 | 700248 | 16..31 | unsigned | :16 | unsigned blecsTRBPUBLETC:16; | <field name="blecsTRBPUBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTRBPUBLETC"/> | pDev->blecsTRBPUBLETC.set(pAcq->blecsTRBPUBLETC); | Tests result per BLETC 2 | Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed | 1 | 16 |  | specific location |  | 46 | if (vme_addr[22..0] >= h"700244") and (vme_addr[22..0] < h"700248") then mux_sel[] = 46; end if; --Tests result per BLETC 2 / comments: Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed
TRBPMBLETC | - | TEST_RES_BP_TC_PER_BLETC_M | TESTS |  |  | 0..15 | unsigned | :16 | unsigned blecsTRBPMBLETC:16; | <field name="blecsTRBPMBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the BPLTC test maskable line (MSBit TC1) 1=passed</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTRBPMBLETC"/> | pDev->blecsTRBPMBLETC.set(pAcq->blecsTRBPMBLETC); |  | Result of the BPLTC test maskable line (MSBit TC1) 1=passed | 1 | 16 |  | write it into
TRHVBLETC | - | TEST_RES_LF_MOD_PER_BLETC | TESTS |  |  | 0..15 | unsigned | :16 | unsigned blecsTRHVBLETC:16; | <field name="blecsTRHVBLETC" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsTRHVBLETC"/> | pDev->blecsTRHVBLETC.set(pAcq->blecsTRHVBLETC); |  | Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed | 1 | 16 |  | otherwise
HVCFCRES | - | TEST_RES_LF_MOD_PER_BLECF | BLECF TEST | 70024C | 700250 | 31..0 | unsigned | :32 | unsigned blecsHVCFCRES:32; | <field name="blecsHVCFCRES" multiplexing-criterion="NONE" persistency="VOLATILE"> <description>Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed</description> <scalar type="long"/> </field> | <data-field-ref-item field-name-ref="blecsHVCFCRES"/> | pDev->blecsHVCFCRES.set(pAcq->blecsHVCFCRES); | HVCFC result per BLECF | Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed | 1 | 32 |  |  |  | 48 | if (vme_addr[22..0] >= h"70024C") and (vme_addr[22..0] < h"700250") then mux_sel[] = 48; end if; --HVCFC result per BLECF / comments: Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed

## LOGGING
LOGGING
FESA Device | all devices
FESA Cycle | ALL
FESA Property | Acquisition
FESA Name | splitting operation | Database name | Logging datatype | units | Entity type | data type | Bit field | Description | STATUS | total in bits | concatenation
blecsSYS_OP |  | SYS_OPERATIONAL | NUMERIC | NO_UNIT | crate | bool | :1 | System is operational (View by the BLECS) | 1 | 1 | CRATE
blecsTEST |  | SYS_UNDER_TEST | NUMERIC | NO_UNIT | crate | bool | :1 | System is under test (All BLM crate of the IP are involved but only this crate is tested) | 1 | 1 | CRATE
blecsT_REQ |  | TEST_PEND_GLOBAL | NUMERIC | NO_UNIT | crate | bool | :1 | Test requested (while take place when the conditions are fullfill) | 1 | 1 | CRATE
blecsHV1 |  | HV_1_COMP_FAULT | NUMERIC | NO_UNIT | crate | bool | :1 | High voltage 1 fault (survey by comparators) | 1 | 1 | CRATE
blecsHV2 |  | HV_2_COMP_FAULT | NUMERIC | NO_UNIT | crate | bool | :1 | High voltage 2 fault (survey by comparators) | 1 | 1 | CRATE
blecsHV12ADC |  | HV_ADC_FAULT | NUMERIC | NO_UNIT | crate | bool | :1 | High voltage 1 or 2 fault (survey by ADC) | 1 | 1 | CRATE
blecsLVCC |  | LV_COMP_FAULT | NUMERIC | NO_UNIT | crate | bool | :1 | Low voltage fault with (survey by comparators) | 1 | 1 | CRATE
blecsLVA |  | LV_ADC_FAULT | NUMERIC | NO_UNIT | crate | bool | :1 | Low voltage fault (survey by ADC) | 1 | 1 | CRATE
blecsBE |  | BEAM_ENGY_FAULT | NUMERIC | NO_UNIT | crate | bool | :1 | Beam Energy fault (From CISV card AB/CO) | 1 | 1 | CRATE
blecsTCKF |  | TURN_CLK_FAULT | NUMERIC | NO_UNIT | crate | bool | :1 | Turn Clock fault (From BOBR card) | 1 | 1 | CRATE
blecsDUMP |  | DUMP_OCCURED | NUMERIC | NO_UNIT | crate | bool | :1 | Dump occurred | 1 | 1 | CRATE
blecsUST |  | TEST_PEND_USR_SYS_TST | NUMERIC | NO_UNIT | crate | bool | :1 | User System test pending | 1 | 1 | CRATE
blecsUCT |  | TEST_PEND_USR_CONSIST | NUMERIC | NO_UNIT | crate | bool | :1 | User Consistency test pending | 1 | 1 | CRATE
blecsUBPBIS |  | TEST_PEND_USR_BP_BIS | NUMERIC | NO_UNIT | crate | bool | :1 | User BPBIS (BLECS to Beam interlock) pending | 1 | 1 | CRATE
blecsEST |  | TEST_PEND_EXP_SYS_TST | NUMERIC | NO_UNIT | crate | bool | :1 | Expert System test pending | 1 | 1 | CRATE
blecsECT |  | TEST_PEND_EXP_CONSIST | NUMERIC | NO_UNIT | crate | bool | :1 | Expert Consistency test pending | 1 | 1 | CRATE
blecsETBPL |  | TEST_PEND_EXP_THRES_BP | NUMERIC | NO_UNIT | crate | bool | :1 | Expert Threshold to BPL pending | 1 | 1 | CRATE
blecsEET |  | TEST_PEND_EXP_ENGY | NUMERIC | NO_UNIT | crate | bool | :1 | Expert Energy Test pending | 1 | 1 | CRATE
blecsEBPBIS |  | TEST_PEND_EXP_BP_BIS | NUMERIC | NO_UNIT | crate | bool | :1 | Expert BPBIS (BLECS to Beam interlock) pending | 1 | 1 | CRATE
blecsEBPTC |  | TEST_PEND_EXP_BP_TC | NUMERIC | NO_UNIT | crate | bool | :1 | Expert BPTC (Beam Permit BLETC to BLECS) pending | 1 | 1 | CRATE
blecsEHVLF |  | TEST_PEND_EXP_LF_MOD | NUMERIC | NO_UNIT | crate | bool | :1 | Expert HVLF test pending | 1 | 1 | CRATE
blecsEHVCFC |  | TEST_PEND_EXP_CFC_TST | NUMERIC | NO_UNIT | crate | bool | :1 | Expert HVCFC (100 pA test) pending | 1 | 1 | CRATE
blecsEHVRDAC |  | TEST_PEND_EXP_RST_DAC | NUMERIC | NO_UNIT | crate | bool | :1 | Expert HVRDAC (Dac Reset) pending | 1 | 1 | CRATE
blecsEHVRGOH |  | TEST_PEND_EXP_RST_GOH | NUMERIC | NO_UNIT | crate | bool | :1 | Expert HVRGOH (GOH Reset) pending | 1 | 1 | CRATE
blecsEMC |  | TEST_PEND_EXP_MAN_CTRL | NUMERIC | NO_UNIT | crate | bool | :1 | Expert Manual control Request pending | 1 | 1 | CRATE
blecsTSTRN |  | TEST_REQUEST_BY_TIMER_NORM | NUMERIC | NO_UNIT | crate | bool | :1 | Internal timer is requesting a test of the system with normal priority | 1 | 1 | CRATE
blecsTSTRP |  | TEST_REQUEST_BY_TIMER_PRIO | NUMERIC | NO_UNIT | crate | bool | :1 | Internal timer is requesting a test of the system with high priority (BeamPermit away next Dump) | 1 | 1 | CRATE
blecsAST |  | TEST_ACTIVE_SYS_TST | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested) | 1 | 1 | CRATE
blecsACT |  | TEST_ACTIVE_CONSIST | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : CONSISTENCY (threshold table control) | 1 | 1 | CRATE
blecsATBPL |  | TEST_ACTIVE_THRES_BP | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : THRESHOLD to BPL | 1 | 1 | CRATE
blecsAET |  | TEST_ACTIVE_ENGY | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : ENERGY TEST | 1 | 1 | CRATE
blecsABPBIS |  | TEST_ACTIVE_BP_BIS | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 1 | 1 | CRATE
blecsABPTC |  | TEST_ACTIVE_BP_TC | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 1 | 1 | CRATE
blecsAHVLF |  | TEST_ACTIVE_LF_MOD | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : HVLF (Modulation of the high voltage and analysis) | 1 | 1 | CRATE
blecsAHVCFC |  | TEST_ACTIVE_CFC_TST | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card) | 1 | 1 | CRATE
blecsAHVRDAC |  | TEST_ACTIVE_RST_DAC | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : HVRDAC (Activation DAC reset mode of the tunnel card) | 1 | 1 | CRATE
blecsAHVRGOH |  | TEST_ACTIVE_RST_GOH | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : HVRGOH (Activation GOH reset mode of the tunnel card) | 1 | 1 | CRATE
blecsMCAL |  | TEST_ALLOWED_MAN_CTRL | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : MANUAL CONTROL ALLOWED | 1 | 1 | CRATE
blecsMCAC |  | TEST_ACTIVE_MAN_CTRL | NUMERIC | NO_UNIT | crate | bool | :1 | Active test : MANUAL CONTROL ACTIVE | 1 | 1 | CRATE
blecsRSTP |  | TEST_PASSED_SYS_TST | NUMERIC | NO_UNIT | crate | bool | :1 | Test Passed : "SYSTEM TEST" | 1 | 1 | CRATE
blecsRCTP |  | TEST_PASSED_CONSIST | NUMERIC | NO_UNIT | crate | bool | :1 | Test Passed : CONSISTENCY (threshold table control) | 1 | 1 | CRATE
blecsRTBPLP |  | TEST_PASSED_THRES_BPL | NUMERIC | NO_UNIT | crate | bool | :1 | Test Passed : THRESHOLD to BPL | 1 | 1 | CRATE
blecsRETP |  | TEST_PASSED_ENGY | NUMERIC | NO_UNIT | crate | bool | :1 | Test Passed : ENERGY TEST | 1 | 1 | CRATE
blecsRBPBISP |  | TEST_PASSED_BP_BIS | NUMERIC | NO_UNIT | crate | bool | :1 | Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 1 | 1 | CRATE
blecsRBPTCP |  | TEST_PASSED_BP_TC | NUMERIC | NO_UNIT | crate | bool | :1 | Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 1 | 1 | CRATE
blecsRHVLFP |  | TEST_PASSED_LF_MOD | NUMERIC | NO_UNIT | crate | bool | :1 | Test Passed : HVLF (Modulation of the high voltage and analysis) | 1 | 1 | CRATE
blecsRHVTCFCP |  | TEST_PASSED_CFC_TEST | NUMERIC | NO_UNIT | crate | bool | :1 | Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card) | 1 | 1 | CRATE
blecsTRR |  | TEST_RESULT_RDY | NUMERIC | NO_UNIT | crate | bool | :1 | New test results are available (Check active test to find out which test) | 1 | 1 | CRATE
blecsUDL |  | BEAM_PERM_TC_LINE_U | NUMERIC | NO_UNIT | crate | bool | :1 | Threshold Comparators Unmaskable beam permit line arriving to the BLECS | 1 | 1 | CRATE
blecsMDL |  | BEAM_PERM_TC_LINE_M | NUMERIC | NO_UNIT | crate | bool | :1 | Threshold Comparators Maskable beam permit line arriving to the BLECS | 1 | 1 | CRATE
blecsUAI |  | BEAM_PERM_U_IN_A | NUMERIC | NO_UNIT | crate | bool | :1 | Beam Permit Unmaskable A Input from the upper Combiner | 1 | 1 | CRATE
blecsUBI |  | BEAM_PERM_U_IN_B | NUMERIC | NO_UNIT | crate | bool | :1 | Beam Permit Unmaskable B Input from the upper Combiner | 1 | 1 | CRATE
blecsMAI |  | BEAM_PERM_M_IN_A | NUMERIC | NO_UNIT | crate | bool | :1 | Beam Permit Maskable A Input from the upper Combiner | 1 | 1 | CRATE
blecsMBI |  | BEAM_PERM_M_IN_B | NUMERIC | NO_UNIT | crate | bool | :1 | Beam Permit Maskable B Input from the upper Combiner | 1 | 1 | CRATE
blecsUAO |  | BEAM_PERM_U_OUT_A | NUMERIC | NO_UNIT | crate | bool | :1 | Beam Permit Unmaskable A Output to lower Combiner (or CIBUS) | 1 | 1 | CRATE
blecsUBO |  | BEAM_PERM_U_OUT_B | NUMERIC | NO_UNIT | crate | bool | :1 | Beam Permit Unmaskable B Output to lower Combiner (or CIBUS) | 1 | 1 | CRATE
blecsMAO |  | BEAM_PERM_M_OUT_A | NUMERIC | NO_UNIT | crate | bool | :1 | Beam Permit Maskable A Output to lower Combiner (or CIBUS) | 1 | 1 | CRATE
blecsMBO |  | BEAM_PERM_M_OUT_B | NUMERIC | NO_UNIT | crate | bool | :1 | Beam Permit Maskable B Output to lower Combiner (or CIBUS) | 1 | 1 | CRATE
blecsUBS |  | BEAM_PERM_INFO_U | NUMERIC | NO_UNIT | crate | bool | :1 | Unmaskable Beam Info (Beam status) | 1 | 1 | CRATE
blecsMBS |  | BEAM_PERM_INFO_M | NUMERIC | NO_UNIT | crate | bool | :1 | Maskable Beam Info (Beam status) | 1 | 1 | CRATE
blecsUBSR |  | BEAM_PERM_CIBUS_U_FAULT | NUMERIC | NO_UNIT | crate | bool | :1 | Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP | 1 | 1 | CRATE
blecsMBSR |  | BEAM_PERM_CIBUS_M_FAULT | NUMERIC | NO_UNIT | crate | bool | :1 | Maskable Beam Info (beam status) had not respond to last BLECS-DUMP | 1 | 1 | CRATE
blecsTCLD |  | TURN_CNT_WN_DUMP_REQ | NUMERIC | Turn | crate | unsigned | :32 | Turn counter value when the last dump was requested (32 bits ) | 1 | 32 | CRATE
blecsBCLD |  | BUNCH_CNT_WN_DUMP_REQ | NUMERIC | Bunch | crate | unsigned | :12 | Bunch counter value when the last dump was requested (12 bits) | 1 | 12 | CRATE
blecsTDPM |  | BEAM_PERM_TIME_TO_PM_FRZE_M | NUMERIC | us | crate | unsigned | :16 | Time from the dump request maskable to PM_freeze (microsecond resolution) | 1 | 16 | CRATE
blecsTDPU |  | BEAM_PERM_TIME_TO_PM_FRZE_U | NUMERIC | us | crate | unsigned | :16 | Time from the dump request unmaskable to PM_freeze (microsecond resolution) | 1 | 16 | CRATE
blecsBEI |  | BEAM_ENGY_IN | NUMERIC | Bits | crate | unsigned | :16 | Beam Energy Input coming from timing distribution (CISV card AB/CO) | 1 | 16 | CRATE
blecsBEO |  | BEAM_ENGY_LEVEL_OUT | NUMERIC | Level | crate | unsigned | :5 | Beam Energy levels output to the Threshold Comparators (send to BLETC cards) | 1 | 5 | CRATE
blecsBEE |  | BEAM_ENGY_ERR | NUMERIC | NO_UNIT | crate | bool | :1 | Beam Energy error (send to BLETC cards) | 1 | 1 | CRATE
blecsSRTC |  | BEAM_ENGY_SOFT_RST_TC | NUMERIC | NO_UNIT | crate | bool | :1 | SoftResetTC (send to BLETC cards) | 1 | 1 | CRATE
blecsBECRCA |  | BEAM_ENGY_CRC_CNT_A | NUMERIC | ERROR | crate | unsigned | :16 | Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 1 | 16 | CRATE
blecsBECRCB |  | BEAM_ENGY_CRC_CNT_B | NUMERIC | ERROR | crate | unsigned | :16 | Number of CRC errors at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 1 | 16 | CRATE
blecsBELFA |  | BEAM_ENGY_LOSTF_CNT_A | NUMERIC | FRAME | crate | unsigned | :16 | Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 1 | 16 | CRATE
blecsBELFB |  | BEAM_ENGY_LOSTF_CNT_B | NUMERIC | FRAME | crate | unsigned | :16 | Number of lost frames (packets) at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 1 | 16 | CRATE
blecsBETOV |  | BEAM_ENGY_TIME_OUT | NUMERIC | TIMEOUT | crate | unsigned | :4 | Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO) | 1 | 16 | CRATE
blecsV1H |  | HV_1_VOLT_HIGH | NUMERIC | NO_UNIT | crate | bool | :1 | High Voltage power supply 1 / voltage higher | 1 | 1 | CRATE
blecsV1L |  | HV_1_VOLT_LOW | NUMERIC | NO_UNIT | crate | bool | :1 | High Voltage power supply 1 / voltage lower | 1 | 1 | CRATE
blecsI1H |  | HV_1_CURR_HIGH | NUMERIC | NO_UNIT | crate | bool | :1 | High Voltage power supply 1 / current higher | 1 | 1 | CRATE
blecsI1L |  | HV_1_CURR_LOW | NUMERIC | NO_UNIT | crate | bool | :1 | High Voltage power supply 1 / current lower | 1 | 1 | CRATE
blecsHV1V |  | HV_1_VOLT_VALUE | NUMERIC | mV | crate | unsigned | :24 | High Voltage power supply 1 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 1 | 24 | CRATE | need filter
blecsV2H |  | HV_2_VOLT_HIGH | NUMERIC | NO_UNIT | crate | bool | :1 | High Voltage power supply 2 / voltage higher | 1 | 1 | CRATE
blecsV2L |  | HV_2_VOLT_LOW | NUMERIC | NO_UNIT | crate | bool | :1 | High Voltage power supply 2 / voltage lower | 1 | 1 | CRATE
blecsI2H |  | HV_2_CURR_HIGH | NUMERIC | NO_UNIT | crate | bool | :1 | High Voltage power supply 2 / current higher | 1 | 1 | CRATE
blecsI2L |  | HV_2_CURR_LOW | NUMERIC | NO_UNIT | crate | bool | :1 | High Voltage power supply 2 / current lower | 1 | 1 | CRATE
blecsHV2V |  | HV_2_VOLT_VALUE | NUMERIC | mV | crate | unsigned | :24 | High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 1 | 24 | CRATE | need filter
blecsHV1I |  | HV_1_CURR_VALUE | NUMERIC | uA | crate | unsigned | :16 | High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA | 1 | 16 | CRATE | need filter
blecsHV2I |  | HV_2_CURR_VALUE | NUMERIC | uA | crate | unsigned | :16 | High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA | 1 | 16 | CRATE | need filter
blecsTCFUAB |  | TURN_CNT_WN_FALL_U_AORB | NUMERIC | Turn | crate | unsigned | :32 | Turn counter when last fall of UA or UB (32 bits) | 1 | 32 | CRATE
blecsTCFMAB |  | TURN_CNT_WN_FALL_M_AORB | NUMERIC | Turn | crate | unsigned | :32 | Turn counter when last fall of MA or MB (32 bits) | 1 | 32 | CRATE
blecsBCFMAB |  | BUNCH_CNT_WN_FALL_M_AORB | NUMERIC | Bunch | crate | unsigned | :12 | Bunch counter when last fall of MA or MB 12 bits counter | 1 | 16 | CRATE
blecsBCFUAB |  | BUNCH_CNT_WN_FALL_U_AORB | NUMERIC | Bunch | crate | unsigned | :12 | Bunch counter when last fall of UA or UB 12 bits counter | 1 | 16 | CRATE
blecsBLETCPRES |  | TEST_RES_BLETC_PRES | NUMERIC | NO_UNIT | crate | unsigned | :16 | Vector of BLETC Board present in the crate used by the tests functions (MSBit TC1 LSBit TC16) | 1 | 16 | CRATE
blecsBLECFPRES |  | TEST_RES_BLECF_PRES | NUMERIC | NO_UNIT | crate | unsigned | :32 | Vector of BLECF Board present in the crate used by the tests functions (MSBit BLECF1(TC1) LSBit BLECF32 (TC16)) | 1 | 32 | CRATE
blecsTRCBLETC |  | TEST_RES_CONSIST_PER_BLETC | NUMERIC | NO_UNIT | crate | unsigned | :16 | Result of the Consistency test per BLETC (MSBit TC1) 1=passed | 1 | 16 | CRATE
blecsTRBPUBLETC |  | TEST_RES_BP_TC_PER_BLETC_U | NUMERIC | NO_UNIT | crate | unsigned | :16 | Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed | 1 | 16 | CRATE
blecsTRBPMBLETC |  | TEST_RES_BP_TC_PER_BLETC_M | NUMERIC | NO_UNIT | crate | unsigned | :16 | Result of the BPLTC test maskable line (MSBit TC1) 1=passed | 1 | 16 | CRATE
blecsTRHVBLETC |  | TEST_RES_LF_MOD_PER_BLETC | NUMERIC | NO_UNIT | crate | unsigned | :16 | Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed | 1 | 16 | CRATE |  |  |  | BLETC number | Setup number
blecsHVCFCRES |  | TEST_RES_CFCTEST_PER_BLECF | NUMERIC | NO_UNIT | crate | unsigned | :32 | Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed | 1 | 32 | CRATE
blecsHVLFGAIN | blecsHVLFGAIN[1][1 to 16] | .BLETC.01:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  | Not automatically generated | BLETC | 1 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][1 to 16] | .BLETC.01:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  | concentrator should generate | BLETC | 1 | 2
blecsHVLFPH | blecsHVLFPH[1][1 to 16] | .BLETC.01:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  | from 2D array to 1D | BLETC | 1 | 1
blecsHVLFPH | blecsHVLFPH[2][1 to 16] | .BLETC.01:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  | then split it per TC card | BLETC | 1 | 2
blecsHVLFRES | blecsHVLFRES[1][1 to 16] | .BLETC.01:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 1 | 1
blecsHVLFRES | blecsHVLFRES[2][1 to 16] | .BLETC.01:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  | CFV-SR1-BLMR | BLETC | 1 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][17 to 32] | .BLETC.02:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  | RESULT_MOD | BLETC | 2 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][17 to 32] | .BLETC.02:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  | 39 | BLETC | 2 | 2
blecsHVLFPH | blecsHVLFPH[1][17 to 32] | .BLETC.02:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  | 40 | BLETC | 2 | 1
blecsHVLFPH | blecsHVLFPH[2][17 to 32] | .BLETC.02:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  | 34 | BLETC | 2 | 2
blecsHVLFRES | blecsHVLFRES[1][17 to 32] | .BLETC.02:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 2 | 1
blecsHVLFRES | blecsHVLFRES[2][17 to 32] | .BLETC.02:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 2 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][33 to 48] | .BLETC.03:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 3 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][33 to 48] | .BLETC.03:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 3 | 2
blecsHVLFPH | blecsHVLFPH[1][33 to 48] | .BLETC.03:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 3 | 1
blecsHVLFPH | blecsHVLFPH[2][33 to 48] | .BLETC.03:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 3 | 2
blecsHVLFRES | blecsHVLFRES[1][33 to 48] | .BLETC.03:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 3 | 1
blecsHVLFRES | blecsHVLFRES[2][33 to 48] | .BLETC.03:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 3 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][49 to 64] | .BLETC.04:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 4 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][49 to 64] | .BLETC.04:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 4 | 2
blecsHVLFPH | blecsHVLFPH[1][49 to 64] | .BLETC.04:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 4 | 1
blecsHVLFPH | blecsHVLFPH[2][49 to 64] | .BLETC.04:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 4 | 2
blecsHVLFRES | blecsHVLFRES[1][49 to 64] | .BLETC.04:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 4 | 1
blecsHVLFRES | blecsHVLFRES[2][49 to 64] | .BLETC.04:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 4 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][65 to 80] | .BLETC.05:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 5 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][65 to 80] | .BLETC.05:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 5 | 2
blecsHVLFPH | blecsHVLFPH[1][65 to 80] | .BLETC.05:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 5 | 1
blecsHVLFPH | blecsHVLFPH[2][65 to 80] | .BLETC.05:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 5 | 2
blecsHVLFRES | blecsHVLFRES[1][65 to 80] | .BLETC.05:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 5 | 1
blecsHVLFRES | blecsHVLFRES[2][65 to 80] | .BLETC.05:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 5 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][81 to 96] | .BLETC.06:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 6 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][81 to 96] | .BLETC.06:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 6 | 2
blecsHVLFPH | blecsHVLFPH[1][81 to 96] | .BLETC.06:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 6 | 1
blecsHVLFPH | blecsHVLFPH[2][81 to 96] | .BLETC.06:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 6 | 2
blecsHVLFRES | blecsHVLFRES[1][81 to 96] | .BLETC.06:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 6 | 1
blecsHVLFRES | blecsHVLFRES[2][81 to 96] | .BLETC.06:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 6 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][97 to 112] | .BLETC.07:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 7 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][97 to 112] | .BLETC.07:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 7 | 2
blecsHVLFPH | blecsHVLFPH[1][97 to 112] | .BLETC.07:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 7 | 1
blecsHVLFPH | blecsHVLFPH[2][97 to 112] | .BLETC.07:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 7 | 2
blecsHVLFRES | blecsHVLFRES[1][97 to 112] | .BLETC.07:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 7 | 1
blecsHVLFRES | blecsHVLFRES[2][97 to 112] | .BLETC.07:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 7 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][113 to 128] | .BLETC.08:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 8 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][113 to 128] | .BLETC.08:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 8 | 2
blecsHVLFPH | blecsHVLFPH[1][113 to 128] | .BLETC.08:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 8 | 1
blecsHVLFPH | blecsHVLFPH[2][113 to 128] | .BLETC.08:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 8 | 2
blecsHVLFRES | blecsHVLFRES[1][113 to 128] | .BLETC.08:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 8 | 1
blecsHVLFRES | blecsHVLFRES[2][113 to 128] | .BLETC.08:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 8 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][129 to 144] | .BLETC.09:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 9 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][129 to 144] | .BLETC.09:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 9 | 2
blecsHVLFPH | blecsHVLFPH[1][129 to 144] | .BLETC.09:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 9 | 1
blecsHVLFPH | blecsHVLFPH[2][129 to 144] | .BLETC.09:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 9 | 2
blecsHVLFRES | blecsHVLFRES[1][129 to 144] | .BLETC.09:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 9 | 1
blecsHVLFRES | blecsHVLFRES[2][129 to 144] | .BLETC.09:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 9 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][145 to 160] | .BLETC.010:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 10 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][145 to 160] | .BLETC.010:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 10 | 2
blecsHVLFPH | blecsHVLFPH[1][145 to 160] | .BLETC.010:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 10 | 1
blecsHVLFPH | blecsHVLFPH[2][145 to 160] | .BLETC.010:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 10 | 2
blecsHVLFRES | blecsHVLFRES[1][145 to 160] | .BLETC.010:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 10 | 1
blecsHVLFRES | blecsHVLFRES[2][145 to 160] | .BLETC.010:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 10 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][161 to 176] | .BLETC.011:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 11 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][161 to 176] | .BLETC.011:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 11 | 2
blecsHVLFPH | blecsHVLFPH[1][161 to 176] | .BLETC.011:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 11 | 1
blecsHVLFPH | blecsHVLFPH[2][161 to 176] | .BLETC.011:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 11 | 2
blecsHVLFRES | blecsHVLFRES[1][161 to 176] | .BLETC.011:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 11 | 1
blecsHVLFRES | blecsHVLFRES[2][161 to 176] | .BLETC.011:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 11 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][177 to 192] | .BLETC.012:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 12 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][177 to 192] | .BLETC.012:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 12 | 2
blecsHVLFPH | blecsHVLFPH[1][177 to 192] | .BLETC.012:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 12 | 1
blecsHVLFPH | blecsHVLFPH[2][177 to 192] | .BLETC.012:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 12 | 2
blecsHVLFRES | blecsHVLFRES[1][177 to 192] | .BLETC.012:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 12 | 1
blecsHVLFRES | blecsHVLFRES[2][177 to 192] | .BLETC.012:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 12 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][193 to 208] | .BLETC.013:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 13 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][193 to 208] | .BLETC.013:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 13 | 2
blecsHVLFPH | blecsHVLFPH[1][193 to 208] | .BLETC.013:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 13 | 1
blecsHVLFPH | blecsHVLFPH[2][193 to 208] | .BLETC.013:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 13 | 2
blecsHVLFRES | blecsHVLFRES[1][193 to 208] | .BLETC.013:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 13 | 1
blecsHVLFRES | blecsHVLFRES[2][193 to 208] | .BLETC.013:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 13 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][209 to 224] | .BLETC.014:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 14 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][209 to 224] | .BLETC.014:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 14 | 2
blecsHVLFPH | blecsHVLFPH[1][209 to 224] | .BLETC.014:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 14 | 1
blecsHVLFPH | blecsHVLFPH[2][209 to 224] | .BLETC.014:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 14 | 2
blecsHVLFRES | blecsHVLFRES[1][209 to 224] | .BLETC.014:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 14 | 1
blecsHVLFRES | blecsHVLFRES[2][209 to 224] | .BLETC.014:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 14 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][225 to 240] | .BLETC.015:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 15 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][225 to 240] | .BLETC.015:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 15 | 2
blecsHVLFPH | blecsHVLFPH[1][225 to 240] | .BLETC.015:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 15 | 1
blecsHVLFPH | blecsHVLFPH[2][225 to 240] | .BLETC.015:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 15 | 2
blecsHVLFRES | blecsHVLFRES[1][225 to 240] | .BLETC.015:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 15 | 1
blecsHVLFRES | blecsHVLFRES[2][225 to 240] | .BLETC.015:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 15 | 2
blecsHVLFGAIN | blecsHVLFGAIN[1][241 to 256] | .BLETC.016:RESULT_MOD_1_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 16 | 1
blecsHVLFGAIN | blecsHVLFGAIN[2][241 to 256] | .BLETC.016:RESULT_MOD_2_GAIN | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 16 | 2
blecsHVLFPH | blecsHVLFPH[1][241 to 256] | .BLETC.016:RESULT_MOD_1_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 1), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 16 | 1
blecsHVLFPH | blecsHVLFPH[2][241 to 256] | .BLETC.016:RESULT_MOD_2_PHASE | VECTORNUMERIC | NO_UNIT | crate | unsigned | :15 | Result of the HVLF modulation test (setup 2), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 1 | 240 | CRATE |  |  | BLETC | 16 | 2
blecsHVLFRES | blecsHVLFRES[1][241 to 256] | .BLETC.016:RESULT_MOD_1 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 1), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 16 | 1
blecsHVLFRES | blecsHVLFRES[2][241 to 256] | .BLETC.016:RESULT_MOD_2 | VECTORNUMERIC | NO_UNIT | crate | bool | :1 | Result of the HVLF modulation test (setup 2), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 1 | 16 | CRATE |  |  | BLETC | 16 | 2
[9 empty] | STATUS | total in bits | concatenation
[8 empty] | Variables to the logging database (bool and unsigned) | 194 | 16441

## SR1
Database name | Logging datatype | units | description | Time Precison (fractional seconds e.g. 0=seconds 3=milliseconds 6=microseconds) | Variable Type | Expected Data Rate | Logging Required | Delta Value | Delta Type (A = Absolute P = Percent) | Dead Time Interval Size | Dead Time Interval Type | Fixed Logging Interval Size | Fixed Logging Interval Type | Rounding (Decimal Places Required) | Path to Hierarchy Node (nodes seperated by '->') | Entity type | crate | IP | #VALUE!
#VALUE! | NUMERIC | NO_UNIT | System is operational (View by the BLECS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate | #VALUE!
#VALUE! | NUMERIC | NO_UNIT | System is under test (All BLM crate of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test requested (while take place when the conditions are fullfill) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 2 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 or 2 fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault with (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy fault (From CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Turn Clock fault (From BOBR card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Dump occurred | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Threshold to BPL pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Energy Test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPTC (Beam Permit BLETC to BLECS) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVLF test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVCFC (100 pA test) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRDAC (Dac Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRGOH (GOH Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Manual control Request pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with normal priority | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with high priority (BeamPermit away next Dump) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRDAC (Activation DAC reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRGOH (Activation GOH reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ALLOWED | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ACTIVE | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : "SYSTEM TEST" | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | New test results are available (Check active test to find out which test) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Unmaskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Maskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter value when the last dump was requested (32 bits ) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter value when the last dump was requested (12 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request maskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request unmaskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bits | Beam Energy Input coming from timing distribution (CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Level | Beam Energy levels output to the Threshold Comparators (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy error (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | SoftResetTC (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | TIMEOUT | Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 1 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of UA or UB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of MA or MB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of MA or MB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of UA or UB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLETC Board present in the crate used by the tests functions (MSBit TC1 LSBit TC16) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLECF Board present in the crate used by the tests functions (MSBit BLECF1(TC1) LSBit BLECF32 (TC16)) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the Consistency test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test maskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | System is operational (View by the BLECS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate | #VALUE!
#VALUE! | NUMERIC | NO_UNIT | System is under test (All BLM crate of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test requested (while take place when the conditions are fullfill) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 2 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 or 2 fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault with (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy fault (From CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Turn Clock fault (From BOBR card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Dump occurred | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Threshold to BPL pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Energy Test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPTC (Beam Permit BLETC to BLECS) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVLF test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVCFC (100 pA test) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRDAC (Dac Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRGOH (GOH Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Manual control Request pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with normal priority | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with high priority (BeamPermit away next Dump) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRDAC (Activation DAC reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRGOH (Activation GOH reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ALLOWED | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ACTIVE | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : "SYSTEM TEST" | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | New test results are available (Check active test to find out which test) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Unmaskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Maskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter value when the last dump was requested (32 bits ) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter value when the last dump was requested (12 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request maskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request unmaskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bits | Beam Energy Input coming from timing distribution (CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Level | Beam Energy levels output to the Threshold Comparators (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy error (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | SoftResetTC (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | TIMEOUT | Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 1 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of UA or UB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of MA or MB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of MA or MB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of UA or UB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLETC Board present in the crate used by the tests functions (MSBit TC1 LSBit TC16) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLECF Board present in the crate used by the tests functions (MSBit BLECF1(TC1) LSBit BLECF32 (TC16)) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the Consistency test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test maskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | System is operational (View by the BLECS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate | #VALUE!
#VALUE! | NUMERIC | NO_UNIT | System is under test (All BLM crate of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test requested (while take place when the conditions are fullfill) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 2 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 or 2 fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault with (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy fault (From CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Turn Clock fault (From BOBR card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Dump occurred | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Threshold to BPL pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Energy Test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPTC (Beam Permit BLETC to BLECS) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVLF test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVCFC (100 pA test) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRDAC (Dac Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRGOH (GOH Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Manual control Request pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with normal priority | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with high priority (BeamPermit away next Dump) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRDAC (Activation DAC reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRGOH (Activation GOH reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ALLOWED | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ACTIVE | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : "SYSTEM TEST" | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | New test results are available (Check active test to find out which test) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Unmaskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Maskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter value when the last dump was requested (32 bits ) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter value when the last dump was requested (12 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request maskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request unmaskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bits | Beam Energy Input coming from timing distribution (CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Level | Beam Energy levels output to the Threshold Comparators (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy error (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | SoftResetTC (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | TIMEOUT | Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 1 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of UA or UB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of MA or MB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of MA or MB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of UA or UB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLETC Board present in the crate used by the tests functions (MSBit TC1 LSBit TC16) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLECF Board present in the crate used by the tests functions (MSBit BLECF1(TC1) LSBit BLECF32 (TC16)) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the Consistency test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test maskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate

## SR7
Database name | Logging datatype | units | description | Time Precison (fractional seconds e.g. 0=seconds 3=milliseconds 6=microseconds) | Variable Type | Expected Data Rate | Logging Required | Delta Value | Delta Type (A = Absolute P = Percent) | Dead Time Interval Size | Dead Time Interval Type | Fixed Logging Interval Size | Fixed Logging Interval Type | Rounding (Decimal Places Required) | Path to Hierarchy Node (nodes seperated by '->') | Entity type | crate | IP | #VALUE!
#VALUE! | NUMERIC | NO_UNIT | System is operational (View by the BLECS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate | #VALUE!
#VALUE! | NUMERIC | NO_UNIT | System is under test (All BLM crate of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test requested (while take place when the conditions are fullfill) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 2 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 or 2 fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault with (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy fault (From CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Turn Clock fault (From BOBR card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Dump occurred | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Threshold to BPL pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Energy Test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPTC (Beam Permit BLETC to BLECS) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVLF test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVCFC (100 pA test) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRDAC (Dac Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRGOH (GOH Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Manual control Request pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with normal priority | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with high priority (BeamPermit away next Dump) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRDAC (Activation DAC reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRGOH (Activation GOH reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ALLOWED | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ACTIVE | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : "SYSTEM TEST" | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | New test results are available (Check active test to find out which test) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Unmaskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Maskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter value when the last dump was requested (32 bits ) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter value when the last dump was requested (12 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request maskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request unmaskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bits | Beam Energy Input coming from timing distribution (CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Level | Beam Energy levels output to the Threshold Comparators (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy error (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | SoftResetTC (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | TIMEOUT | Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 1 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of UA or UB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of MA or MB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of MA or MB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of UA or UB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLETC Board present in the crate used by the tests functions (MSBit TC1 LSBit TC16) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLECF Board present in the crate used by the tests functions (MSBit BLECF1(TC1) LSBit BLECF32 (TC16)) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the Consistency test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test maskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | System is operational (View by the BLECS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate | #VALUE!
#VALUE! | NUMERIC | NO_UNIT | System is under test (All BLM crate of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test requested (while take place when the conditions are fullfill) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 2 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 or 2 fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault with (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy fault (From CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Turn Clock fault (From BOBR card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Dump occurred | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Threshold to BPL pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Energy Test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPTC (Beam Permit BLETC to BLECS) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVLF test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVCFC (100 pA test) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRDAC (Dac Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRGOH (GOH Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Manual control Request pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with normal priority | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with high priority (BeamPermit away next Dump) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRDAC (Activation DAC reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRGOH (Activation GOH reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ALLOWED | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ACTIVE | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : "SYSTEM TEST" | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | New test results are available (Check active test to find out which test) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Unmaskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Maskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter value when the last dump was requested (32 bits ) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter value when the last dump was requested (12 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request maskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request unmaskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bits | Beam Energy Input coming from timing distribution (CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Level | Beam Energy levels output to the Threshold Comparators (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy error (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | SoftResetTC (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | TIMEOUT | Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 1 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of UA or UB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of MA or MB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of MA or MB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of UA or UB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLETC Board present in the crate used by the tests functions (MSBit TC1 LSBit TC16) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLECF Board present in the crate used by the tests functions (MSBit BLECF1(TC1) LSBit BLECF32 (TC16)) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the Consistency test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test maskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | System is operational (View by the BLECS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate | #VALUE!
#VALUE! | NUMERIC | NO_UNIT | System is under test (All BLM crate of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test requested (while take place when the conditions are fullfill) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 2 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 or 2 fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault with (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy fault (From CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Turn Clock fault (From BOBR card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Dump occurred | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Threshold to BPL pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Energy Test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPTC (Beam Permit BLETC to BLECS) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVLF test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVCFC (100 pA test) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRDAC (Dac Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRGOH (GOH Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Manual control Request pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with normal priority | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with high priority (BeamPermit away next Dump) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRDAC (Activation DAC reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRGOH (Activation GOH reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ALLOWED | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ACTIVE | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : "SYSTEM TEST" | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | New test results are available (Check active test to find out which test) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Unmaskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Maskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter value when the last dump was requested (32 bits ) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter value when the last dump was requested (12 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request maskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request unmaskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bits | Beam Energy Input coming from timing distribution (CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Level | Beam Energy levels output to the Threshold Comparators (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy error (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | SoftResetTC (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | TIMEOUT | Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 1 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of UA or UB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of MA or MB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of MA or MB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of UA or UB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLETC Board present in the crate used by the tests functions (MSBit TC1 LSBit TC16) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLECF Board present in the crate used by the tests functions (MSBit BLECF1(TC1) LSBit BLECF32 (TC16)) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the Consistency test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test maskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | System is operational (View by the BLECS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate | #VALUE!
#VALUE! | NUMERIC | NO_UNIT | System is under test (All BLM crate of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test requested (while take place when the conditions are fullfill) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 2 fault (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High voltage 1 or 2 fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault with (survey by comparators) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Low voltage fault (survey by ADC) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy fault (From CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Turn Clock fault (From BOBR card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Dump occurred | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | User BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert System test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Consistency test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Threshold to BPL pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Energy Test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPBIS (BLECS to Beam interlock) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert BPTC (Beam Permit BLETC to BLECS) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVLF test pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVCFC (100 pA test) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRDAC (Dac Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert HVRGOH (GOH Reset) pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Expert Manual control Request pending | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with normal priority | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Internal timer is requesting a test of the system with high priority (BeamPermit away next Dump) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : "SYSTEM TEST" (All BLM crates of the IP are involved but only this crate is tested) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVTCFC (Activation Test CFC "100pA" mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRDAC (Activation DAC reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : HVRGOH (Activation GOH reset mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ALLOWED | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Active test : MANUAL CONTROL ACTIVE | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : "SYSTEM TEST" | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : CONSISTENCY (threshold table control) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : THRESHOLD to BPL | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : ENERGY TEST | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPBIS BPBIS (Beam Permit lines fron the last combiner to the LBIS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : BPTC (Threshold comparator Beam Permit to last Combiner before CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVLF (Modulation of the high voltage and analysis) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Test Passed : HVTCFC (Activation Test CFC mode of the tunnel card) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | New test results are available (Check active test to find out which test) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Unmaskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Threshold Comparators Maskable beam permit line arriving to the BLECS | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Input from the upper Combiner | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Unmaskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable A Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Permit Maskable B Output to lower Combiner (or CIBUS) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (Beam status) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Unmaskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Maskable Beam Info (beam status) had not respond to last BLECS-DUMP | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter value when the last dump was requested (32 bits ) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter value when the last dump was requested (12 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request maskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | us | Time from the dump request unmaskable to PM_freeze (microsecond resolution) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bits | Beam Energy Input coming from timing distribution (CISV card AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Level | Beam Energy levels output to the Threshold Comparators (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Beam Energy error (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | SoftResetTC (send to BLETC cards) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | ERROR | Number of CRC errors at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel A) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | FRAME | Number of lost frames (packets) at the reception of the beam energy (channel B) from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | TIMEOUT | Number of TIMEOUTS at the reception of the beam energy value from the CISV card (AB/CO) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 1 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 1 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / voltage lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current higher | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | High Voltage power supply 2 / current lower | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | mV | High Voltage power supply 2 / voltage value (24 bits) [HV = 3000/2**24*code] in V | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 1 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | uA | High Voltage power supply 2 / current value (16 bits) [I = 20/2**16*code] in mA | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of UA or UB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Turn | Turn counter when last fall of MA or MB (32 bits) | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of MA or MB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | Bunch | Bunch counter when last fall of UA or UB 12 bits counter | 3 | NORMAL | ~1 per minute | Y |  | A |  |  | 1 | MINUTE | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLETC Board present in the crate used by the tests functions (MSBit TC1 LSBit TC16) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Vector of BLECF Board present in the crate used by the tests functions (MSBit BLECF1(TC1) LSBit BLECF32 (TC16)) | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the Consistency test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test unmaskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the BPLTC test maskable line (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVLF modulation test per BLETC (MSBit TC1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | NUMERIC | NO_UNIT | Result of the HVCFC test (100pA) per BLECF (MSBit BLECF1) 1=passed | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.01, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.02, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.03, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.04, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.05, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.06, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.07, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.08, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.09, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.010, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.011, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.012, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.013, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.014, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.015, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Gain per channel (MaxRS/Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Phase per channel (between MaxRS and Modulation Voltage), one numeric value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 1), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate
#VALUE! | VECTORNUMERIC | NO_UNIT | Result of the HVLF modulation test (setup 2), BLETC.016, Result per channel (Passed/Failed), one boolean value for each of the 16 channels | 3 | NORMAL | ~1 per hour | Y |  | A |  |  | 1 | HOUR | 10 | #VALUE! | crate

## delated variables
blecsUTBPL | TEST_PEND_USR_THRES_BPL | NUMERIC | NO_UNIT | crate | bool | :1 | User Threshold to BPL pending | 1 | 1 | CRATE
blecsUET | TEST_PEND_USR_ENGY | NUMERIC | NO_UNIT | crate | bool | :1 | User Energy Test pending | 1 | 1 | CRATE