## Installation overview
BLECF installation (tunnel electronic)
| ARC |  |  |  |  | straight section
| BLECF |  |  |  |  | BLECF |  |  |  |  | Source test planning
| Dismounted | Calibrated | Installed | Checked and working | Comments | Dismounted | Calibrated | Installed | Checked and working | Comments
1 LEFT | 22 | 22 | 22 | 22 |  | 19 | 19 | 19 | 19 |  | 4
1 RIGHT | 23 | 23 | 23 | 23 |  | 19 | 19 | 19 | 19 |  | 3
2 LEFT | 22 | 22 | 22 | 22 |  | 22 | 22 | 22 | 22 |  | 3
2 RIGHT | 23 | 23 | 23 | 23 |  | 21 | 21 | 21 | 21 |  | 8
3 LEFT | 22 | 22 | 22 | 22 |  | 18 | 18 | 18 | 18 |  | 8
3 RIGHT | 23 | 23 | 23 | 23 |  | 18 | 18 | 18 | 18 |  | 2
4 LEFT | 22 | 22 | 22 | 22 |  | 8 | 8 | 8 | 8 |  | 2
4 RIGHT | 23 | 23 | 23 | 23 |  | 8 | 8 | 8 | 8 |  | 6
5 LEFT | 22 | 22 | 22 | 22 |  | 19 | 19 | 19 | 19 |  | 6
5 RIGHT | 23 | 23 | 23 | 23 |  | 19 | 19 | 19 | 19 |  | 7
6 LEFT | 22 | 22 | 22 | 22 |  | 21 | 21 | 21 | 21 |  | 7
6 RIGHT | 23 | 23 | 23 | 23 |  | 21 | 21 | 21 | 21 |  | 1
7 LEFT | 22 | 22 | 22 | 22 |  | 28 | 28 | 28 | 28 |  | 1
7 RIGHT | 23 | 23 | 23 | 23 |  | 28 | 28 | 28 | 28 |  | 5
8 LEFT | 22 | 22 | 22 | 22 |  | 19 | 19 | 19 | 19 |  | 5
8 RIGHT | 23 | 23 | 23 | 23 |  | 21 | 21 | 20 | 21 | C11.R8 (BLECF_0300) not installed | 4
Total | 360 | 360 | 360 | 360 | 0 | 309 | 309 | 308 | 309
in % |  | 100 | 100 | 100 | 0 |  | 100 | 99.6764 | 100
Missing |  | 0 | 0 | 0 | 360 |  | 0 | 1 | 0
Sum BLECF complete installation |  |  |  |  | 669
Installed BLECF |  |  |  |  | 669 | 100
Missing BLECF complete installation |  |  |  |  | 0 | 0

## Point 1
| Tunnel - position | [14 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L1 |  | BLECF_ | 285 | 0x | 011D | 6 | Arc | BYPLM.A33L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
2 | BJBHT.A32L1 |  | BLECF_ | 280 | 0x | 0118 | 6 | Arc | BYPLM.A32L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 00 | [9 empty] | No
3 | BJBHT.A31L1 |  | BLECF_ | 276 | 0x | 0114 | 6 | Arc | BYPLM.A31L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
4 | BJBHT.A30L1 |  | BLECF_ | 201 | 0x | 00C9 | 6 | Arc | BYPLM.A30L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
5 | BJBHT.A29L1 |  | BLECF_ | 141 | 0x | 008D | 6 | Arc | BYPLM.A29L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
6 | BJBHT.A28L1 |  | BLECF_ | 249 | 0x | 00F9 | 6 | Arc | BYPLM.A28L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | DAC 0 --> CFC reset and call for verification
7 | BJBHT.A27L1 |  | BLECF_ | 241 | 0x | 00F1 | 6 | Arc | BYPLM.A27L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
8 | BJBHT.A26L1 |  | BLECF_ | 282 | 0x | 011A | 6 | Arc | BYPLM.A26L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Ch6 blocked and all DAC 0 --> CFC reset and call for verification
9 | BJBHT.A25L1 |  | BLECF_ | 218 | 0x | 00DA | 6 | Arc | BYPLM.A25L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | DAC high --> Reset needed
10 | BJBHT.A24L1 |  | BLECF_ | 204 | 0x | 00CC | 6 | Arc | BYPLM.A24L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Reset needed
11 | BJBHT.A23L1 |  | BLECF_ | 212 | 0x | 00D4 | 6 | Arc | BYPLM.A23L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
12 | BJBHT.A22L1 |  | BLECF_ | 284 | 0x | 011C | 6 | Arc | BYPLM.A22L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
13 | BJBHT.A21L1 |  | BLECF_ | 22 | 0x | 0016 | 6 | Arc | BYPLM.A21L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
14 | BJBHT.A20L1 |  | BLECF_ | 754 | 0x | 02F2 | 6 | Arc | BYPLM.A20L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
15 | BJBHT.A19L1 |  | BLECF_ | 35 | 0x | 0023 | 6 | Arc | BYPLM.A19L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
16 | BJBHT.A18L1 |  | BLECF_ | 400 | 0x | 0190 | 6 | Arc | BYPLM.A18L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
17 | BJBHT.A17L1 |  | BLECF_ | 36 | 0x | 0024 | 6 | Arc | BYPLM.A17L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
18 | BJBHT.A16L1 |  | BLECF_ | 64 | 0x | 0040 | 6 | Arc | BYPLM.A16L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Reset needed
19 | BJBHT.A15L1 |  | BLECF_ | 59 | 0x | 003B | 6 | Arc | BYPLM.A15L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
20 | BJBHT.A14L1 |  | BLECF_ | 58 | 0x | 003A | 6 | Arc | BYPLM.A14L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
21 | BJBHT.A13L1 |  | BLECF_ | 77 | 0x | 004D | 8 | Arc | BYPLM.A13L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
22 | BJBHT.A12L1 |  | BLECF_ | 79 | 0x | 004F | 6 | Arc | BYPLM.A12L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
| BJBHT.B11L1 |  | BLECF_ | 667 | 0x | 029B | 8 | RR13 | BY02 | 1 | CL | 2 | 9 | B | 13-14 | SR1 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
24 | BJBHT.A11L1 |  | BLECF_ | 594 | 0x | 0252 | 8 | RR13 | BY02 | 1 | CL | 2 | 8 | B | 11-12 | SR1 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
25 | BJBHT.B10L1 |  | BLECF_ | 409 | 0x | 0199 | 6 | RR13 | BY02 | 1 | CL | 2 | 7 | B | 9-10 | SR1 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
26 | BJBHT.A10L1 |  | BLECF_ | 550 | 0x | 0226 | 2 | RR13 | BY02 | 1 | CL | 2 | 6 | B | 7-8 | SR1 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
27 | BJBHT.A9L1 |  | BLECF_ | 518 | 0x | 0206 | 6 | RR13 | BY02 | 1 | CL | 2 | 5 | B | 5-6 | SR1 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Reset needed
28 | BJBHT.B8L1 |  | BLECF_ | 457 | 0x | 01C9 | 6 | RR13 | BY02 | 1 | CL | 2 | 4 | B | 3-4 | SR1 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
29 | BJBHT.A8L1 |  | BLECF_ | 588 | 0x | 024C | 6 | RR13 | BY02 | 1 | CL | 2 | 3 | B | 1-2 | SR1 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
30 | BJBHT.B7L1 |  | BLECF_ | 322 | 0x | 0142 | 7 | RR13 | BY02 | 1 | CL | 2 | 2 | A | 23-24 | SR1 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
31 | BJBHT.A7L1 |  | BLECF_ | 194 | 0x | 00C2 | 1 | RR13 | BY02 | 1 | CL | 2 | 1 | A | 21-22 | SR1 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
32 | BJBHT.A6L1 |  | BLECF_ | 88 | 0x | 0058 | 6 | RR13 | BY02 | 1 | CL | 1 | 10 | A | 19-20 | SR1 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
33 | BJBHT.B5L1 |  | BLECF_ | 530 | 0x | 0212 | 7 | RR13 | BY02 | 1 | CL | 1 | 9 | A | 17-18 | SR1 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
34 | BJBHT.A5L1 |  | BLECF_ | 597 | 0x | 0255 | 1 | RR13 | BY02 | 1 | CL | 1 | 8 | A | 15-16 | SR1 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
35 | BJBHT.C4L1 |  | BLECF_ | 398 | 0x | 018E | 6 | RR13 | BY02 | 1 | CL | 1 | 7 | A | 13-14 | SR1 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
36 | BJBHT.B4L1 |  | BLECF_ | 374 | 0x | 0176 | 4 | RR13 | BY02 | 1 | CL | 1 | 6 | A | 11-12 | SR1 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
37 | BJBHT.A4L1 |  | BLECF_ | 401 | 0x | 0191 | 4 | RR13 | BY02 | 1 | CL | 1 | 5 | A | 9-10 | SR1 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
38 | BJBHT.A3L1 |  | BLECF_ | 593 | 0x | 0251 | 5 | RR13 | BY02 | 1 | CL | 1 | 4 | A | 7-8 | SR1 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | HV dividere not installed -----> HV divider installed 17.2.15 | Access Tuesday 16.2
39 | BJBHT.A2L1 |  | BLECF_ | 592 | 0x | 0250 | 8 | RR13 | BY02 | 1 | CL | 1 | 3 | A | 5-6 | SR1 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | 1102566
40 | BJBHT.B1L1 |  | BLECF_ | 635 | 0x | 027B | 5 | RR13 | BY02 | 1 | CL | 1 | 2 | A | 3-4 | SR1 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | 1102564 BJBAP.B1L1=UX15
41 | BJBHT.A1L1 | yes | BLECF_ | 251 | 0x | 00FB | 1 | RR13 | BY02 | 1 | CL | 1 | 1 | A | 1-2 | SR1 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | HV missing (problem with CFC!) BLECF exchanged
42 | BJBHT.A1R1 | yes | BLECF_ | 377 | 0x | 0179 | 1 | RR17 | BY02 | 1 | CR | 1 | 1 | A | 1-2 | SR1 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Pins of Ch2-8 disconnected in filter box and Signalbox!!!! | Access Tuesday 16.2
43 | BJBHT.B1R1 |  | BLECF_ | 331 | 0x | 014B | 5 | RR17 | BY02 | 1 | CR | 1 | 2 | A | 3-4 | SR1 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
44 | BJBHT.A2R1 |  | BLECF_ | 501 | 0x | 01F5 | 8 | RR17 | BY02 | 1 | CR | 1 | 3 | A | 5-6 | SR1 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
45 | BJBHT.A3R1 |  | BLECF_ | 636 | 0x | 027C | 5 | RR17 | BY02 | 1 | CR | 1 | 4 | A | 7-8 | SR1 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | HV dividere not installed -----> HV divider installed 17.2.15 | Access Tuesday 16.2
46 | BJBHT.A4R1 |  | BLECF_ | 676 | 0x | 02A4 | 4 | RR17 | BY02 | 1 | CR | 1 | 5 | A | 9-10 | SR1 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
47 | BJBHT.B4R1 |  | BLECF_ | 653 | 0x | 028D | 4 | RR17 | BY02 | 1 | CR | 1 | 6 | A | 11-12 | SR1 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
48 | BJBHT.C4R1 |  | BLECF_ | 376 | 0x | 0178 | 6 | RR17 | BY02 | 1 | CR | 1 | 7 | A | 13-14 | SR1 | BY05 | C | 11 | 1-2 | BY02 | R | 3 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
49 | BJBHT.A5R1 |  | BLECF_ | 437 | 0x | 01B5 | 1 | RR17 | BY02 | 1 | CR | 1 | 8 | A | 15-16 | SR1 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | HV missing --> cable was badly connected
50 | BJBHT.B5R1 |  | BLECF_ | 442 | 0x | 01BA | 7 | RR17 | BY02 | 1 | CR | 1 | 9 | A | 17-18 | SR1 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | check if cable is NES connected
51 | BJBHT.A6R1 |  | BLECF_ | 355 | 0x | 0163 | 6 | RR17 | BY02 | 1 | CR | 1 | 10 | A | 19-20 | SR1 | BY05 | C | 11 | 7-8 | BY02 | R | 3 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | check if cable is NES connected & Reset
52 | BJBHT.A7R1 |  | BLECF_ | 323 | 0x | 0143 | 1 | RR17 | BY02 | 1 | CR | 2 | 1 | A | 21-22 | SR1 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
53 | BJBHT.B7R1 |  | BLECF_ | 512 | 0x | 0200 | 7 | RR17 | BY02 | 1 | CR | 2 | 2 | A | 23-24 | SR1 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
54 | BJBHT.A8R1 |  | BLECF_ | 325 | 0x | 0145 | 6 | RR17 | BY02 | 1 | CR | 2 | 3 | B | 1-2 | SR1 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
55 | BJBHT.B8R1 |  | BLECF_ | 647 | 0x | 0287 | 6 | RR17 | BY02 | 1 | CR | 2 | 4 | B | 3-4 | SR1 | BY05 | C | 10 | 3-4 | BY02 | R | 3 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
56 | BJBHT.A9R1 |  | BLECF_ | 85 | 0x | 0055 | 6 | RR17 | BY02 | 1 | CR | 2 | 5 | B | 5-6 | SR1 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | check if cable is NES connected and installation & Reset
57 | BJBHT.A10R1 |  | BLECF_ | 338 | 0x | 0152 | 2 | RR17 | BY02 | 1 | CR | 2 | 6 | B | 7-8 | SR1 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
58 | BJBHT.B10R1 |  | BLECF_ | 567 | 0x | 0237 | 6 | RR17 | BY02 | 1 | CR | 2 | 7 | B | 9-10 | SR1 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Levels are slowly blinking , HV missing and a lot of noise on all channels; Channel 2,4,5,8 is blocked to 0
59 | BJBHT.A11R1 |  | BLECF_ | 349 | 0x | 015D | 8 | RR17 | BY02 | 1 | CR | 2 | 8 | B | 11-12 | SR1 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | reset needed
60 | BJBHT.B11R1 |  | BLECF_ | 327 | 0x | 0147 | 8 | RR17 | BY03 | 1 | CR | 2 | 9 | B | 13-14 | SR2 | BY06 | C | 9 | 1-2 | BY02 | R | 3 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
61 | BJBHT.A12R1 |  | BLECF_ | 539 | 0x | 021B | 6 | Arc | BYPLM.A12R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | O.K | O.K | OK | OK | O.K | O.K | O.K | [10 empty] | No
62 | BJBHT.A13R1 |  | BLECF_ | 361 | 0x | 0169 | 8 | Arc | BYPLM.A13R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | O.K | O.K | OK | OK | O.K | O.K | O.K | [10 empty] | No
63 | BJBHT.A14R1 |  | BLECF_ | 570 | 0x | 023A | 6 | Arc | BYPLM.A14R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | O.K | O.K | OK | OK | O.K | O.K | O.K | [10 empty] | No
64 | BJBHT.A15R1 |  | BLECF_ | 306 | 0x | 0132 | 6 | Arc | BYPLM.A15R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
65 | BJBHT.A16R1 |  | BLECF_ | 311 | 0x | 0137 | 6 | Arc | BYPLM.A16R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | No Power on the Main
66 | BJBHT.A17R1 |  | BLECF_ | 303 | 0x | 012F | 6 | Arc | BYPLM.A17R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Reset !
67 | BJBHT.A18R1 |  | BLECF_ | 515 | 0x | 0203 | 6 | Arc | BYPLM.A18R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
68 | BJBHT.A19R1 |  | BLECF_ | 481 | 0x | 01E1 | 6 | Arc | BYPLM.A19R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
69 | BJBHT.A20R1 |  | BLECF_ | 81 | 0x | 0051 | 6 | Arc | BYPLM.A20R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Reset !
70 | BJBHT.A21R1 |  | BLECF_ | 522 | 0x | 020A | 6 | Arc | BYPLM.A21R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
71 | BJBHT.A22R1 |  | BLECF_ | 373 | 0x | 0175 | 6 | Arc | BYPLM.A22R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
72 | BJBHT.A23R1 |  | BLECF_ | 423 | 0x | 01A7 | 6 | Arc | BYPLM.A23R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
73 | BJBHT.A24R1 |  | BLECF_ | 352 | 0x | 0160 | 6 | Arc | BYPLM.A24R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
74 | BJBHT.A25R1 |  | BLECF_ | 500 | 0x | 01F4 | 6 | Arc | BYPLM.A25R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
75 | BJBHT.A26R1 |  | BLECF_ | 479 | 0x | 01DF | 6 | Arc | BYPLM.A26R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
76 | BJBHT.A27R1 |  | BLECF_ | 478 | 0x | 01DE | 6 | Arc | BYPLM.A27R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | O.K | O.K | OK | OK | O.K | O.K | O.K | [10 empty] | No |  | all connected channels go to 0! HV needed to be checked
77 | BJBHT.A28R1 |  | BLECF_ | 317 | 0x | 013D | 6 | Arc | BYPLM.A28R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
78 | BJBHT.A29R1 |  | BLECF_ | 310 | 0x | 0136 | 6 | Arc | BYPLM.A29R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Cable disconnected from crate to workaround Problem with the WorldFip reset
79 | BJBHT.A30R1 |  | BLECF_ | 329 | 0x | 0149 | 6 | Arc | BYPLM.A30R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | O.K | O.K | ↑↑ | O.K | O.K | O.K | O.K | [10 empty] | No
80 | BJBHT.A31R1 |  | BLECF_ | 495 | 0x | 01EF | 6 | Arc | BYPLM.A31R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
81 | BJBHT.A32R1 |  | BLECF_ | 384 | 0x | 0180 | 6 | Arc | BYPLM.A32R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
82 | BJBHT.A33R1 |  | BLECF_ | 417 | 0x | 01A1 | 6 | Arc | BYPLM.A33R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
83 | BJBHT.A34R1 |  | BLECF_ | 321 | 0x | 0141 | 6 | Arc | BYPLM.A34R1 | 1 | R |  |  |  | 7-9 | SR1 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
|  |  |  | 81 | 0x | 0051
Door
[27 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[27 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[7 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[7 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[7 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[27 empty] | O.K. | No problems observed
Card number wrong while checking on the 26.11
[46 empty] | 1102564 | BJBAP.B1L1=UX15
[46 empty] | 1102565 | BJBAP.A6L1=RI132
[46 empty] | 1102566 | BJBAP.A2L1=RB14
[46 empty] | 1102567 | BJBAP.A3L1=UJ14
[46 empty] | 1102568 | BJBAP.B5R1=RI171
[46 empty] | 1102569 | BJBAP.A6R1=RI171
[46 empty] | 1102570 | BJBAP.A9R1=RI18
[46 empty] | 1103757 | BJBAP.B6R1=RI171
[46 empty] | 1103758 | BJBAP.B6L1=RI132
Note on U79: effinger: 13.08.08 need to check if the fibre are changed to 5-6 --> checked 18.8.08 connected to 7-8
Note on U14: effinger: 13.08.08 need to check if the fibre are changed to 5-6 --> 18.8.08 was connected to 5&6 reconnected to 7&8
Note on E15: effinger: Card replaced because channels 3 -6 blocked to negative: BLECF_0283 replaced by BLECF_0212
Note on G84: Ewald Effinger: 0x013D = 317 BLECF_0317 is installed
Note on G76: Ewald Effinger: 0x0051 = 51 BLECF_0051 is installed
Note on G57: Ewald Effinger: 0x0163 to be checked with list of Ion
Note on AD83: Ewald Effinger: LF at 17:08 at 27/08/14 LF at 9:10 28.8 still the same
Note on AD68: Ewald Effinger: LF at 17:08 at 27/08/14 LF at 9:10 28.8 still the same
Note on AD69: Ewald Effinger: LF at 17:08 at 27/08/14 LF at 9:10 28.8 still the same
Note on AD70: Ewald Effinger: LF at 17:08 at 27/08/14 LF at 9:10 28.8 still the same
Note on G42: Ewald Effinger: 0x0191 to be check in Ions list ---> same number in the list of Ion!
Note on AT83: Ewald Effinger: After disconneting the HV on the first BLM noise dissapeared. Disconncted HV number 4 noise dispeared
Note on AU83: Ewald Effinger: Exchange of two HV connectors from the russian team, noise disappeared!
Note on AU64: Ewald Effinger: Test with Andrey: All signal cables discconected on the signal box and there was still noise! HV disconnected on the return and all channels have been negative and didn't recuperate. Next test disconnetion of NG 18 cable on the electronic side.
Note on AR50: Ewald Effinger: Installed XmasStop 2011/12
Note on AR51: Ewald Effinger: Installed XmasStop 2011/12
Note on AD86: Ewald Effinger: After reset (15/12/14) it looked o.k!
Note on AE22: Ewald Effinger: ↑↑
Note on AU44: Ewald Effinger: Level blinking -> HV o.k?, ground o.k? !!! New cable NES-18 wasn't connected! Now it is o.k !
Note on AU37: Ewald Effinger: Wrong number -> cheking of card number in the tunnel
Note on AU28: Ewald Effinger: Level blinking -> HV o.k?, ground o.k? !!! Isolation of a HV cable was toucinh the machine ground!
Note on AU22: Ewald Effinger: Wrong number -> cheking of card number in the tunnel
Note on AU18: Ewald Effinger: Wrong number -> cheking of card number in the tunnel
Note on AU19: Ewald Effinger: CFC reset needed -> call to verify (DAC reset needed (didn't reset with HV request))
Note on AU38: Ewald Effinger: After reset o.k. !!!
Note on AU41: Ewald Effinger: After reset o.k. !!!
Note on AU46: Ewald Effinger: After reset o.k. !!!
Note on AE38: Ewald Effinger: ↑↑↑
Note on AD38: Ewald Effinger: ↑↑↑
Note on AD41: Ewald Effinger: ↑↑↑
Note on AE41: Ewald Effinger: ↑↑↑
Note on AD46: Ewald Effinger: ↑↑↑
Note on AE46: Ewald Effinger: ↑↑↑
Note on AU45: Ewald Effinger: Ch4 and Ch5 (BLMQI.01L1.B1I30 & BLMQI.01L1.B2E10) are not modulating!!!!
Note on AU42: Ewald Effinger: CFC reset needed -> call to verify
Note on AU40: Ewald Effinger: CFC reset needed -> call to verify
Note on AU13: Ewald Effinger: CFC reset needed (High offset)-> call to verify
Note on AU61: Ewald Effinger: Channel 2 is blocked to 0 -> CFC ERROR is not on!!!, (DAC above 155) All the other channels are noisy!
Note on AU62: Ewald Effinger: Channel 3,4,5 is blocked to 0 -> CFC ERROR is not on!!! All the other channels are noisy!
Note on AU57: Ewald Effinger: Channel 3,4,5 is blocked to 0 -> CFC ERROR is not on!!! All the other channels are noisy!
Note on AU56: Ewald Effinger: Channel 3,4,5 is blocked to 0 -> CFC ERROR is not on!!! All the other channels are noisy!
Note on AU32: Ewald Effinger: Wrong number -> cheking of card number in the tunnel
Note on AU14: Ewald Effinger: Wrong number -> cheking of card number in the tunnel
Note on AU7: Ewald Effinger: Wrong number -> cheking of card number in the tunnel
Note on AU55: Ewald Effinger: After reset o.k. !!!
Note on AL62: Ewald Effinger: ch3-6 level was blinking. Gone but still noise
Note on AR62: Ewald Effinger: 1102570
Note on E46: Ewald Effinger: BLECF_0639
Note on AU48: Ewald Effinger: All Levels are blinking, a lot of noise on all channels Ch3,4,6,7 high offset

## Point 2
| Tunnel - position | [14 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L2 |  | BLECF_ | 399 | 0x | 018F | 6 | Arc | BYPLM.A33L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
2 | BJBHT.A32L2 |  | BLECF_ | 348 | 0x | 015C | 6 | Arc | BYPLM.A32L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
3 | BJBHT.A31L2 |  | BLECF_ | 497 | 0x | 01F1 | 6 | Arc | BYPLM.A31L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | CFC reset
4 | BJBHT.A30L2 |  | BLECF_ | 432 | 0x | 01B0 | 6 | Arc | BYPLM.A30L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
5 | BJBHT.A29L2 |  | BLECF_ | 520 | 0x | 0208 | 6 | Arc | BYPLM.A29L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | CFC reset
6 | BJBHT.A28L2 |  | BLECF_ | 439 | 0x | 01B7 | 6 | Arc | BYPLM.A28L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
7 | BJBHT.A27L2 |  | BLECF_ | 342 | 0x | 0156 | 6 | Arc | BYPLM.A27L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
8 | BJBHT.A26L2 |  | BLECF_ | 460 | 0x | 01CC | 6 | Arc | BYPLM.A26L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
9 | BJBHT.A25L2 |  | BLECF_ | 388 | 0x | 0184 | 6 | Arc | BYPLM.A25L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
10 | BJBHT.A24L2 |  | BLECF_ | 440 | 0x | 01B8 | 6 | Arc | BYPLM.A24L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
11 | BJBHT.A23L2 |  | BLECF_ | 468 | 0x | 01D4 | 6 | Arc | BYPLM.A23L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
12 | BJBHT.A22L2 |  | BLECF_ | 456 | 0x | 01C8 | 6 | Arc | BYPLM.A22L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
13 | BJBHT.A21L2 |  | BLECF_ | 111 | 0x | 006F | 6 | Arc | BYPLM.A21L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
14 | BJBHT.A20L2 |  | BLECF_ | 11 | 0x | 000B | 6 | Arc | BYPLM.A20L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
15 | BJBHT.A19L2 |  | BLECF_ | 715 | 0x | 02CB | 6 | Arc | BYPLM.A19L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
16 | BJBHT.A18L2 |  | BLECF_ | 366 | 0x | 016E | 6 | Arc | BYPLM.A18L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
17 | BJBHT.A17L2 |  | BLECF_ | 427 | 0x | 01AB | 6 | Arc | BYPLM.A17L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
18 | BJBHT.A16L2 |  | BLECF_ | 385 | 0x | 0181 | 6 | Arc | BYPLM.A16L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
19 | BJBHT.A15L2 |  | BLECF_ | 444 | 0x | 01BC | 6 | Arc | BYPLM.A15L2 | 2 | L |  |  |  | 5-6 | SR2 | BY05 | F | 4 | 5-6 | BY02 | L | 1 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
20 | BJBHT.A14L2 |  | BLECF_ | 380 | 0x | 017C | 6 | Arc | BYPLM.A14L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  |  |  | 1202897 | UA23 | BJBAP.C6L2=RH23
21 | BJBHT.A13L2 |  | BLECF_ | 347 | 0x | 015B | 6 | Arc | BYPLM.A13L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  |  |  | 1202898 | UA27 | BJBAP.A5R2=RA27
22 | BJBHT.A12L2 |  | BLECF_ | 200 | 0x | 00C8 | 8 | Arc | BYPLM.A12L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  |  |  | 1202899 | UA27 | BJBAP.B6R2=RA27
[48 empty] | 1202900 | UA27 | BJBAP.A11R2=R28
23 | BJBHT.C11L2 |  | BLECF_ | 583 | 0x | 0247 | 6 | UA23 | BY02 | 2 | CL | 3 | 2 | C | 19-20 | SR2 | BY05 | C | 4 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
24 | BJBHT.B11L2 | yes | BLECF_ | 343 | 0x | 0157 | 6 | UA23 | BY02 | 2 | CL | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
25 | BJBHT.A11L2 |  | BLECF_ | 378 | 0x | 017A | 6 | UA23 | BY02 | 2 | CL | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
26 | BJBHT.B10L2 |  | BLECF_ | 98 | 0x | 0062 | 6 | UA23 | BY02 | 2 | CL | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
27 | BJBHT.A10L2 |  | BLECF_ | 363 | 0x | 016B | 7 | UA23 | BY02 | 2 | CL | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
28 | BJBHT.A9L2 |  | BLECF_ | 517 | 0x | 0205 | 6 | UA23 | BY02 | 2 | CL | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
29 | BJBHT.C8L2 | yes | BLECF_ | 357 | 0x | 0165 | 6 | UA23 | BY02 | 2 | CL | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
30 | BJBHT.B8L2 | yes | BLECF_ | 333 | 0x | 014D | 6 | UA23 | BY02 | 2 | CL | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
31 | BJBHT.A8L2 | yes | BLECF_ | 491 | 0x | 01EB | 6 | UA23 | BY02 | 2 | CL | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | BLMEL.08L2.B2I23_MBA BJBAP.A8L2_7 no modulation LIC 11.02.15 (wrong connected goes to signalbox c8 channel 8) ->> done o.k. (also tested with the battery)
32 | BJBHT.A7L2 | yes | BLECF_ | 362 | 0x | 016A | 6 | UA23 | BY02 | 2 | CL | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | no mod 11.2 BLMQI.07L2.B2I20_MQM BJBAP.A7L2_3 Ch3 no modulation : BLMQI.07L2.B2I20_MQM 10.2
33 | BJBHT.C6L2 | yes | BLECF_ | 189 | 0x | 00BD | 6 | UA23 | BY02 | 2 | CL | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | High voltage return missing
34 | BJBHT.B6L2 | yes | BLECF_ | 344 | 0x | 0158 | 6 | UA23 | BY02 | 2 | CL | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
35 | BJBHT.A6L2 | yes | BLECF_ | 474 | 0x | 01DA | 6 | UA23 | BY02 | 2 | CL | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Noisy check instsallation --> reset needed ---> cabling redone (separating Signal gnd and HV gnd)
36 | BJBHT.A5L2 |  | BLECF_ | 328 | 0x | 0148 | 6 | UA23 | BY02 | 2 | CL | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
37 | BJBHT.D4L2 |  | BLECF_ | 371 | 0x | 0173 | 8 | UA23 | BY02 | 2 | CL | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
38 | BJBHT.C4L2 | yes | BLECF_ | 181 | 0x | 00B5 | 1 | UA23 | BY02 | 2 | CL | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No | ; | Ch3,Ch4,Ch5,Ch6 no modulation ; BLMEI.04L2.B2I10_MBRC; BLMEI.04L2.B2I20_MBRC; BLMTI.04L2.B1E10_TCTPH.4L2.B1; BLMTI.04L2.B1E10_TCTPV.4L2.B1
39 | BJBHT.B4L2 | yes | BLECF_ | 320 | 0x | 0140 | 4 | UA23 | BY02 | 2 | CL | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Noisy check instsallation --> reset needed
40 | BJBHT.A4L2 | yes | BLECF_ | 516 | 0x | 0204 | 7 | UA23 | BY02 | 2 | CL | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | BLMTI.04L2.B1E10_TCDD.4L2 BJBAP.A4L2_4 (it is connected to 5 instead of 4 ) -> Connection changed !! o.k. !!!!!
41 | BJBHT.A3L2 |  | BLECF_ | 324 | 0x | 0144 | 4 | UA23 | BY02 | 2 | CL | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | HV Missing
42 | BJBHT.A2L2 |  | BLECF_ | 188 | 0x | 00BC | 8 | UA23 | BY02 | 2 | CL | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | >
43 | BJBHT.B1L2 |  | BLECF_ | 510 | 0x | 01FE | 5 | UA23 | BY02 | 2 | CL | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | BLMEI.01L2.B1B10_BKGD BJBAP.B1L2_8 ---
44 | BJBHT.A1L2 | yes | BLECF_ | 476 | 0x | 01DC | 1 | UA23 | BY02 | 2 | CL | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
45 | BJBHT.A1R2 | yes | BLECF_ | 706 | 0x | 02C2 | 1 | UA27 | BY03 | 2 | CR | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
46 | BJBHT.B1R2 |  | BLECF_ | 606 | 0x | 025E | 5 | UA27 | BY03 | 2 | CR | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
47 | BJBHT.A2R2 |  | BLECF_ | 675 | 0x | 02A3 | 8 | UA27 | BY03 | 2 | CR | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
48 | BJBHT.A3R2 |  | BLECF_ | 641 | 0x | 0281 | 4 | UA27 | BY03 | 2 | CR | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Noisy check of installation (same noise with NG18 disconnected on the BJBAT side ---> Problem of the NG18 cable)
49 | BJBHT.A4R2 |  | BLECF_ | 719 | 0x | 02CF | 5 | UA27 | BY03 | 2 | CR | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
50 | BJBHT.B4R2 | yes | BLECF_ | 738 | 0x | 02E2 | 2 | UA27 | BY03 | 2 | CR | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Problem CFC (repaired by Ion) reinstallation CFC -> reset done
51 | BJBHT.C4R2 | yes | BLECF_ | 652 | 0x | 028C | 1 | UA27 | BY03 | 2 | CR | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
52 | BJBHT.D4R2 |  | BLECF_ | 665 | 0x | 0299 | 8 | UA27 | BY03 | 2 | CR | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
53 | BJBHT.A5R2 |  | BLECF_ | 740 | 0x | 02E4 | 6 | UA27 | BY03 | 2 | CR | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | Verify new NES18 is connected (there exist already a NES18)
54 | BJBHT.A6R2 | yes | BLECF_ | 666 | 0x | 029A | 1 | UA27 | BY03 | 2 | CR | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
55 | BJBHT.B6R2 | yes | BLECF_ | 722 | 0x | 02D2 | 7 | UA27 | BY03 | 2 | CR | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | Verify NES18 is connected
56 | BJBHT.A7R2 |  | BLECF_ | 658 | 0x | 0292 | 6 | UA27 | BY03 | 2 | CR | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
57 | BJBHT.A8R2 |  | BLECF_ | 626 | 0x | 0272 | 6 | UA27 | BY03 | 2 | CR | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
58 | BJBHT.B8R2 | yes | BLECF_ | 406 | 0x | 0196 | 6 | UA27 | BY03 | 2 | CR | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
59 | BJBHT.C8R2 |  | BLECF_ | 4 | 0x | 0004 | 6 | UA27 | BY03 | 2 | CR | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
60 | BJBHT.A9R2 |  | BLECF_ | 279 | 0x | 0117 | 6 | UA27 | BY03 | 2 | CR | 2 | 6 | C | 7 | SR2 | BY05 | C | 10 | 7 | BY02 | R | 3 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
61 | BJBHT.A10R2 |  | BLECF_ | 97 | 0x | 0061 | 6 | UA27 | BY03 | 2 | CR | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
62 | BJBHT.B10R2 |  | BLECF_ | 721 | 0x | 02D1 | 6 | UA27 | BY03 | 2 | CR | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
63 | BJBHT.A11R2 |  | BLECF_ | 128 | 0x | 0080 | 6 | UA27 | BY03 | 2 | CR | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No |  | Verify NES18 is connected
64 | BJBHT.B11R2 | yes | BLECF_ | 627 | 0x | 0273 | 6 | UA27 | BY03 | 2 | CR | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
65 | BJBHT.C11R2 |  | BLECF_ | 689 | 0x | 02B1 | 6 | UA27 | BY03 | 2 | CR | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | DAC Values high (Ch3,Ch8) --- CFC Reset needed --> call to verify --> ch3,8 to be checked again
66 | BJBHT.A12R2 |  | BLECF_ | 681 | 0x | 02A9 | 8 | Arc | BYPLM.A12R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
67 | BJBHT.A13R2 |  | BLECF_ | 687 | 0x | 02AF | 6 | Arc | BYPLM.A13R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
68 | BJBHT.A14R2 |  | BLECF_ | 684 | 0x | 02AC | 6 | Arc | BYPLM.A14R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Some channels are not displayed in the Fix display --. To be checked after database update --> Mobile monitors
69 | BJBHT.A15R2 |  | BLECF_ | 484 | 0x | 01E4 | 6 | Arc | BYPLM.A15R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
70 | BJBHT.A16R2 |  | BLECF_ | 686 | 0x | 02AE | 6 | Arc | BYPLM.A16R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
71 | BJBHT.A17R2 |  | BLECF_ | 143 | 0x | 008F | 6 | Arc | BYPLM.A17R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
72 | BJBHT.A18R2 |  | BLECF_ | 463 | 0x | 01CF | 6 | Arc | BYPLM.A18R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | CFC reset ---> call for verification
73 | BJBHT.A19R2 |  | BLECF_ | 379 | 0x | 017B | 6 | Arc | BYPLM.A19R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
74 | BJBHT.A20R2 |  | BLECF_ | 382 | 0x | 017E | 6 | Arc | BYPLM.A20R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
75 | BJBHT.A21R2 |  | BLECF_ | 451 | 0x | 01C3 | 6 | Arc | BYPLM.A21R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | CFC reset ---> call for verification
76 | BJBHT.A22R2 |  | BLECF_ | 404 | 0x | 0194 | 6 | Arc | BYPLM.A22R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
77 | BJBHT.A23R2 |  | BLECF_ | 55 | 0x | 0037 | 6 | Arc | BYPLM.A23R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
78 | BJBHT.A24R2 |  | BLECF_ | 418 | 0x | 01A2 | 6 | Arc | BYPLM.A24R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | DAC Values high --- CFC Reset needed --> call to verify
79 | BJBHT.A25R2 |  | BLECF_ | 390 | 0x | 0186 | 6 | Arc | BYPLM.A25R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
80 | BJBHT.A26R2 |  | BLECF_ | 214 | 0x | 00D6 | 6 | Arc | BYPLM.A26R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
81 | BJBHT.A27R2 |  | BLECF_ | 694 | 0x | 02B6 | 6 | Arc | BYPLM.A27R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
82 | BJBHT.A28R2 |  | BLECF_ | 685 | 0x | 02AD | 6 | Arc | BYPLM.A28R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
83 | BJBHT.A29R2 |  | BLECF_ | 725 | 0x | 02D5 | 6 | Arc | BYPLM.A29R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
84 | BJBHT.A30R2 |  | BLECF_ | 195 | 0x | 00C3 | 6 | Arc | BYPLM.A30R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
85 | BJBHT.A31R2 |  | BLECF_ | 690 | 0x | 02B2 | 6 | Arc | BYPLM.A31R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
86 | BJBHT.A32R2 |  | BLECF_ | 82 | 0x | 0052 | 6 | Arc | BYPLM.A32R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
87 | BJBHT.A33R2 |  | BLECF_ | 751 | 0x | 02EF | 6 | Arc | BYPLM.A33R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
88 | BJBHT.A34R2 |  | BLECF_ | 724 | 0x | 02D4 | 6 | Arc | BYPLM.A34R2 | 2 | R |  |  |  | 7-9 | SR2 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
89 |  |  | BLECF_ | 288 | 0x | 0120 | 7 | Arc | BYPLM.A34R2 | 3 | R |  |  |  | 7-10 | SR3 | BY06 | I | 13 | 7-9 | BY03 | R | 3 | 20 | 16 | 1-9 | [19 empty] | Mobile System removed (3.2.15) | o
| Door | [44 empty] | 1202897 |  | BJBAP.C6L2=RH23
[27 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k | [8 empty] | 1202898 |  | BJBAP.A5R2=RA27
[27 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current) | [8 empty] | 1202899 |  | BJBAP.B6R2=RA27
[7 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing | [17 empty] | 1202900 |  | BJBAP.A11R2=R28
[7 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[7 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[27 empty] | O.K. | No problems observed
[45 empty] | L2 | BLMEI.04L2.B2I10_MBRC | BLMEI.04L2.B2I10_MBRC:LOSS_RS09
[45 empty] | L2 | BLMEI.04L2.B2I20_MBRC | BLMEI.04L2.B2I20_MBRC:LOSS_RS09
[45 empty] | L2 | BLMEL.08L2.B2I23_MBA | BLMEL.08L2.B2I23_MBA:LOSS_RS09
[45 empty] | L2 | BLMQI.07L2.B2I20_MQM | BLMQI.07L2.B2I20_MQM:LOSS_RS09
[45 empty] | L2 | BLMTI.04L2.B1E10_TCDD.4L2 | BLMTI.04L2.B1E10_TCDD.4L2:LOSS_RS09
[45 empty] | L2 | BLMTI.04L2.B1E10_TCTPH.4L2.B1 | BLMTI.04L2.B1E10_TCTPH.4L2.B1:LOSS_RS09
[45 empty] | L2 | BLMTI.04L2.B1E10_TCTPV.4L2.B1 | BLMTI.04L2.B1E10_TCTPV.4L2.B1:LOSS_RS09
Note on G81: effinger: Checked-----> BLECF_0382 installed! Wrong number 0x017E (382) instead of 0x0188 (392) ; to be check in the tunnel; Card number checked: --> BLECF_0382 installed; sent number 382
Note on U23: effinger: Fibre 7 & 8 broken, spare fibre used! Still the case after LS1!!!!
Note on AD48: Ewald Effinger: CFC working checked with the fiber connection of card 1 Problem must me in the surface! Fibers where not plugged in properly
Note on AD46: Ewald Effinger: CFC working problem at the surface Fibers where not plugged in properly
Note on AD43: Ewald Effinger: CFC working problem at the surface Fibers where not plugged in properly
Note on AD39: Ewald Effinger: Fiber mostprobably internaly not changed to the new configuration!! ! Verify with Steve what he has changed! LF due to bad connection (10) in the patch panel (insight) Fiber 9-10 swaped on the patch panel Fiber mixed up in the TC.
Note on AD36: Ewald Effinger: Fiber mostprobably internaly not changed to the new configuration!! ! Verify with Steve what he has changed!
Note on F32: Ewald Effinger: 0x0168 installed! In the list from Ion the old number was 0050
Note on AT43: Ewald Effinger: HV ok. By using a HV jumper!
Note on AJ29: Ewald Effinger: cable was missing, hidden in the cable tray.
Note on AU42: Ewald Effinger: Missing BLM LIC ,
Note on AU10: Ewald Effinger: A lot of noise an the 6 connected channels! Andrey made a reset and checked the HV. All loocked fine.
Note on AS42: Ewald Effinger: Reset done by Andrey
Note on AT41: Ewald Effinger: Andrey could not check since the cable is on the cealing! WE wait till all is installed.
Note on AD78: Ewald Effinger: 0 LF @16:30 1/10/2014 after a CPU reset still 0 LF after 23 hours!
Note on E84: Ewald Effinger: was BLECF_0085
Note on E93: Ewald Effinger: The old number was 360 ??
Note on AB53: Ewald Effinger: ↑↑↑
Note on AF53: Ewald Effinger: ↑↑↑
Note on AF63: Ewald Effinger: ↑↑↑
Note on AG63: Ewald Effinger: ↑↑↑
Note on AF64: Ewald Effinger: ↑↑↑
Note on AG64: Ewald Effinger: ↑↑↑
Note on AH63: Ewald Effinger: ↑↑↑
Note on AH64: Ewald Effinger: ↑↑↑
Note on AJ69: Ewald Effinger: Noise on all channels! Cable changed by NES18
Note on AT55: Ewald Effinger: HV was missing
Note on AT56: Ewald Effinger: HV was missing. Repaired and o.k. Rechecked and HV missing since 2.2.15 Andrey checked on the HV divider and 3 V a present.
Note on AT57: Ewald Effinger: HV was missing
Note on AD65: Ewald Effinger: ↑↑↑
Note on AE66: Ewald Effinger: ↑↑↑
Note on U66: Ewald Effinger: Second fiber : C 9 - 7
Note on P66: Ewald Effinger: Second fibre: C 19
Note on AT69: Ewald Effinger: HV was missing, repaired!
Note on AR47: Ewald Effinger: Installed XmasStop 2011/12
Note on AR45: Ewald Effinger: Installed XmasStop 2011/12
Note on AR53: Ewald Effinger: Installed XmasStop 2011/12
Note on AR52: Ewald Effinger: Installed XmasStop 2011/12
Note on AU55: Ewald Effinger: BLMs where missing the russian team installed them 14.10
Note on AU56: Ewald Effinger: BLMs where missing the russian team installed them 14.10
Note on AU57: Ewald Effinger: BLMs where missing the russian team installed them 14.10
Note on AU53: Ewald Effinger: (Fibers wrongly plugged !)
Note on AU63: Ewald Effinger: Order of fiber in the TC needed to be adjusted
Note on AU64: Ewald Effinger: Fiber where inverted insight the TC. Ok now!
Note on AU65: Ewald Effinger: Fibre between BY02 and BY05 brocken
Note on AU66: Ewald Effinger: Fibre between tunnel and BY05 brocken
Note on AU95: Ewald Effinger: HV connector ?
Note on AU94: Ewald Effinger: BLM exchanged (P1027) due to problem with HV connector
Note on AU88: Ewald Effinger: HV missing. O.k. now!
Note on AU79: Ewald Effinger: Noise on the channel 6 , o.k. now and no problem while RSM BMLEI.B2E10_Drift (modulation signal is very small) --- > Chamber was removed
Note on AR59: Ewald Effinger: Already 1 NES18 Installed in the XmasStop 2011/12
Note on AU69: Ewald Effinger: Noise on all channels! Cable changed by NES18
Note on AU61: Ewald Effinger: Noise on all channels! Cable changed by NES18
Note on AU59: Ewald Effinger: Noise on all channels! Cable changed by NES18
Note on AU41: Ewald Effinger: Voltage missing, because of missing BLM (MKI support) -.o.k
Note on AU48: Ewald Effinger: cable exchanged
Note on C45: Ewald Effinger: Fliter already existed. Replaced with new filter!
Note on AT38: Ewald Effinger: HV was missing
Note on AT44: Ewald Effinger: HV was missing
Note on AJ43: Ewald Effinger: Level blinking on all channels
Note on AJ44: Ewald Effinger: Noise on all channels --> Cabling scheme changed!
Note on C60: Ewald Effinger: needed
Note on C38: Ewald Effinger: old filter
Note on C34: Ewald Effinger: old filter should be rplaced e
Note on AJ40: Ewald Effinger: All Levels were blinking
Note on AU43: Ewald Effinger: Noisy check instsallation and reset needed (separating Signal gnd and HV gnd)

## Point 3
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L3 | BLECF_ | 619 | 0x | 026B | 6 | Arc | BYPLM.A33L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Ch 5 high offset!! 0 DAC
2 | BJBHT.A32L3 | BLECF_ | 601 | 0x | 0259 | 6 | Arc | BYPLM.A32L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
3 | BJBHT.A31L3 | BLECF_ | 727 | 0x | 02D7 | 6 | Arc | BYPLM.A31L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
4 | BJBHT.A30L3 | BLECF_ | 705 | 0x | 02C1 | 6 | Arc | BYPLM.A30L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Problem with RSM (Huge peek after conecting the source) All concetors verified!
5 | BJBHT.A29L3 | BLECF_ | 646 | 0x | 0286 | 6 | Arc | BYPLM.A29L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
6 | BJBHT.A28L3 | BLECF_ | 702 | 0x | 02BE | 6 | Arc | BYPLM.A28L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
7 | BJBHT.A27L3 | BLECF_ | 607 | 0x | 025F | 6 | Arc | BYPLM.A27L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
8 | BJBHT.A26L3 | BLECF_ | 623 | 0x | 026F | 6 | Arc | BYPLM.A26L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
9 | BJBHT.A25L3 | BLECF_ | 717 | 0x | 02CD | 6 | Arc | BYPLM.A25L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
10 | BJBHT.A24L3 | BLECF_ | 319 | 0x | 013F | 6 | Arc | BYPLM.A24L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
11 | BJBHT.A23L3 | BLECF_ | 745 | 0x | 02E9 | 6 | Arc | BYPLM.A23L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
12 | BJBHT.A22L3 | BLECF_ | 645 | 0x | 0285 | 6 | Arc | BYPLM.A22L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
13 | BJBHT.A21L3 | BLECF_ | 741 | 0x | 02E5 | 6 | Arc | BYPLM.A21L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
14 | BJBHT.A20L3 | BLECF_ | 614 | 0x | 0266 | 6 | Arc | BYPLM.A20L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
15 | BJBHT.A19L3 | BLECF_ | 620 | 0x | 026C | 6 | Arc | BYPLM.A19L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
16 | BJBHT.A18L3 | BLECF_ | 720 | 0x | 02D0 | 6 | Arc | BYPLM.A18L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
17 | BJBHT.A17L3 | BLECF_ | 668 | 0x | 029C | 6 | Arc | BYPLM.A17L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
18 | BJBHT.A16L3 | BLECF_ | 604 | 0x | 025C | 6 | Arc | BYPLM.A16L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
19 | BJBHT.A15L3 | BLECF_ | 622 | 0x | 026E | 6 | Arc | BYPLM.A15L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
20 | BJBHT.A14L3 | BLECF_ | 305 | 0x | 0131 | 6 | Arc | BYPLM.A14L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
21 | BJBHT.A13L3 | BLECF_ | 600 | 0x | 0258 | 10 | Arc | BYPLM.A13L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | RSM strange signal seen! All connection have been checked ! To be remeasured
22 | BJBHT.A12L3 | BLECF_ | 703 | 0x | 02BF | 10 | Arc | BYPLM.A12L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
23 | BJBHT.B11L3 | BLECF_ | 657 | 0x | 0291 | 7 | UJ33 | BY02 | 3 | L | 2 | 8 | C | 11-12 | SR3 | BY05 | C | 3 | 11-12 | BY02 | L | 1 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
24 | BJBHT.A11L3 | BLECF_ | 654 | 0x | 028E | 12 | UJ33 | BY02 | 3 | L | 2 | 7 | C | 9-10 | SR3 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | All channels are noisy (cable length)
25 | BJBHT.A10L3 | BLECF_ | 574 | 0x | 023E | 11 | UJ33 | BY02 | 3 | L | 2 | 6 | C | 7-8 | SR3 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
26 | BJBHT.B9L3 | BLECF_ | 674 | 0x | 02A2 | 11 | UJ33 | BY02 | 3 | L | 2 | 5 | C | 5-6 | SR3 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
27 | BJBHT.A9L3 | BLECF_ | 609 | 0x | 0261 | 10 | UJ33 | BY02 | 3 | L | 2 | 4 | C | 3-4 | SR3 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
28 | BJBHT.B8L3 | BLECF_ | 571 | 0x | 023B | 6 | UJ33 | BY02 | 3 | L | 2 | 3 | C | 1-2 | SR3 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | CH 1 and Ch5 blocked to 0 --> now high offset !? (Reset nedded !) Card needs to be checked -> same offset with NG18 removed! -> o.k. now
29 | BJBHT.A8L3 | BLECF_ | 103 | 0x | 0067 | 6 | UJ33 | BY02 | 3 | L | 2 | 2 | B | 23-24 | SR3 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
30 | BJBHT.B7L3 | BLECF_ | 576 | 0x | 0240 | 7 | UJ33 | BY02 | 3 | L | 2 | 1 | B | 21-22 | SR3 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | HV was o.k after the first check Ch1-Ch4 high offset!!! CFC reset!!!
31 | BJBHT.A7L3 | BLECF_ | 637 | 0x | 027D | 1 | UJ33 | BY02 | 3 | L | 1 | 10 | B | 19-20 | SR3 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | HV divider ? Hv divifer was missing
32 | BJBHT.C6L3 | BLECF_ | 211 | 0x | 00D3 | 6 | UJ33 | BY02 | 3 | L | 1 | 9 | B | 17-18 | SR3 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
33 | BJBHT.B6L3 | BLECF_ | 629 | 0x | 0275 | 4 | UJ33 | BY02 | 3 | L | 1 | 8 | B | 15-16 | SR3 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Ch1 to Ch5 high offset! CFC reset -> call for verification
34 | BJBHT.A6L3 | BLECF_ | 536 | 0x | 0218 | 4 | UJ33 | BY02 | 3 | L | 1 | 7 | B | 13-14 | SR3 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
35 | BJBHT.D5L3 | BLECF_ | 175 | 0x | 00AF | 2 | UJ33 | BY02 | 3 | L | 1 | 6 | B | 11-12 | SR3 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
36 | BJBHT.C5L3 | BLECF_ | 577 | 0x | 0241 | 8 | UJ33 | BY02 | 3 | L | 1 | 5 | B | 9-10 | SR3 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
37 | BJBHT.B5L3 | BLECF_ | 575 | 0x | 023F | 6 | UJ33 | BY02 | 3 | L | 1 | 4 | B | 7-8 | SR3 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
38 | BJBHT.A5L3 | BLECF_ | 602 | 0x | 025A | 6 | UJ33 | BY02 | 3 | L | 1 | 3 | B | 5-6 | SR3 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
39 | BJBHT.B4L3 | BLECF_ | 711 | 0x | 02C7 | 2 | UJ33 | BY02 | 3 | L | 1 | 2 | B | 3-4 | SR3 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
40 | BJBHT.A4L3 | BLECF_ | 729 | 0x | 02D9 | 8 | UJ33 | BY02 | 3 | L | 1 | 1 | B | 1-2 | SR3 | BY05 | C | 1 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
41 | BJBHT.A4R3 | BLECF_ | 543 | 0x | 021F | 8 | UJ33 | BY03 | 3 | R | 1 | 1 | C | 13-14 | SR3 | BY05 | C | 4 | 1-2 | BY02 | R | 3 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
42 | BJBHT.B4R3 | BLECF_ | 677 | 0x | 02A5 | 2 | UJ33 | BY03 | 3 | R | 1 | 2 | C | 15-16 | SR3 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
43 | BJBHT.A5R3 | BLECF_ | 182 | 0x | 00B6 | 6 | UJ33 | BY03 | 3 | R | 1 | 3 | C | 17-18 | SR3 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
44 | BJBHT.B5R3 | BLECF_ | 326 | 0x | 0146 | 6 | UJ33 | BY03 | 3 | R | 1 | 4 | C | 19-20 | SR3 | BY05 | C | 4 | 7-8 | BY02 | C | 2 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
45 | BJBHT.C5R3 | BLECF_ | 487 | 0x | 01E7 | 8 | UJ33 | BY03 | 3 | R | 1 | 5 | C | 21-22 | SR3 | BY05 | C | 4 | 9-10 | BY02 | R | 3 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
46 | BJBHT.D5R3 | BLECF_ | 393 | 0x | 0189 | 2 | UJ33 | BY03 | 3 | R | 1 | 6 | C | 23-24 | SR3 | BY05 | C | 4 | 11-12 | BY02 | C | 2 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
47 | BJBHT.A6R3 | BLECF_ | 541 | 0x | 021D | 4 | UJ33 | BY03 | 3 | R | 1 | 7 | D | 1-2 | SR3 | BY05 | C | 5 | 1-2 | BY02 | C | 2 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
48 | BJBHT.B6R3 | BLECF_ | 341 | 0x | 0155 | 4 | UJ33 | BY03 | 3 | R | 1 | 8 | D | 3-4 | SR3 | BY05 | C | 5 | 3-4 | BY02 | C | 2 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
49 | BJBHT.C6R3 | BLECF_ | 339 | 0x | 0153 | 6 | UJ33 | BY03 | 3 | R | 1 | 9 | D | 5-6 | SR3 | BY05 | C | 5 | 5-6 | BY02 | R | 3 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
50 | BJBHT.A7R3 | BLECF_ | 367 | 0x | 016F | 1 | UJ33 | BY03 | 3 | R | 1 | 10 | D | 7-8 | SR3 | BY05 | C | 5 | 7-8 | BY02 | C | 2 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Divider missing, Ion ? Done!!
51 | BJBHT.B7R3 | BLECF_ | 709 | 0x | 02C5 | 7 | UJ33 | BY03 | 3 | R | 2 | 1 | D | 9-10 | SR3 | BY05 | C | 5 | 9-10 | BY02 | R | 3 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | no modulation signal
52 | BJBHT.A8R3 | BLECF_ | 502 | 0x | 01F6 | 6 | UJ33 | BY03 | 3 | R | 2 | 2 | D | 11-12 | SR3 | BY05 | C | 5 | 11-12 | BY02 | C | 2 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
53 | BJBHT.B8R3 | BLECF_ | 183 | 0x | 00B7 | 6 | UJ33 | BY03 | 3 | R | 2 | 3 | D | 13-14 | SR3 | BY05 | C | 6 | 1-2 | BY02 | R | 3 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
54 | BJBHT.A9R3 | BLECF_ | 375 | 0x | 0177 | 10 | UJ33 | BY03 | 3 | R | 2 | 4 | D | 15-16 | SR3 | BY05 | C | 6 | 3-4 | BY02 | C | 2 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
55 | BJBHT.B9R3 | BLECF_ | 177 | 0x | 00B1 | 11 | UJ33 | BY03 | 3 | R | 2 | 5 | D | 17-18 | SR3 | BY05 | C | 6 | 5-6 | BY02 | R | 3 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | no |  | Checked ? Connected !! | 1300380 | UJ33 | BJBAP.B9R3=R38
56 | BJBHT.A10R3 | BLECF_ | 186 | 0x | 00BA | 11 | UJ33 | BY03 | 3 | R | 2 | 6 | D | 19-20 | SR3 | BY05 | C | 6 | 7-8 | BY02 | R | 3 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | RESET needed! -> done o.k. !!!!
57 | BJBHT.A11R3 | BLECF_ | 562 | 0x | 0232 | 12 | UJ33 | BY03 | 3 | R | 2 | 7 | D | 21-22 | SR3 | BY05 | C | 6 | 9-10 | BY02 | C | 2 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
58 | BJBHT.B11R3 | BLECF_ | 459 | 0x | 01CB | 7 | UJ33 | BY03 | 3 | R | 2 | 8 | D | 23-24 | SR3 | BY05 | C | 6 | 11-12 | BY02 | R | 3 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | no |  | no modulation signal
59 | BJBHT.A12R3 | BLECF_ | 68 | 0x | 0044 | 10 | Arc | BYPLM.A12R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | card is constanty reseted!!! ----> Reset PCB excahged, now it is working!
60 | BJBHT.A13R3 | BLECF_ | 527 | 0x | 020F | 10 | Arc | BYPLM.A13R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | was 611, removed to install at 12r6 !!! no modulation signal on 4 BLM2I !!!
61 | BJBHT.A14R3 | BLECF_ | 504 | 0x | 01F8 | 6 | Arc | BYPLM.A14R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
62 | BJBHT.A15R3 | BLECF_ | 499 | 0x | 01F3 | 6 | Arc | BYPLM.A15R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
63 | BJBHT.A16R3 | BLECF_ | 514 | 0x | 0202 | 6 | Arc | BYPLM.A16R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
64 | BJBHT.A17R3 | BLECF_ | 493 | 0x | 01ED | 6 | Arc | BYPLM.A17R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
65 | BJBHT.A18R3 | BLECF_ | 496 | 0x | 01F0 | 6 | Arc | BYPLM.A18R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
66 | BJBHT.A19R3 | BLECF_ | 314 | 0x | 013A | 6 | Arc | BYPLM.A19R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | Ch 3 shows CFC error (but level o.k) Reset and verification, (ch3 is blocked!? Steve)
67 | BJBHT.A20R3 | BLECF_ | 523 | 0x | 020B | 6 | Arc | BYPLM.A20R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
68 | BJBHT.A21R3 | BLECF_ | 185 | 0x | 00B9 | 6 | Arc | BYPLM.A21R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
69 | BJBHT.A22R3 | BLECF_ | 387 | 0x | 0183 | 6 | Arc | BYPLM.A22R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
70 | BJBHT.A23R3 | BLECF_ | 485 | 0x | 01E5 | 6 | Arc | BYPLM.A23R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
71 | BJBHT.A24R3 | BLECF_ | 309 | 0x | 0135 | 6 | Arc | BYPLM.A24R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
72 | BJBHT.A25R3 | BLECF_ | 184 | 0x | 00B8 | 6 | Arc | BYPLM.A25R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
73 | BJBHT.A26R3 | BLECF_ | 337 | 0x | 0151 | 6 | Arc | BYPLM.A26R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
74 | BJBHT.A27R3 | BLECF_ | 519 | 0x | 0207 | 6 | Arc | BYPLM.A27R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
75 | BJBHT.A28R3 | BLECF_ | 448 | 0x | 01C0 | 6 | Arc | BYPLM.A28R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no |  | no modulation signal
76 | BJBHT.A29R3 | BLECF_ | 489 | 0x | 01E9 | 6 | Arc | BYPLM.A29R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
77 | BJBHT.A30R3 | BLECF_ | 492 | 0x | 01EC | 6 | Arc | BYPLM.A30R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
78 | BJBHT.A31R3 | BLECF_ | 488 | 0x | 01E8 | 6 | Arc | BYPLM.A31R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
79 | BJBHT.A32R3 | BLECF_ | 407 | 0x | 0197 | 6 | Arc | BYPLM.A32R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
80 | BJBHT.A33R3 | BLECF_ | 453 | 0x | 01C5 | 6 | Arc | BYPLM.A33R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
81 | BJBHT.A34R3 | BLECF_ | 109 | 0x | 006D | 6 | Arc | BYPLM.A34R3 | 3 | R |  |  |  | 6-7 | SR3 | BY05 | I | 12 | 6-7 | BY02 | R | 3 | 20 | 16 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | no
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on AD88: Ewald Effinger: was on 7-8
Note on AD85: Ewald Effinger: LF cleaned and reconnected than it was o.k.
Note on AE66: Ewald Effinger: Fibers mixed in the TC
Note on AL10: Ewald Effinger: Signal was disconnected
Note on D15: Ewald Effinger: was BLECF_0730
Note on D30: Ewald Effinger: was 0754
Note on D34: Ewald Effinger: was 574
Note on AT32: Ewald Effinger: Comment from Matti after RSM: • Signals from monitors connected to BJBHT.A11L3 and BJBHT.A9L3 are swapped Fiber must be exchanged!!!! Fiber exchanged by Steph 15/10 o.k. ! RSM
Note on AT29: Ewald Effinger: Comment from Matti after RSM: • Signals from monitors connected to BJBHT.A11L3 and BJBHT.A9L3 are swapped Fiber must be exchanged!!!! Fiber exchanged by Steph 15/10 o.k. ! RSM ! Divider missing
Note on AT15: Ewald Effinger: Comment from Matti after RSM: • Signals from monitors connected to BJBHT.A23L3 and BJBHT.A22L3 are swapped Fiber should be inversed in the surface! (RSM confirms the swap) Fiber exchanged by Steph 15/10 o.k. ! RSM
Note on AT16: Ewald Effinger: Comment from Matti after RSM: • Signals from monitors connected to BJBHT.A23L3 and BJBHT.A22L3 are swapped Fiber should be inversed in the surface! (RSM confirms the swap) Fiber exchanged by Steph 15/10 o.k. ! RSM
Note on AS26: Ewald Effinger: HV distribution need to be checked!
Note on F66: Ewald Effinger: was 472
Note on AE60: Ewald Effinger: ↑↑↑
Note on AF60: Ewald Effinger: ↑↑↑
Note on AG60: Ewald Effinger: ↑↑↑
Note on AE63: Ewald Effinger: ↑↑↑
Note on AF63: Ewald Effinger: ↑↑↑
Note on AG63: Ewald Effinger: ↑↑↑
Note on AG47: Ewald Effinger: 1
Note on AQ64: Ewald Effinger: Installed in september 2011
Note on AS66: Ewald Effinger: Problem HV
Note on AS57: Ewald Effinger: bad connect
Note on AS67: Ewald Effinger: 1 wrong connection, 1 Connector badly made
Note on AS82: Ewald Effinger: Cable not connected! Reset done on the card!
Note on AS64: Ewald Effinger: HV problem: mix up of signal cables NES18 and NG18 connected
Note on AS63: Ewald Effinger: HV problem
Note on D61: Ewald Effinger: BLECF_0410 excahnged by BLECF_0177 due to problem on HV input.
Note on AS45: Ewald Effinger: Cable bad connected!
Note on AT9: Ewald Effinger: HV was disconnected
Note on AT10: Ewald Effinger: Signal and HV was disconnected
Note on AT26: Ewald Effinger: O.k. now, only the feedback was not connected correctly! RSM showed o.k!
Note on AT45: Ewald Effinger: HV missing o.k. cable not correctly connected
Note on AT31: Ewald Effinger: Divider missing (ion installed!)
Note on D33: Ewald Effinger: BLECF_0672 exchange with BLECF_0571

## Point 4
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L4 | BLECF_ | 452 | 0x | 01C4 | 6 | Arc | BYPLM.A33L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8
2 | BJBHT.A32L4 | BLECF_ | 748 | 0x | 02EC | 6 | Arc | BYPLM.A32L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | [19 empty] | Ch4 goes to 0 !!! All DAC values are 0!!! CFC Reset needed -- > call to verify
3 | BJBHT.A31L4 | BLECF_ | 616 | 0x | 0268 | 6 | Arc | BYPLM.A31L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8
4 | BJBHT.A30L4 | BLECF_ | 662 | 0x | 0296 | 6 | Arc | BYPLM.A30L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16
5 | BJBHT.A29L4 | BLECF_ | 38 | 0x | 0026 | 6 | Arc | BYPLM.A29L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8
6 | BJBHT.A28L4 | BLECF_ | 76 | 0x | 004C | 6 | Arc | BYPLM.A28L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16
7 | BJBHT.A27L4 | BLECF_ | 572 | 0x | 023C | 6 | Arc | BYPLM.A27L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8
8 | BJBHT.A26L4 | BLECF_ | 21 | 0x | 0015 | 6 | Arc | BYPLM.A26L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16
9 | BJBHT.A25L4 | BLECF_ | 613 | 0x | 0265 | 6 | Arc | BYPLM.A25L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8
10 | BJBHT.A24L4 | BLECF_ | 655 | 0x | 028F | 6 | Arc | BYPLM.A24L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | [19 empty] | Ch 5 shows CFC error (but level o.k) Reset and verification(is blocked!? Steve)
11 | BJBHT.A23L4 | BLECF_ | 631 | 0x | 0277 | 6 | Arc | BYPLM.A23L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8
12 | BJBHT.A22L4 | BLECF_ | 743 | 0x | 02E7 | 6 | Arc | BYPLM.A22L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16
13 | BJBHT.A21L4 | BLECF_ | 664 | 0x | 0298 | 6 | Arc | BYPLM.A21L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8
14 | BJBHT.A20L4 | BLECF_ | 471 | 0x | 01D7 | 6 | Arc | BYPLM.A20L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16
15 | BJBHT.A19L4 | BLECF_ | 552 | 0x | 0228 | 6 | Arc | BYPLM.A19L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8
16 | BJBHT.A18L4 | BLECF_ | 395 | 0x | 018B | 6 | Arc | BYPLM.A18L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16
17 | BJBHT.A17L4 | BLECF_ | 615 | 0x | 0267 | 6 | Arc | BYPLM.A17L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | [19 empty] | High DAC -> CFC reset -> call for verification
18 | BJBHT.A16L4 | BLECF_ | 591 | 0x | 024F | 6 | Arc | BYPLM.A16L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16
19 | BJBHT.A15L4 | BLECF_ | 642 | 0x | 0282 | 6 | Arc | BYPLM.A15L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8
20 | BJBHT.A14L4 | BLECF_ | 670 | 0x | 029E | 6 | Arc | BYPLM.A14L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16
21 | BJBHT.A13L4 | BLECF_ | 599 | 0x | 0257 | 6 | Arc | BYPLM.A13L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8
22 | BJBHT.A12L4 | BLECF_ | 659 | 0x | 0293 | 6 | Arc | BYPLM.A12L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16
23 | BJBHT.A5L4 | BLECF_ | 251 | 0x | 00FB | 0 | UA43 | BY12 | 4 | CL | 1 | 9 | A | 17-18 | SX4 | BY04 | C | 2 | 5-6 | BY02 | C | 2 | 4 | 1 | 9-16 | [16 empty] | NES18 |  |  | BLMs are connected to A5R4; BLECF_0251 to be removed!!!!
24 | BJBHT.A11L4 | BLECF_ | 643 | 0x | 0283 | 6 | UA43 | BY12 | 4 | CL | 1 | 8 | A | 15-16 | SX4 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 7 | 4 | 1-8 | [19 empty] | High DAC -> CFC reset -> call for verification
25 | BJBHT.A10L4 | BLECF_ | 578 | 0x | 0242 | 6 | UA43 | BY12 | 4 | CL | 1 | 7 | A | 13-14 | SX4 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 6 | 3 | 9-16 | [16 empty] | NES18 |  |  |  | BJBAP.A10L4
26 | BJBHT.A9L4 | BLECF_ | 408 | 0x | 0198 | 6 | UA43 | BY12 | 4 | CL | 1 | 6 | A | 11-12 | SX4 | BY05 | C | 1 | 11-12 | BY02 | L | 1 | 6 | 3 | 1-8
27 | BJBHT.A8L4 | BLECF_ | 651 | 0x | 028B | 6 | UA43 | BY12 | 4 | CL | 1 | 5 | A | 9-10 | SX4 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16
28 | BJBHT.A7L4 | BLECF_ | 386 | 0x | 0182 | 6 | UA43 | BY12 | 4 | CL | 1 | 4 | A | 7-8 | SX4 | BY05 | C | 1 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | [19 empty] | High DAC -> CFC reset -> call for verification
29 | BJBHT.A6L4 | BLECF_ | 157 | 0x | 009D | 6 | UA43 | BY12 | 4 | CL | 1 | 3 | A | 5-6 | SX4 | BY05 | C | 1 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | [19 empty] | High DAC -> CFC reset -> call for verification
30 | BJBHT.C5L4 | BLECF_ | 744 | 0x | 02E8 | 8 | UA43 | BY12 | 4 | CL | 1 | 2 | A | 3-4 | SX4 | BY05 | C | 1 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8 | [19 empty] | High DAC -> CFC reset -> call for verification
31 | BJBHT.B5L4 | BLECF_ | 624 | 0x | 0270 | 2 | UA43 | BY12 | 4 | CL | 1 | 1 | A | 1-2 | SX4 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8
32 | BJBHT.B5R4 | BLECF_ | 8 | 0x | 0008 | 2 | UA47 | BY09 | 4 | CR | 1 | 1 | A | 1-2 | SX4 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | [21 empty] | 1401043 | UA43 | BJBAP.C5L4=RA43
33 | BJBHT.C5R4 | BLECF_ | 5 | 0x | 0005 | 8 | UA47 | BY09 | 4 | CR | 1 | 2 | A | 3-4 | SX4 | BY05 | C | 12 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8 | [16 empty] | NES18 |  |  | NES 18 connected ? Needs to be checked! | BJBAP.C5R4 | 1401044 | UA43 | BJBAP.A10L4=R42
34 | BJBHT.A6R4 | BLECF_ | 316 | 0x | 013C | 6 | UA47 | BY09 | 4 | CR | 1 | 3 | A | 5-6 | SX4 | BY05 | C | 12 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | [21 empty] | 1401045 | UA47 | BJBAP.C5R4=RA47
35 | BJBHT.A7R4 | BLECF_ | 235 | 0x | 00EB | 6 | UA47 | BY09 | 4 | CR | 1 | 4 | A | 7-8 | SX4 | BY05 | C | 12 | 7-8 | BY02 | R | 3 | 5 | 2 | 1-8 | [16 empty] | NES18 |  |  | CFC reset needed -> call to verify & NES 18 connected ? Needs to be checked! | BJBAP.A7R4 | 1401046 | UA47 | BJBAP.A7R4=RA47
36 | BJBHT.A8R4 | BLECF_ | 584 | 0x | 0248 | 6 | UA47 | BY09 | 4 | CR | 1 | 5 | A | 9-10 | SX4 | BY05 | C | 12 | 9-10 | BY02 | R | 3 | 5 | 2 | 9-16 | [19 empty] | High DAC -> CFC reset -> call for verification
37 | BJBHT.A9R4 | BLECF_ | 509 | 0x | 01FD | 6 | UA47 | BY09 | 4 | CR | 1 | 6 | A | 11-12 | SX4 | BY05 | C | 12 | 11-12 | BY02 | R | 3 | 6 | 3 | 1-8
38 | BJBHT.A10R4 | BLECF_ | 7 | 0x | 0007 | 6 | UA47 | BY09 | 4 | CR | 1 | 7 | A | 13-14 | SX4 | BY05 | C | 11 | 1-2 | BY02 | R | 3 | 6 | 3 | 9-16 | [19 empty] | CFC reset needed -> call to verify
39 | BJBHT.A11R4 | BLECF_ | 10 | 0x | 000A | 6 | UA47 | BY09 | 4 | CR | 1 | 8 | A | 15-16 | SX4 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 7 | 4 | 1-8
40 | BJBHT.A5R4 | BLECF_ | 119 | 0x | 0077 | 0 | UA47 | BY09 | 4 | CR | 1 | 9 | A | 17-18 | SX4 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 13 | 9 | 9-16 | [19 empty] | Channel 3 ( IC from A5L4 will be connected!
41 | BJBHT.A12R4 | BLECF_ | 217 | 0x | 00D9 | 6 | Arc | BYPLM.A12R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 7 | 4 | 9-16 | OK | OK | OK | OK | OK | OK | OK
42 | BJBHT.A13R4 | BLECF_ | 531 | 0x | 0213 | 6 | Arc | BYPLM.A13R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | OK | OK | OK | OK | OK | OK | OK
43 | BJBHT.A14R4 | BLECF_ | 568 | 0x | 0238 | 6 | Arc | BYPLM.A14R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | OK | OK | OK | OK | OK | OK | OK
44 | BJBHT.A15R4 | BLECF_ | 57 | 0x | 0039 | 6 | Arc | BYPLM.A15R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | OK | OK | OK | OK | OK | OK | OK
45 | BJBHT.A16R4 | BLECF_ | 57 | 0x | 0039 | 6 | Arc | BYPLM.A16R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | OK | OK | OK | OK | OK | OK | OK
46 | BJBHT.A17R4 | BLECF_ | 403 | 0x | 0193 | 6 | Arc | BYPLM.A17R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | OK | ↑↑↑ | OK | OK | ↑↑↑ | OK | OK | [12 empty] | BLECF_0403 exchanged with BLECF_0251 form A5L4, same problems remains Error B and Error C -----> BLETC needs to be exchanged
47 | BJBHT.A18R4 | BLECF_ | 558 | 0x | 022E | 6 | Arc | BYPLM.A18R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | OK | OK | OK | OK | OK | OK | OK
48 | BJBHT.A19R4 | BLECF_ | 304 | 0x | 0130 | 6 | Arc | BYPLM.A19R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | OK | OK | OK | OK | OK | OK | OK
49 | BJBHT.A20R4 | BLECF_ | 585 | 0x | 0249 | 6 | Arc | BYPLM.A20R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | OK | OK | OK | OK | OK | OK | OK
50 | BJBHT.A21R4 | BLECF_ | 315 | 0x | 013B | 6 | Arc | BYPLM.A21R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | OK | OK | OK | OK | OK | OK | OK
51 | BJBHT.A22R4 | BLECF_ | 340 | 0x | 0154 | 6 | Arc | BYPLM.A22R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | OK | OK | OK | OK | OK | OK | OK
52 | BJBHT.A23R4 | BLECF_ | 537 | 0x | 0219 | 6 | Arc | BYPLM.A23R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | OK | OK | OK | OK | OK | OK | OK
53 | BJBHT.A24R4 | BLECF_ | 405 | 0x | 0195 | 6 | Arc | BYPLM.A24R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [12 empty] | High DAC -> CFC reset -> call for verification
54 | BJBHT.A25R4 | BLECF_ | 483 | 0x | 01E3 | 6 | Arc | BYPLM.A25R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | OK | OK | OK | OK | OK | OK | OK
55 | BJBHT.A26R4 | BLECF_ | 438 | 0x | 01B6 | 6 | Arc | BYPLM.A26R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | OK | OK | OK | OK | OK | OK | OK
56 | BJBHT.A27R4 | BLECF_ | 547 | 0x | 0223 | 6 | Arc | BYPLM.A27R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | OK | OK | OK | OK | OK | OK | OK
57 | BJBHT.A28R4 | BLECF_ | 551 | 0x | 0227 | 6 | Arc | BYPLM.A28R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | OK | OK | OK | OK | OK | OK | OK
58 | BJBHT.A29R4 | BLECF_ | 83 | 0x | 0053 | 6 | Arc | BYPLM.A29R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | OK | OK | OK | OK | OK | OK | OK
59 | BJBHT.A30R4 | BLECF_ | 121 | 0x | 0079 | 6 | Arc | BYPLM.A30R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | OK | OK | OK | OK | OK | OK | OK
60 | BJBHT.A31R4 | BLECF_ | 40 | 0x | 0028 | 6 | Arc | BYPLM.A31R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | OK | OK | OK | OK | OK | OK | OK
61 | BJBHT.A32R4 | BLECF_ | 230 | 0x | 00E6 | 6 | Arc | BYPLM.A32R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [12 empty] | Ch 2 shows CFC error (but level o.k) Reset and verification(is blocked!? Steve)
62 | BJBHT.A33R4 | BLECF_ | 582 | 0x | 0246 | 6 | Arc | BYPLM.A33R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [12 empty] | High DAC -> CFC reset -> call for verification
63 | BJBHT.A34R4 | BLECF_ | 389 | 0x | 0185 | 6 | Arc | BYPLM.A34R4 | 4 | R |  |  |  | 7-9 | SX4 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | OK | OK | OK | OK | OK | OK | OK
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on D12: Ewald Effinger: Number from the calibration list of Ion
Note on D41: Ewald Effinger: Number from the calibration list of Ion
Note on D68: Ewald Effinger: Number from the calibration list of Ion
Note on D66: Ewald Effinger: Number from the calibration list of Ion
Note on D55: Ewald Effinger: Number from the calibration list of Ion
Note on D52: Ewald Effinger: Number (BLECF_0134) from the calibration list of Ion. Installed BLECF_0057
Note on F52: Ewald Effinger: BLECF_0057 (x0039) installed
Note on AD46: Ewald Effinger: ↑↑↑ o.k 9.12.2014
Note on AS56: Ewald Effinger: HV problem
Note on AS24: Ewald Effinger: HV Problem ; bad connection
Note on AS25: Ewald Effinger: HV Problem ; bad connection
Note on AS23: Ewald Effinger: HV Problem ; bad connection
Note on AS21: Ewald Effinger: HV Problem ; bad connection
Note on AS20: Ewald Effinger: HV Problem ; bad connection
Note on AS15: Ewald Effinger: HV Problem ; bad connection
Note on AS13: Ewald Effinger: HV Problem ; bad connection
Note on AC8: Ewald Effinger: ↑↑↑: fibre bad connected
Note on AD6: Ewald Effinger: ↑↑↑: fibre bad connected
Note on AS9: Ewald Effinger: hv problem
Note on AS16: Ewald Effinger: HV Problem ; bad connection
Note on AS17: Ewald Effinger: HV Problem ; bad connection
Note on AD9: Ewald Effinger: ↑↑↑: GOH modul was not correctly plugged
Note on AD34: Ewald Effinger: ↑↑↑
Note on AC19: Ewald Effinger: power problem ↑↑↑
Note on AD19: Ewald Effinger: ↑↑↑
Note on AC23: Ewald Effinger: some LF seen, counter reset at 18.12.2014 - 16:49
Note on D28: Ewald Effinger: BLECF_0251
Note on AT9: Ewald Effinger: GOH2 ready signal off! GOH replugged then o.k!
Note on AT30: Ewald Effinger: NES 18 connected ? Needs to be checked!
Note on AT34: Ewald Effinger: BLM Channel 7 is missing ? Checked with expert application and display, problem disapearded
Note on AT33: Ewald Effinger: BLM Channel 7 is missing ? Checked with expert application and display, problem disapearded
Note on AT29: Ewald Effinger: BLM Channel 7 is missing ? Checked with expert application and display, problem disapearded
Note on AT42: Ewald Effinger: BLM Channel 7 is missing ? Checked with expert application and display, problem disapearded
Note on AQ28: Ewald Effinger: NES18 pulled but to the wrond desitination C5L5
Note on AT69: Ewald Effinger: BLM Channel 7 is missing ? Checked with expert application and display, problem disapearded
Note on AT60: Ewald Effinger: BLM Channel 7 is missing ? Checked with expert application and display, problem disapearded
Note on AT21: Ewald Effinger: BLM Channel 7 is missing ? Checked with expert application and display, problem disapearded
Note on AT35: Ewald Effinger: BLM Channel 7 is missing ? Checked with expert application and display, problem disapearded
Note on D53: Ewald Effinger: was BLECF_0403 with BLECF_0251 form A5L4 same problems remain Error B and Error C

## Point 5
| Tunnel - position | [14 empty] | Surface position | [19 empty] | LEVEL
| HV box | Filter | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | Comment | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | HV | 1024 | Comments
Pos. | position |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L5 |  | BLECF_ | 335 | 0x | 014F | 6 | Arc | BYPLM.A33L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16
2 | BJBHT.A32L5 |  | BLECF_ | 330 | 0x | 014A | 6 | Arc | BYPLM.A32L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | Test in tunnel needed! | [19 empty] | Reset needed
3 | BJBHT.A31L5 |  | BLECF_ | 650 | 0x | 028A | 6 | Arc | BYPLM.A31L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | Test in tunnel needed! | [19 empty] | Reset needed
4 | BJBHT.A30L5 |  | BLECF_ | 414 | 0x | 019E | 6 | Arc | BYPLM.A30L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8
5 | BJBHT.A29L5 |  | BLECF_ | 307 | 0x | 0133 | 6 | Arc | BYPLM.A29L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | [20 empty] | 220 V ? Was working !
6 | BJBHT.A28L5 |  | BLECF_ | 447 | 0x | 01BF | 6 | Arc | BYPLM.A28L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8
7 | BJBHT.A27L5 |  | BLECF_ | 226 | 0x | 00E2 | 6 | Arc | BYPLM.A27L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16
8 | BJBHT.A26L5 |  | BLECF_ | 680 | 0x | 02A8 | 6 | Arc | BYPLM.A26L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | Test in tunnel needed! | [19 empty] | Reset needed
9 | BJBHT.A25L5 |  | BLECF_ | 424 | 0x | 01A8 | 6 | Arc | BYPLM.A25L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16
10 | BJBHT.A24L5 |  | BLECF_ | 419 | 0x | 01A3 | 6 | Arc | BYPLM.A24L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8
11 | BJBHT.A23L5 |  | BLECF_ | 34 | 0x | 0022 | 6 | Arc | BYPLM.A23L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | Test in tunnel needed!
12 | BJBHT.A22L5 |  | BLECF_ | 625 | 0x | 0271 | 6 | Arc | BYPLM.A22L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | Test in tunnel needed! | [19 empty] | Reset needed
13 | BJBHT.A21L5 |  | BLECF_ | 506 | 0x | 01FA | 6 | Arc | BYPLM.A21L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | Test in tunnel needed! | [19 empty] | Ch 2 shows CFC error after Reset (is blocked!? Steve)
14 | BJBHT.A20L5 |  | BLECF_ | 589 | 0x | 024D | 6 | Arc | BYPLM.A20L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8
15 | BJBHT.A19L5 |  | BLECF_ | 546 | 0x | 0222 | 6 | Arc | BYPLM.A19L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16
16 | BJBHT.A18L5 |  | BLECF_ | 473 | 0x | 01D9 | 6 | Arc | BYPLM.A18L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | Test in tunnel needed!
17 | BJBHT.A17L5 |  | BLECF_ | 475 | 0x | 01DB | 6 | Arc | BYPLM.A17L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16
18 | BJBHT.A16L5 |  | BLECF_ | 196 | 0x | 00C4 | 6 | Arc | BYPLM.A16L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8
19 | BJBHT.A15L5 |  | BLECF_ | 346 | 0x | 015A | 6 | Arc | BYPLM.A15L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16
20 | BJBHT.A14L5 |  | BLECF_ | 48 | 0x | 0030 | 6 | Arc | BYPLM.A14L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8
21 | BJBHT.A13L5 |  | BLECF_ | 640 | 0x | 0280 | 6 | Arc | BYPLM.A13L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | Test in tunnel needed!
22 | BJBHT.A12L5 |  | BLECF_ | 477 | 0x | 01DD | 6 | Arc | BYPLM.A12L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8
23 | BJBHT.B11L5 |  | BLECF_ | 753 | 0x | 02F1 | 6 | RR53 | BY02 | 5 | CL | 2 | 9 | B | 13-14 | SR5 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16
24 | BJBHT.A11L5 |  | BLECF_ | 581 | 0x | 0245 | 7 | RR53 | BY02 | 5 | CL | 2 | 8 | B | 11-12 | SR5 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8
25 | BJBHT.B10L5 |  | BLECF_ | 14 | 0x | 000E | 6 | RR53 | BY02 | 5 | CL | 2 | 7 | B | 9-10 | SR5 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 7 | 4 | 1-8
26 | BJBHT.A10L5 |  | BLECF_ | 124 | 0x | 007C | 2 | RR53 | BY02 | 5 | CL | 2 | 6 | B | 7-8 | SR5 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16
27 | BJBHT.A9L5 |  | BLECF_ | 757 | 0x | 02F5 | 6 | RR53 | BY02 | 5 | CL | 2 | 5 | B | 5-6 | SR5 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16
28 | BJBHT.B8L5 |  | BLECF_ | 190 | 0x | 00BE | 6 | RR53 | BY02 | 5 | CL | 2 | 4 | B | 3-4 | SR5 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8
29 | BJBHT.A8L5 |  | BLECF_ | 9 | 0x | 0009 | 6 | RR53 | BY02 | 5 | CL | 2 | 3 | B | 1-2 | SR5 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | [17 empty] | NES18
30 | BJBHT.A7L5 |  | BLECF_ | 525 | 0x | 020D | 6 | RR53 | BY02 | 5 | CL | 2 | 2 | A | 23-24 | SR5 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 9-16
31 | BJBHT.B6L5 |  | BLECF_ | 632 | 0x | 0278 | 7 | RR53 | BY02 | 5 | CL | 2 | 1 | A | 21-22 | SR5 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 1-8 | [17 empty] | NES18
32 | BJBHT.A6L5 |  | BLECF_ | 535 | 0x | 0217 | 1 | RR53 | BY02 | 5 | CL | 1 | 10 | A | 19-20 | SR5 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16 | Test in tunnel needed!
33 | BJBHT.B5L5 |  | BLECF_ | 313 | 0x | 0139 | 6 | RR53 | BY02 | 5 | CL | 1 | 9 | A | 17-18 | SR5 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | [17 empty] | NES18
34 | BJBHT.A5L5 |  | BLECF_ | 521 | 0x | 0209 | 2 | RR53 | BY02 | 5 | CL | 1 | 8 | A | 15-16 | SR5 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8
35 | BJBHT.C4L5 |  | BLECF_ | 18 | 0x | 0012 | 6 | RR53 | BY02 | 5 | CL | 1 | 7 | A | 13-14 | SR5 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | [17 empty] | NES18
36 | BJBHT.B4L5 |  | BLECF_ | 370 | 0x | 0172 | 7 | RR53 | BY02 | 5 | CL | 1 | 6 | A | 11-12 | SR5 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16
37 | BJBHT.A4L5 |  | BLECF_ | 3 | 0x | 0003 | 2 | RR53 | BY02 | 5 | CL | 1 | 5 | A | 9-10 | SR5 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | [17 empty] | NES18
38 | BJBHT.A3L5 |  | BLECF_ | 72 | 0x | 0048 | 5 | RR53 | BY02 | 5 | CL | 1 | 4 | A | 7-8 | SR5 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | [17 empty] | NES18 |  |  | NES18
39 | BJBHT.A2L5 |  | BLECF_ | 288 | 0x | 0120 | 8 | RR53 | BY02 | 5 | CL | 1 | 3 | A | 5-6 | SR5 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | [17 empty] | NES18 |  |  | Exchange BLMQI.02L5.B1I23_MQXB (channel 7 ) since it behaves strange while modulation ---> CFC needs to be checked | Access Monday 17.2
40 | BJBHT.B1L5 |  | BLECF_ | 586 | 0x | 024A | 4 | RR53 | BY02 | 5 | CL | 1 | 2 | A | 3-4 | SR5 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | [20 empty] | reset needed since HV DAC reset didn't work
41 | BJBHT.A1L5 | yes | BLECF_ | 579 | 0x | 0243 | 2 | RR53 | BY02 | 5 | CL | 1 | 1 | A | 1-2 | SR5 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | [20 empty] | BLMEL.01L5.B2I10_BPMSW.1L5 (was not modulating? To be verified ) | Access Monday 17.2 | BLMES.01L5.B2E10_BPMSW.1L5
42 | BJBHT.A1R5 | yes | BLECF_ | 193 | 0x | 00C1 | 2 | RR57 | BY02 | 5 | CR | 1 | 1 | A | 1-2 | SR5 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | [20 empty] | BLMEL.01R5.B2I10_BPMSW.1R5 | Access Monday 17.2
43 | BJBHT.B1R5 |  | BLECF_ | 749 | 0x | 02ED | 6 | RR57 | BY02 | 5 | CR | 1 | 2 | A | 3-4 | SR5 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | [17 empty] | NES18 |  |  | NES18
44 | BJBHT.A2R5 |  | BLECF_ | 671 | 0x | 029F | 6 | RR57 | BY02 | 5 | CR | 1 | 3 | A | 5-6 | SR5 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8
45 | BJBHT.A3R5 |  | BLECF_ | 750 | 0x | 02EE | 6 | RR57 | BY02 | 5 | CR | 1 | 4 | A | 7-8 | SR5 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16
46 | BJBHT.A4R5 |  | BLECF_ | 726 | 0x | 02D6 | 2 | RR57 | BY02 | 5 | CR | 1 | 5 | A | 9-10 | SR5 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8
47 | BJBHT.B4R5 |  | BLECF_ | 663 | 0x | 0297 | 8 | RR57 | BY02 | 5 | CR | 1 | 6 | A | 11-12 | SR5 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16
48 | BJBHT.C4R5 |  | BLECF_ | 710 | 0x | 02C6 | 6 | RR57 | BY02 | 5 | CR | 1 | 7 | A | 13-14 | SR5 | BY05 | C | 11 | 1-2 | BY02 | R | 3 | 4 | 1 | 1-8 | [22 empty] | 1518964 | RR53 | BJBAP.A3L5=R542
49 | BJBHT.A5R5 |  | BLECF_ | 598 | 0x | 0256 | 2 | RR57 | BY02 | 5 | CR | 1 | 8 | A | 15-16 | SR5 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | [22 empty] | 1518965 | RR53 | BJBAP.A4L5=R532
50 | BJBHT.B5R5 |  | BLECF_ | 697 | 0x | 02B9 | 6 | RR57 | BY02 | 5 | CR | 1 | 9 | A | 17-18 | SR5 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | [20 empty] | Channel 7 (constant 10 blmbits @40us) |  | 1518966 | RR53 | BJBAP.C4L5=R532
51 | BJBHT.A6R5 |  | BLECF_ | 648 | 0x | 0288 | 2 | RR57 | BY02 | 5 | CR | 1 | 10 | A | 19-20 | SR5 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16 | [22 empty] | 1518967 | RR53 | BJBAP.B5L5=R532
52 | BJBHT.B6R5 |  | BLECF_ | 656 | 0x | 0290 | 6 | RR57 | BY02 | 5 | CR | 2 | 1 | A | 21-22 | SR5 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 5 | 2 | 1-8 | [17 empty] | NES18 |  |  | NES18 |  | 1518968 | RR53 | BJBAP.B6L5=R532
53 | BJBHT.A7R5 |  | BLECF_ | 383 | 0x | 017F | 6 | RR57 | BY02 | 5 | CR | 2 | 2 | A | 23-24 | SR5 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16 | [22 empty] | 1518969 | RR53 | BJBAP.A8L5=R52
54 | BJBHT.A8R5 |  | BLECF_ | 596 | 0x | 0254 | 6 | RR57 | BY02 | 5 | CR | 2 | 3 | B | 1-2 | SR5 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | [22 empty] | 1518970 | RR57 | BJBAP.A1R5=R561
55 | BJBHT.B8R5 |  | BLECF_ | 660 | 0x | 0294 | 6 | RR57 | BY02 | 5 | CR | 2 | 4 | B | 3-4 | SR5 | BY05 | C | 10 | 3-4 | BY02 | R | 3 | 6 | 3 | 1-8 | [22 empty] | 1518971 | RR57 | BJBAP.B6R5=R571
56 | BJBHT.A9R5 |  | BLECF_ | 149 | 0x | 0095 | 6 | RR57 | BY02 | 5 | CR | 2 | 5 | B | 5-6 | SR5 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 9-16 | [22 empty] | 1518972 | RR57 | BJBAP.B10R5=R58
57 | BJBHT.A10R5 |  | BLECF_ | 605 | 0x | 025D | 2 | RR57 | BY02 | 5 | CR | 2 | 6 | B | 7-8 | SR5 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | [22 empty] | 1520398 | RR53 | BJBAP.B9L5=R52
58 | BJBHT.B10R5 |  | BLECF_ | 673 | 0x | 02A1 | 6 | RR57 | BY02 | 5 | CR | 2 | 7 | B | 9-10 | SR5 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 7 | 4 | 1-8 | [17 empty] | NES18
59 | BJBHT.A11R5 |  | BLECF_ | 621 | 0x | 026D | 8 | RR57 | BY02 | 5 | CR | 2 | 8 | B | 11-12 | SR5 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8
60 | BJBHT.B11R5 |  | BLECF_ | 603 | 0x | 025B | 6 | RR57 | BY02 | 5 | CR | 2 | 9 | B | 13-14 | SR5 | BY05 | C | 9 | 1-2 | BY02 | R | 3 | 7 | 4 | 9-16 | [20 empty] | All channles up to 100 blmbits @40us
61 | BJBHT.A12R5 |  | BLECF_ | 699 | 0x | 02BB | 6 | Arc | BYPLM.A12R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | Test in tunnel needed!
62 | BJBHT.A13R5 |  | BLECF_ | 728 | 0x | 02D8 | 6 | Arc | BYPLM.A13R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16
63 | BJBHT.A14R5 |  | BLECF_ | 698 | 0x | 02BA | 6 | Arc | BYPLM.A14R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8
64 | BJBHT.A15R5 |  | BLECF_ | 737 | 0x | 02E1 | 6 | Arc | BYPLM.A15R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16
65 | BJBHT.A16R5 |  | BLECF_ | 733 | 0x | 02DD | 6 | Arc | BYPLM.A16R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8
66 | BJBHT.A17R5 |  | BLECF_ | 634 | 0x | 027A | 6 | Arc | BYPLM.A17R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16
67 | BJBHT.A18R5 |  | BLECF_ | 732 | 0x | 02DC | 6 | Arc | BYPLM.A18R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8
68 | BJBHT.A19R5 |  | BLECF_ | 742 | 0x | 02E6 | 6 | Arc | BYPLM.A19R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | [20 empty] | Channel 4 (peaks up to 50 blmbits @40us)
69 | BJBHT.A20R5 |  | BLECF_ | 695 | 0x | 02B7 | 6 | Arc | BYPLM.A20R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | Test in tunnel needed!
70 | BJBHT.A21R5 |  | BLECF_ | 688 | 0x | 02B0 | 6 | Arc | BYPLM.A21R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | [20 empty] | Reset needed
71 | BJBHT.A22R5 |  | BLECF_ | 739 | 0x | 02E3 | 6 | Arc | BYPLM.A22R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | [20 empty] | Reset needed
72 | BJBHT.A23R5 |  | BLECF_ | 691 | 0x | 02B3 | 6 | Arc | BYPLM.A23R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | Test in tunnel needed!
73 | BJBHT.A24R5 |  | BLECF_ | 617 | 0x | 0269 | 6 | Arc | BYPLM.A24R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | [20 empty] | Reset needed
74 | BJBHT.A25R5 |  | BLECF_ | 736 | 0x | 02E0 | 6 | Arc | BYPLM.A25R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | [20 empty] | Reset needed
75 | BJBHT.A26R5 |  | BLECF_ | 683 | 0x | 02AB | 6 | Arc | BYPLM.A26R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | [20 empty] | Reset needed
76 | BJBHT.A27R5 |  | BLECF_ | 490 | 0x | 01EA | 6 | Arc | BYPLM.A27R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | [20 empty] | Reset needed
77 | BJBHT.A28R5 |  | BLECF_ | 50 | 0x | 0032 | 6 | Arc | BYPLM.A28R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8
78 | BJBHT.A29R5 |  | BLECF_ | 701 | 0x | 02BD | 6 | Arc | BYPLM.A29R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | [20 empty] | Reset needed
79 | BJBHT.A30R5 |  | BLECF_ | 704 | 0x | 02C0 | 6 | Arc | BYPLM.A30R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8
80 | BJBHT.A31R5 |  | BLECF_ | 693 | 0x | 02B5 | 6 | Arc | BYPLM.A31R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | [20 empty] | Reset needed
81 | BJBHT.A32R5 |  | BLECF_ | 735 | 0x | 02DF | 6 | Arc | BYPLM.A32R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | [20 empty] | Reset needed
82 | BJBHT.A33R5 |  | BLECF_ | 731 | 0x | 02DB | 6 | Arc | BYPLM.A33R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16
83 | BJBHT.A34R5 |  | BLECF_ | 661 | 0x | 0295 | 6 | Arc | BYPLM.A34R5 | 5 | R |  |  |  | 7-9 | SR5 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8
Door
[26 empty] | ↑↑↑ | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[27 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[7 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[7 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[7 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[27 empty] | O.K. | No problems observed
Note on AB68: effinger: both jumper not correctly pluged, Card o.k.,
Note on AB76: effinger: jumper not correctly pluged, Card o.k.,
Note on AB79: effinger: jumper not correctly pluged, Card o.k.,
Note on AT80: Ewald Effinger: HV not correctly connected.
Note on AP51: Ewald Effinger: Checked for 5 min, I didn't see anymore the level on this channel blinkimg.
Note on AT85: Ewald Effinger: HV connector with bad
Note on AP85: Ewald Effinger: Level blinking for ch1 -6
Note on E84: Ewald Effinger: Number from the calibration list of Ion
Note on AS44: Ewald Effinger: Installed XmasStop 2011/12
Note on AV43: Ewald Effinger: Installed XmasStop 2011/12
Note on AV49: Ewald Effinger: Installed XmasStop 2011/12
Note on AV58: Ewald Effinger: Installed XmasStop 2011/12
Note on E37: Ewald Effinger: Number from the calibration list of Ion
Note on E40: Ewald Effinger: Number from the calibration list of Ion
Note on E43: Ewald Effinger: Number from the calibration list of Ion
Note on E44: Ewald Effinger: Number from the calibration list of Ion blecf_0042 exchanged BLECF_0288
Note on E58: Ewald Effinger: Number from the calibration list of Ion
Note on AT76: Ewald Effinger: Problem with HV
Note on AT15: Ewald Effinger: Problem HV: BLM dismanteled!
Note on AE37: Ewald Effinger: ↑↑↑ due to a broken fiber from the patch down to upstairs. Spare fiber used!
Note on U37: Ewald Effinger: C - 4 - 3 to B - 15
Note on P37: Ewald Effinger: C - 4 - 3 to B - 15
Note on AT14: Ewald Effinger: Adrey checked all connections, no o.k.
Note on AT13: Ewald Effinger: Adrey checked all connections, no o.k.
Note on AT10: Ewald Effinger: HV problem now o.k.
Note on AT12: Ewald Effinger: HV problem now o.k.
Note on AT17: Ewald Effinger: Ion and Adrien
Note on AT22: Ewald Effinger: Problem HV: BLM dismanteled!
Note on AT25: Ewald Effinger: Ion and Adrien : BLM reinstalled o.k.
Note on AT8: Ewald Effinger: HV problem
Note on AE9: Ewald Effinger: ↑↑↑ After removing and repluging the CFC it worked!
Note on AF9: Ewald Effinger: ↑↑↑
Note on AT26: Ewald Effinger: HV problems
Note on AT45: Ewald Effinger: HV problem
Note on AV85: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV83: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV81: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV80: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV35: Ewald Effinger: Ch4 no Modulation :BLMBI.08L5B0T10_MBB_MBA -> cable badly connected!
Note on AT48: Ewald Effinger: NG18 cable was not connected
Note on AQ12: Ewald Effinger: Ch7: Level was o.k CFC error was on! Card checked in the lab, since no improvement after a reset. Card was o.k. -> Card reinstalled and all wasd o.k. Again 2 resets done and all was o.k.
Note on AV88: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV87: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV82: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV78: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV77: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV17: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV16: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV6: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV7: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)
Note on AV12: Ewald Effinger: Blinking (all channels peaks up to 50 blmbits @40us)

## Point 6
| Tunnel - position | [14 empty] | Surface position | [19 empty] | LEVEL
| HV box | Filter | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | Comment | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L6 |  | BLECF_ | 533 | 0x | 0215 | 6 | Arc | BYPLM.A33L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
2 | BJBHT.A32L6 |  | BLECF_ | 117 | 0x | 0075 | 6 | Arc | BYPLM.A32L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
3 | BJBHT.A31L6 |  | BLECF_ | 232 | 0x | 00E8 | 6 | Arc | BYPLM.A31L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
4 | BJBHT.A30L6 |  | BLECF_ | 93 | 0x | 005D | 6 | Arc | BYPLM.A30L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | card o.k
5 | BJBHT.A29L6 |  | BLECF_ | 163 | 0x | 00A3 | 6 | Arc | BYPLM.A29L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
6 | BJBHT.A28L6 |  | BLECF_ | 549 | 0x | 0225 | 6 | Arc | BYPLM.A28L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
7 | BJBHT.A27L6 |  | BLECF_ | 166 | 0x | 00A6 | 6 | Arc | BYPLM.A27L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | dac ovrflow
8 | BJBHT.A26L6 |  | BLECF_ | 266 | 0x | 010A | 6 | Arc | BYPLM.A26L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
9 | BJBHT.A25L6 |  | BLECF_ | 162 | 0x | 00A2 | 6 | Arc | BYPLM.A25L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
10 | BJBHT.A24L6 |  | BLECF_ | 168 | 0x | 00A8 | 6 | Arc | BYPLM.A24L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
11 | BJBHT.A23L6 |  | BLECF_ | 177 | 0x | 00B1 | 6 | Arc | BYPLM.A23L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
12 | BJBHT.A22L6 |  | BLECF_ | 197 | 0x | 00C5 | 6 | Arc | BYPLM.A22L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
13 | BJBHT.A21L6 |  | BLECF_ | 170 | 0x | 00AA | 6 | Arc | BYPLM.A21L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
14 | BJBHT.A20L6 |  | BLECF_ | 179 | 0x | 00B3 | 6 | Arc | BYPLM.A20L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
15 | BJBHT.A19L6 |  | BLECF_ | 178 | 0x | 00B2 | 6 | Arc | BYPLM.A19L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
16 | BJBHT.A18L6 |  | BLECF_ | 207 | 0x | 00CF | 6 | Arc | BYPLM.A18L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
17 | BJBHT.A17L6 |  | BLECF_ | 125 | 0x | 007D | 6 | Arc | BYPLM.A17L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
18 | BJBHT.A16L6 |  | BLECF_ | 708 | 0x | 02C4 | 6 | Arc | BYPLM.A16L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
19 | BJBHT.A15L6 |  | BLECF_ | 299 | 0x | 012B | 6 | Arc | BYPLM.A15L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
20 | BJBHT.A14L6 |  | BLECF_ | 281 | 0x | 0119 | 6 | Arc | BYPLM.A14L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Ch5: BLMQI.14L6.B1E30_MQ was connected on ch8 ; to be verified
21 | BJBHT.A13L6 |  | BLECF_ | 268 | 0x | 010C | 6 | Arc | BYPLM.A13L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
22 | BJBHT.A12L6 |  | BLECF_ | 172 | 0x | 00AC | 6 | Arc | BYPLM.A12L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
23 | UD62 |  | BLECF_ | 187 | 0x | 00BB | 8 | UJ63 | BY03 | 6 | CL | 3 | 1 | B | 17-18 | SR6 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Ch4 : BLMDS.9760.B2C22_6.983_DUMP no modulation !!! Not connected in data base!!!!!!
24 | TD62.2 | yes | BLECF_ | 255 | 0x | 00FF | 8 | UJ63 | BY03 | 6 | CL | 2 | 10 | B | 15-16 | SR6 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
25 | TD62.1 | yes | BLECF_ | 511 | 0x | 01FF | 8 | UJ63 | BY03 | 6 | CL | 2 | 9 | B | 13-14 | SR6 | BY05 | C | 4 | 1-2 | BY02 | C | 2 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
26 | BJBHT.A11L6 |  | BLECF_ | 108 | 0x | 006C | 6 | UJ63 | BY03 | 6 | CL | 2 | 8 | B | 11-12 | SR6 | BY05 | C | 3 | 11-12 | BY02 | L | 1 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
27 | BJBHT.A10L6 |  | BLECF_ | 91 | 0x | 005B | 6 | UJ63 | BY03 | 6 | CL | 2 | 7 | B | 9-10 | SR6 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
28 | BJBHT.A9L6 |  | BLECF_ | 161 | 0x | 00A1 | 6 | UJ63 | BY03 | 6 | CL | 2 | 6 | B | 7-8 | SR6 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
29 | BJBHT.A8L6 |  | BLECF_ | 115 | 0x | 0073 | 6 | UJ63 | BY03 | 6 | CL | 2 | 5 | B | 5-6 | SR6 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [8 empty] | NES18 | No
30 | BJBHT.C5L6 |  | BLECF_ | 49 | 0x | 0031 | 1 | UJ63 | BY03 | 6 | CL | 2 | 4 | B | 3-4 | SR6 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [9 empty] | No |  | HV dividere not installed (behind DFB)
31 | BJBHT.B5L6 |  | BLECF_ | 429 | 0x | 01AD | 8 | UJ63 | BY03 | 6 | CL | 2 | 3 | B | 1-2 | SR6 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Ch5: BLMQI.05L6.B2I10_MQY also wit Batterie no signal -BLECF needes to be checked
32 | BJBHT.A5L6 |  | BLECF_ | 107 | 0x | 006B | 3 | UJ63 | BY03 | 6 | CL | 2 | 2 | A | 23-24 | SR6 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
33 | BJBHT.K4L6 |  | BLECF_ | 223 | 0x | 00DF | 8 | UJ63 | BY03 | 6 | CL | 2 | 1 | A | 21-22 | SR6 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
34 | BJBHT.J4L6 | yes | BLECF_ | 548 | 0x | 0224 | 5 | UJ63 | BY03 | 6 | CL | 1 | 10 | A | 19-20 | SR6 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
35 | BJBHT.I4L6 |  | BLECF_ | 508 | 0x | 01FC | 5 | UJ63 | BY03 | 6 | CL | 1 | 9 | A | 17-18 | SR6 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
36 | BJBHT.H4L6 |  | BLECF_ | 336 | 0x | 0150 | 3 | UJ63 | BY03 | 6 | CL | 1 | 8 | A | 15-16 | SR6 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [9 empty] | No
37 | BJBHT.G4L6 |  | BLECF_ | 66 | 0x | 0042 | 3 | UJ63 | BY03 | 6 | CL | 1 | 7 | A | 13-14 | SR6 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [9 empty] | No
38 | BJBHT.F4L6 |  | BLECF_ | 610 | 0x | 0262 | 2 | UJ63 | BY03 | 6 | CL | 1 | 6 | A | 11-12 | SR6 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
39 | BJBHT.E4L6 |  | BLECF_ | 392 | 0x | 0188 | 2 | UJ63 | BY03 | 6 | CL | 1 | 5 | A | 9-10 | SR6 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
40 | BJBHT.D4L6 |  | BLECF_ | 612 | 0x | 0264 | 8 | UJ63 | BY03 | 6 | CL | 1 | 4 | A | 7-8 | SR6 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
41 | BJBHT.C4L6 |  | BLECF_ | 12 | 0x | 000C | 8 | UJ63 | BY03 | 6 | CL | 1 | 3 | A | 5-6 | SR6 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | No
42 | BJBHT.B4L6 |  | BLECF_ | 436 | 0x | 01B4 | 7 | UJ63 | BY03 | 6 | CL | 1 | 2 | A | 3-4 | SR6 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
43 | BJBHT.A4L6 | yes | BLECF_ | 144 | 0x | 0090 | 7 | UJ63 | BY03 | 6 | CL | 1 | 1 | A | 1-2 | SR6 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
44 | BJBHT.A4R6 | yes | BLECF_ | 102 | 0x | 0066 | 7 | UJ67 | BY03 | 6 | CR | 1 | 1 | A | 1-2 | SR6 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Problem with Filter CH5/6/7
45 | BJBHT.B4R6 |  | BLECF_ | 116 | 0x | 0074 | 7 | UJ67 | BY03 | 6 | CR | 1 | 2 | A | 3-4 | SR6 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
46 | BJBHT.C4R6 | yes | BLECF_ | 156 | 0x | 009C | 8 | UJ67 | BY03 | 6 | CR | 1 | 3 | A | 5-6 | SR6 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
47 | BJBHT.D4R6 |  | BLECF_ | 202 | 0x | 00CA | 8 | UJ67 | BY03 | 6 | CR | 1 | 4 | A | 7-8 | SR6 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
48 | BJBHT.E4R6 | yes | BLECF_ | 167 | 0x | 00A7 | 2 | UJ67 | BY03 | 6 | CR | 1 | 5 | A | 9-10 | SR6 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
49 | BJBHT.F4R6 |  | BLECF_ | 545 | 0x | 0221 | 2 | UJ67 | BY03 | 6 | CR | 1 | 6 | A | 11-12 | SR6 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
50 | BJBHT.G4R6 |  | BLECF_ | 219 | 0x | 00DB | 3 | UJ67 | BY03 | 6 | CR | 1 | 7 | A | 13-14 | SR6 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Ch1: BLMDI.3243.B1E10_MKBH CFC needs to be check
51 | BJBHT.H4R6 |  | BLECF_ | 45 | 0x | 002D | 3 | UJ67 | BY03 | 6 | CR | 1 | 8 | A | 15-16 | SR6 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
52 | BJBHT.I4R6 | yes | BLECF_ | 173 | 0x | 00AD | 5 | UJ67 | BY03 | 6 | CR | 1 | 9 | A | 17-18 | SR6 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | CFC Reset
53 | BJBHT.J4R6 | yes | BLECF_ | 262 | 0x | 0106 | 4 | UJ67 | BY03 | 6 | CR | 1 | 10 | A | 19-20 | SR6 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
54 | BJBHT.K4R6 |  | BLECF_ | 198 | 0x | 00C6 | 8 | UJ67 | BY03 | 6 | CR | 2 | 1 | A | 21-22 | SR6 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Ch6: BLMQI.04R6.B1E30_MQY ; also with Batterie no signal -BLECF needes to be checked
55 | BJBHT.A5R6 |  | BLECF_ | 289 | 0x | 0121 | 3 | UJ67 | BY03 | 6 | CR | 2 | 2 | A | 23-24 | SR6 | BY05 | C | 11 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
56 | BJBHT.B5R6 |  | BLECF_ | 165 | 0x | 00A5 | 8 | UJ67 | BY03 | 6 | CR | 2 | 3 | B | 1-2 | SR6 | BY05 | C | 10 | 1-2 | BY02 | R | 3 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Ch7: BLMDI.4162.B1T10_MQY.5R6 CFC needs to be check
57 | BJBHT.C5R6 |  | BLECF_ | 88 | 0x | 0058 | 1 | UJ67 | BY03 | 6 | CR | 2 | 4 | B | 3-4 | SR6 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [9 empty] | No
58 | BJBHT.A8R6 |  | BLECF_ | 159 | 0x | 009F | 6 | UJ67 | BY03 | 6 | CR | 2 | 5 | B | 5-6 | SR6 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
59 | BJBHT.A9R6 |  | BLECF_ | 542 | 0x | 021E | 6 | UJ67 | BY03 | 6 | CR | 2 | 6 | B | 7-8 | SR6 | BY05 | C | 10 | 7-8 | BY02 | R | 3 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
60 | BJBHT.A10R6 |  | BLECF_ | 164 | 0x | 00A4 | 6 | UJ67 | BY03 | 6 | CR | 2 | 7 | B | 9-10 | SR6 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
61 | BJBHT.A11R6 |  | BLECF_ | 844 | 0x | 034C | 6 | UJ67 | BY03 | 6 | CR | 2 | 8 | B | 11-12 | SR6 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
62 | TD68.1 | yes | BLECF_ | 391 | 0x | 0187 | 8 | UJ67 | BY03 | 6 | CR | 2 | 9 | B | 13-14 | SR6 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Problem with Filter CH3/7
63 | TD68.2 | yes | BLECF_ | 56 | 0x | 0038 | 8 | UJ67 | BY03 | 6 | CR | 2 | 10 | B | 15-16 | SR6 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Problem with Filter CH4
64 | UD68 |  | BLECF_ | 480 | 0x | 01E0 | 8 | UJ67 | BY03 | 6 | CR | 3 | 1 | B | 17-18 | SR6 | BY05 | C | 9 | 5-6 | BY02 | C | 2 | 20 | 16 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
65 | BJBHT.A12R6 |  | BLECF_ | 611 | 0x | 0263 | 6 | Arc | BYPLM.A12R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | was 527, new comes from (13r3)
66 | BJBHT.A13R6 |  | BLECF_ | 730 | 0x | 02DA | 6 | Arc | BYPLM.A13R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [9 empty] | No |  | Ch4 : BLMBI.14R6.B0T10_MBA-MBB also with Batterie no signal -BLECF needes to be checked
67 | BJBHT.A14R6 |  | BLECF_ | 532 | 0x | 0214 | 6 | Arc | BYPLM.A14R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
68 | BJBHT.A15R6 |  | BLECF_ | 569 | 0x | 0239 | 6 | Arc | BYPLM.A15R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [9 empty] | No
69 | BJBHT.A16R6 |  | BLECF_ | 274 | 0x | 0112 | 6 | Arc | BYPLM.A16R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
70 | BJBHT.A17R6 |  | BLECF_ | 301 | 0x | 012D | 6 | Arc | BYPLM.A17R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
71 | BJBHT.A18R6 |  | BLECF_ | 19 | 0x | 0013 | 6 | Arc | BYPLM.A18R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
72 | BJBHT.A19R6 |  | BLECF_ | 563 | 0x | 0233 | 6 | Arc | BYPLM.A19R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 0 | [9 empty] | No
73 | BJBHT.A20R6 |  | BLECF_ | 566 | 0x | 0236 | 6 | Arc | BYPLM.A20R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
74 | BJBHT.A21R6 |  | BLECF_ | 295 | 0x | 0127 | 6 | Arc | BYPLM.A21R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | CH1, CH3, CH6: BLMBI.21R6.B0T20_MBA-MBB ,BLMQI.21R6.B1E30_MQ , BLMQI.21R6.B2I30_MQ also with Batterie no signal -BLECF needes to be checked
75 | BJBHT.A22R6 |  | BLECF_ | 242 | 0x | 00F2 | 6 | Arc | BYPLM.A22R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
76 | BJBHT.A23R6 |  | BLECF_ | 540 | 0x | 021C | 6 | Arc | BYPLM.A23R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
77 | BJBHT.A24R6 |  | BLECF_ | 234 | 0x | 00EA | 6 | Arc | BYPLM.A24R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
78 | BJBHT.A25R6 |  | BLECF_ | 555 | 0x | 022B | 6 | Arc | BYPLM.A25R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
79 | BJBHT.A26R6 |  | BLECF_ | 564 | 0x | 0234 | 6 | Arc | BYPLM.A26R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
80 | BJBHT.A27R6 |  | BLECF_ | 557 | 0x | 022D | 6 | Arc | BYPLM.A27R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
81 | BJBHT.A28R6 |  | BLECF_ | 63 | 0x | 003F | 6 | Arc | BYPLM.A28R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
82 | BJBHT.A29R6 |  | BLECF_ | 561 | 0x | 0231 | 6 | Arc | BYPLM.A29R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
83 | BJBHT.A30R6 |  | BLECF_ | 565 | 0x | 0235 | 6 | Arc | BYPLM.A30R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No |  | Ch4: BLMBI.31R6.B0T10_MBB-MBA; also with Batterie no signal -BLECF needes to be checked
84 | BJBHT.A31R6 |  | BLECF_ | 210 | 0x | 00D2 | 6 | Arc | BYPLM.A31R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
85 | BJBHT.A32R6 |  | BLECF_ | 154 | 0x | 009A | 6 | Arc | BYPLM.A32R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
86 | BJBHT.A33R6 |  | BLECF_ | 538 | 0x | 021A | 6 | Arc | BYPLM.A33R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
87 | BJBHT.A34R6 |  | BLECF_ | 176 | 0x | 00B0 | 6 | Arc | BYPLM.A34R6 | 6 | R |  |  |  | 7-9 | SR6 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | No
| Door | [47 empty] | 1600246 | UJ63 | BJBAP.C4L6=RB64
|  |  | BLECF_ | 176 |  | CFC card repaired after HV issue | [20 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k | [11 empty] | 1600258 | UJ63 | BJBAP.E4L6=RB64
[27 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current) | [11 empty] | 1600760 | UJ63 | BJBAP.A8L6=R623
[7 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing | [20 empty] | 1600761 | UJ63 | BJBAP.A10L6=R623
[7 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[7 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[27 empty] | O.K. | No problems observed
[46 empty] | R6 | BLMBI.14R6.B0T10_MBA-MBB | BLMBI.14R6.B0T10_MBA-MBB:LOSS_RS09
[46 empty] | R6 | BLMBI.21R6.B0T20_MBA-MBB | BLMBI.21R6.B0T20_MBA-MBB:LOSS_RS09
[46 empty] | R6 | BLMBI.31R6.B0T10_MBB-MBA | BLMBI.31R6.B0T10_MBB-MBA:LOSS_RS09
[46 empty] | R6 | BLMQI.04R6.B1E30_MQY | BLMQI.04R6.B1E30_MQY:LOSS_RS09
[46 empty] | R6 | BLMQI.21R6.B1E30_MQ | BLMQI.21R6.B1E30_MQ:LOSS_RS09
[46 empty] | R6 | BLMQI.21R6.B2I30_MQ | BLMQI.21R6.B2I30_MQ:LOSS_RS09
[46 empty] | L6 | BLMQI.05L6.B2I10_MQY | BLMQI.05L6.B2I10_MQY:LOSS_RS09
[46 empty] | L6 | BLMQI.14L6.B1E30_MQ | BLMQI.14L6.B1E30_MQ:LOSS_RS09
[46 empty] | 43 | BLMDI.3243.B1E10_MKBH | BLMDI.3243.B1E10_MKBH:LOSS_RS09
[46 empty] | 60 | BLMDS.9760.B2C22_6.983_DUMP | BLMDS.9760.B2C22_6.983_DUMP:LOSS_RS09
[46 empty] | 62 | BLMDI.4162.B1T10_MQY.5R6 | BLMDI.4162.B1T10_MQY.5R6:LOSS_RS09
Note on E22: effinger: BLECF_0279 repalced by BLECF_0708 at the 5.6.08 by Ion
Note on AV28: Ewald Effinger: Problem with HV (20mA)!!!! After checking the powersupply the 20mA disapeard! The cable of the first HV suppl was not connected. Filter Capacitor of the HV box in the Dump was burned and therefore short circuited
Note on C68: Ewald Effinger: All Channels 1uF/100k
Note on C69: Ewald Effinger: All Channels 1uF/100k
Note on C29: Ewald Effinger: All Channels 1uF/100k
Note on C30: Ewald Effinger: All Channels 1uF/100k
Note on AV58: Ewald Effinger: Problem with Filter CH2
Note on AV39: Ewald Effinger: Channel 3 blocked Check cabling (Card needed to be unplugged to perform a proper reset)
Note on AV40: Ewald Effinger: CFC Reset (with switch was o.k)
Note on E90: Ewald Effinger: BLECF_0565 ch4 and ch8 repaired since the input was burnded -> While checking

## Point 7
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES-18 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L7 | BLECF_ | 69 | 0x | 0045 | 6 | Arc | BYPLM.A33L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | [17 empty] | NO
2 | BJBHT.A32L7 | BLECF_ | 122 | 0x | 007A | 6 | Arc | BYPLM.A32L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | [17 empty] | NO |  | CH 6 no modulation : BLMQI.32L7.B1E30_MQ (was connected on ch7)
3 | BJBHT.A31L7 | BLECF_ | 426 | 0x | 01AA | 6 | Arc | BYPLM.A31L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | [17 empty] | NO
4 | BJBHT.A30L7 | BLECF_ | 580 | 0x | 0244 | 6 | Arc | BYPLM.A30L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | [17 empty] | NO |  | Power supply ? Reset! --> Transformer broken -> TR. Exchanged o.k!!!
5 | BJBHT.A29L7 | BLECF_ | 123 | 0x | 007B | 6 | Arc | BYPLM.A29L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | [17 empty] | NO
6 | BJBHT.A28L7 | BLECF_ | 20 | 0x | 0014 | 6 | Arc | BYPLM.A28L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | [17 empty] | NO
7 | BJBHT.A27L7 | BLECF_ | 140 | 0x | 008C | 6 | Arc | BYPLM.A27L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | [17 empty] | NO
8 | BJBHT.A26L7 | BLECF_ | 126 | 0x | 007E | 6 | Arc | BYPLM.A26L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | [17 empty] | NO
9 | BJBHT.A25L7 | BLECF_ | 43 | 0x | 002B | 6 | Arc | BYPLM.A25L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | [17 empty] | NO
10 | BJBHT.A24L7 | BLECF_ | 454 | 0x | 01C6 | 6 | Arc | BYPLM.A24L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | [17 empty] | NO
11 | BJBHT.A23L7 | BLECF_ | 17 | 0x | 0011 | 6 | Arc | BYPLM.A23L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | [17 empty] | NO |  | Fiber connection to be checked --> o.k.! (LF goes with the CFC channel-> CFC needs to be checked) -> GOH repalced on the CFC
12 | BJBHT.A22L7 | BLECF_ | 47 | 0x | 002F | 6 | Arc | BYPLM.A22L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | [17 empty] | NO
13 | BJBHT.A21L7 | BLECF_ | 87 | 0x | 0057 | 6 | Arc | BYPLM.A21L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | [17 empty] | NO
14 | BJBHT.A20L7 | BLECF_ | 74 | 0x | 004A | 6 | Arc | BYPLM.A20L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | [17 empty] | NO
15 | BJBHT.A19L7 | BLECF_ | 466 | 0x | 01D2 | 6 | Arc | BYPLM.A19L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | [17 empty] | NO
16 | BJBHT.A18L7 | BLECF_ | 208 | 0x | 00D0 | 6 | Arc | BYPLM.A18L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | [7 empty] | FF | [9 empty] | NO |  | Reset and channel 7 to be checked if DAC value is correct , same after reset ---> CFC needs to be checked ---> spare card installed BLECF_0208
17 | BJBHT.A17L7 | BLECF_ | 334 | 0x | 014E | 6 | Arc | BYPLM.A17L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | [17 empty] | NO
18 | BJBHT.A16L7 | BLECF_ | 23 | 0x | 0017 | 6 | Arc | BYPLM.A16L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | [17 empty] | NO
19 | BJBHT.A15L7 | BLECF_ | 534 | 0x | 0216 | 6 | Arc | BYPLM.A15L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | [17 empty] | NO |  | Reset needed , same after reset ---> CFC needs to be checked --> CFC checked and o.k. | chris saw problems with TC
20 | BJBHT.A14L7 | BLECF_ | 461 | 0x | 01CD | 6 | Arc | BYPLM.A14L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | [17 empty] | NO |  | Reset needed , same after reset ---> CFC needs to be checked --> CFC checked and o.k.
21 | BJBHT.A13L7 | BLECF_ | 449 | 0x | 01C1 | 6 | Arc | BYPLM.A13L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | [17 empty] | NO
22 | BJBHT.A12L7 | BLECF_ | 472 | 0x | 01D8 | 6 | Arc | BYPLM.A12L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | [17 empty] | NO
23 | BJBHT.B11L7 | BLECF_ | 469 | 0x | 01D5 | 8 | RR73 | BY01 | 7 | CL | 3 | 8 | F | 1-2 | SR7 | BY05 | C | 5 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | [17 empty] | NO
24 | BJBHT.A11L7 | BLECF_ | 679 | 0x | 02A7 | 8 | RR73 | BY01 | 7 | CL | 3 | 7 | E | 5-6 | SR7 | BY05 | C | 5 | 5-6 | BY02 | E | 4 | 5 | 2 | 9-16 | [17 empty] | NO
25 | BJBHT.A10L7 | BLECF_ | 544 | 0x | 0220 | 8 | RR73 | BY01 | 7 | CL | 3 | 6 | E | 3-4 | SR7 | BY05 | C | 5 | 3-4 | BY02 | L | 1 | 7 | 4 | 1-8 | [17 empty] | NO |  | Reset needed
26 | BJBHT.B9L7 | BLECF_ | 678 | 0x | 02A6 | 8 | RR73 | BY01 | 7 | CL | 3 | 5 | E | 1-2 | SR7 | BY05 | C | 5 | 1-2 | BY02 | L | 1 | 6 | 3 | 9-16 | [7 empty] | 00 | [9 empty] | NO
27 | BJBHT.A9L7 | BLECF_ | 559 | 0x | 022F | 8 | RR73 | BY01 | 7 | CL | 3 | 4 | D | 23-24 | SR7 | BY05 | C | 4 | 11-12 | BY02 | E | 4 | 5 | 2 | 1-8 | [17 empty] | NO |  | 1. Reset ->verification 2. Replacement of the CFC ---> CFC needs to be checked! --> card checked o.k. | Surface TC
28 | BJBHT.B8L7 | BLECF_ | 649 | 0x | 0289 | 6 | RR73 | BY01 | 7 | CL | 3 | 3 | D | 21-22 | SR7 | BY05 | C | 4 | 9-10 | BY02 | L | 1 | 6 | 3 | 1-8 | [17 empty] | NO
29 | BJBHT.A8L7 | BLECF_ | 462 | 0x | 01CE | 3 | RR73 | BY01 | 7 | CL | 3 | 2 | D | 19-20 | SR7 | BY05 | C | 4 | 7-8 | BY02 | E | 4 | 4 | 1 | 9-16 | [7 empty] | 00 | [9 empty] | NO
30 | BJBHT.B7L7 | BLECF_ | 513 | 0x | 0201 | 7 | RR73 | BY01 | 7 | CL | 3 | 1 | D | 17-18 | SR7 | BY05 | C | 4 | 5-6 | BY02 | L | 1 | 5 | 2 | 9-16 | [17 empty] | NO
31 | BJBHT.A7L7 | BLECF_ | 411 | 0x | 019B | 2 | RR73 | BY01 | 7 | CL | 2 | 10 | D | 15-16 | SR7 | BY05 | C | 4 | 3-4 | BY02 | E | 4 | 4 | 1 | 1-8 | [17 empty] | NO
32 | BJBHT.I6L7 | BLECF_ | 587 | 0x | 024B | 7 | RR73 | BY01 | 7 | CL | 2 | 9 | D | 13-14 | SR7 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 5 | 2 | 1-8 | [17 empty] | NO
33 | BJBHT.H6L7 | BLECF_ | 118 | 0x | 0076 | 4 | RR73 | BY01 | 7 | CL | 2 | 8 | D | 11-12 | SR7 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 11 | 8 | 9-16 | [17 empty] | NO
34 | BJBHT.G6L7 | BLECF_ | 191 | 0x | 00BF | 4 | RR73 | BY01 | 7 | CL | 2 | 7 | D | 9-10 | SR7 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 11 | 8 | 1-8 | [17 empty] | NO
35 | BJBHT.F6L7 | BLECF_ | 434 | 0x | 01B2 | 7 | RR73 | BY01 | 7 | CL | 2 | 6 | D | 7-8 | SR7 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 10 | 7 | 9-16 | [17 empty] | NO
36 | BJBHT.E6L7 | BLECF_ | 638 | 0x | 027E | 7 | RR73 | BY01 | 7 | CL | 2 | 5 | D | 5-6 | SR7 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 10 | 7 | 1-8 | [17 empty] | NO
37 | BJBHT.D6L7 | BLECF_ | 364 | 0x | 016C | 3 | RR73 | BY01 | 7 | CL | 2 | 4 | D | 3-4 | SR7 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 9 | 6 | 9-16 | [17 empty] | NO
38 | BJBHT.C6L7 | BLECF_ | 360 | 0x | 0168 | 3 | RR73 | BY01 | 7 | CL | 2 | 3 | D | 1-2 | SR7 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 9 | 6 | 1-8 | [17 empty] | NO
39 | BJBHT.B6L7 | BLECF_ | 682 | 0x | 02AA | 7 | RR73 | BY01 | 7 | CL | 2 | 2 | C | 23-24 | SR7 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 8 | 5 | 9-16 | [7 empty] | 00 | [9 empty] | NO
40 | BJBHT.A6L7 | BLECF_ | 265 | 0x | 0109 | 7 | RR73 | BY01 | 7 | CL | 2 | 1 | C | 21-22 | SR7 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8 | [17 empty] | NO
41 | BJBHT.E5L7 | BLECF_ | 96 | 0x | 0060 | 6 | RR73 | BY01 | 7 | CL | 1 | 10 | C | 19-20 | SR7 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 4 | 1 | 9-16 | [16 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
42 | BJBHT.D5L7 | BLECF_ | 94 | 0x | 005E | 6 | RR73 | BY01 | 7 | CL | 1 | 9 | C | 17-18 | SR7 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 7 | 4 | 9-16 | [7 empty] | 00 | [9 empty] | NO
43 | BJBHT.C5L7 | BLECF_ | 273 | 0x | 0111 | 6 | RR73 | BY01 | 7 | CL | 1 | 8 | C | 15-16 | SR7 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8 | [16 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
44 | BJBHT.B5L7 | BLECF_ | 590 | 0x | 024E | 8 | RR73 | BY01 | 7 | CL | 1 | 7 | C | 13-14 | SR7 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 6 | 3 | 9-16 | [17 empty] | NO
45 | BJBHT.A5L7 | BLECF_ | 244 | 0x | 00F4 | 8 | RR73 | BY01 | 7 | CL | 1 | 6 | C | 11-12 | SR7 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 1-8 | [17 empty] | NO
46 | BJBHT.E4L7 | BLECF_ | 372 | 0x | 0174 | 6 | RR73 | BY01 | 7 | CL | 1 | 5 | C | 9-10 | SR7 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 4 | 1 | 1-8 | [17 empty] | NO
47 | BJBHT.D4L7 | BLECF_ | 445 | 0x | 01BD | 6 | RR73 | BY01 | 7 | CL | 1 | 4 | C | 7-8 | SR7 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | [17 empty] | NO
48 | BJBHT.C4L7 | BLECF_ | 369 | 0x | 0171 | 6 | RR73 | BY01 | 7 | CL | 1 | 3 | C | 5-6 | SR7 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | [17 empty] | NO
49 | BJBHT.B4L7 | BLECF_ | 633 | 0x | 0279 | 6 | RR73 | BY01 | 7 | CL | 1 | 2 | C | 3-4 | SR7 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | [17 empty] | NO
50 | BJBHT.A4L7 | BLECF_ | 308 | 0x | 0134 | 6 | RR73 | BY01 | 7 | CL | 1 | 1 | C | 1-2 | SR7 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | [17 empty] | NO
51 | BJBHT.A4R7 | BLECF_ | 443 | 0x | 01BB | 6 | RR77 | BY02 | 7 | CR | 1 | 1 | C | 1-2 | SR7 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | [7 empty] | 0 | [9 empty] | NO
52 | BJBHT.B4R7 | BLECF_ | 455 | 0x | 01C7 | 6 | RR77 | BY02 | 7 | CR | 1 | 2 | C | 3-4 | SR7 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | [7 empty] | 0 | [9 empty] | NO
53 | BJBHT.C4R7 | BLECF_ | 351 | 0x | 015F | 6 | RR77 | BY02 | 7 | CR | 1 | 3 | C | 5-6 | SR7 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | [7 empty] | 0 | [9 empty] | NO
54 | BJBHT.D4R7 | BLECF_ | 431 | 0x | 01AF | 6 | RR77 | BY02 | 7 | CR | 1 | 4 | C | 7-8 | SR7 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | [17 empty] | NO
55 | BJBHT.E4R7 | BLECF_ | 420 | 0x | 01A4 | 6 | RR77 | BY02 | 7 | CR | 1 | 5 | C | 9-10 | SR7 | BY05 | C | 12 | 9-10 | BY02 | R | 3 | 4 | 1 | 1-8 | [17 empty] | NO
56 | BJBHT.A5R7 | BLECF_ | 713 | 0x | 02C9 | 8 | RR77 | BY02 | 7 | CR | 1 | 6 | C | 11-12 | SR7 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 1-8 | [17 empty] | NO |  | 2c9
57 | BJBHT.B5R7 | BLECF_ | 446 | 0x | 01BE | 8 | RR77 | BY02 | 7 | CR | 1 | 7 | C | 13-14 | SR7 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 15 | 11 | 9-16 | [17 empty] | NO
58 | BJBHT.C5R7 | BLECF_ | 421 | 0x | 01A5 | 6 | RR77 | BY02 | 7 | CR | 1 | 8 | C | 15-16 | SR7 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | [7 empty] | 0 | [8 empty] | NES18 | NO
59 | BJBHT.D5R7 | BLECF_ | 415 | 0x | 019F | 6 | RR77 | BY02 | 7 | CR | 1 | 9 | C | 17-18 | SR7 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 16 | 12 | 9-16 | [17 empty] | NO
60 | BJBHT.E5R7 | BLECF_ | 368 | 0x | 0170 | 6 | RR77 | BY02 | 7 | CR | 1 | 10 | C | 19-20 | SR7 | BY05 | C | 11 | 7-8 | BY02 | R | 3 | 4 | 1 | 9-16 | [17 empty] | NO
61 | BJBHT.A6R7 | BLECF_ | 192 | 0x | 00C0 | 4 | RR77 | BY02 | 7 | CR | 2 | 1 | C | 21-22 | SR7 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 17 | 13 | 1-8 | [7 empty] | 0 | [9 empty] | NO |  | cd
62 | BJBHT.B6R7 | BLECF_ | 221 | 0x | 00DD | 4 | RR77 | BY02 | 7 | CR | 2 | 2 | C | 23-24 | SR7 | BY05 | C | 11 | 11-12 | BY02 | C | 2 | 17 | 13 | 9-16 | [17 empty] | NO |  | dd
63 | BJBHT.C6R7 | BLECF_ | 228 | 0x | 00E4 | 5 | RR77 | BY02 | 7 | CR | 2 | 3 | D | 1-2 | SR7 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 18 | 14 | 1-8 | [17 empty] | NO |  | (e4) HV cable needed jumper on the place where BLM was removed!
64 | BJBHT.D6R7 | BLECF_ | 245 | 0x | 00F5 | 5 | RR77 | BY02 | 7 | CR | 2 | 4 | D | 3-4 | SR7 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 18 | 14 | 9-16 | [17 empty] | NO |  | f5
65 | BJBHT.E6R7 | BLECF_ | 31 | 0x | 001F | 8 | RR77 | BY02 | 7 | CR | 2 | 5 | D | 5-6 | SR7 | BY05 | C | 10 | 5-6 | BY02 | C | 2 | 19 | 15 | 1-8 | [17 empty] | NO |  | 1f
66 | BJBHT.F6R7 | BLECF_ | 13 | 0x | 000D | 8 | RR77 | BY02 | 7 | CR | 2 | 6 | D | 7-8 | SR7 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 19 | 15 | 9-16 | [17 empty] | NO |  | CH 1 and CH2 no modulation: BLMTL.06R7.B1E10_TCLA.B6R7.B1 and BLMEL.06R7.B2I10_TCHSS.6R7.B2
67 | BJBHT.G6R7 | BLECF_ | 433 | 0x | 01B1 | 4 | RR77 | BY02 | 7 | CR | 2 | 7 | D | 9-10 | SR7 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 20 | 16 | 1-8 | [17 empty] | NO
68 | BJBHT.H6R7 | BLECF_ | 482 | 0x | 01E2 | 4 | RR77 | BY02 | 7 | CR | 2 | 8 | D | 11-12 | SR7 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 20 | 16 | 9-16 | [7 empty] | 0 | [9 empty] | NO
69 | BJBHT.I6R7 | BLECF_ | 464 | 0x | 01D0 | 7 | RR77 | BY02 | 7 | CR | 2 | 9 | D | 13-14 | SR7 | BY05 | C | 9 | 1-2 | BY02 | R | 3 | 5 | 2 | 1-8 | [17 empty] | NO
70 | BJBHT.A7R7 | BLECF_ | 359 | 0x | 0167 | 2 | RR77 | BY02 | 7 | CR | 2 | 10 | D | 15-16 | SR7 | BY05 | C | 9 | 3-4 | BY02 | E | 4 | 13 | 9 | 1-8 | [17 empty] | NO |  | BNC cable between HV divider and signal box had a shortcut!
71 | BJBHT.B7R7 | BLECF_ | 470 | 0x | 01D6 | 7 | RR77 | BY02 | 7 | CR | 3 | 1 | D | 17-18 | SR7 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 5 | 2 | 9-16 | [17 empty] | NO |  | Ch 7 and CH8 no modulation BLMTI.07R7.B1E10_TCLA.A7R7.B1 and BLMEI.07R7.B1E10_TCLA.B7R7.B1
72 | BJBHT.A8R7 | BLECF_ | 700 | 0x | 02BC | 3 | RR77 | BY02 | 7 | CR | 3 | 2 | D | 19-20 | SR7 | BY05 | C | 9 | 7-8 | BY02 | E | 4 | 13 | 9 | 9-16 | [17 empty] | NO
73 | BJBHT.B8R7 | BLECF_ | 358 | 0x | 0166 | 6 | RR77 | BY02 | 7 | CR | 3 | 3 | D | 21-22 | SR7 | BY05 | C | 9 | 9-10 | BY02 | R | 3 | 6 | 3 | 1-8 | [17 empty] | NO |  | 17d
74 | BJBHT.A9R7 | BLECF_ | 505 | 0x | 01F9 | 8 | RR77 | BY02 | 7 | CR | 3 | 4 | D | 23-24 | SR7 | BY05 | C | 9 | 11-12 | BY02 | E | 4 | 14 | 10 | 1-8 | [17 empty] | NO
75 | BJBHT.B9R7 | BLECF_ | 345 | 0x | 0159 | 8 | RR77 | BY02 | 7 | CR | 3 | 5 | E | 1-2 | SR7 | BY05 | C | 8 | 1-2 | BY02 | R | 3 | 6 | 3 | 9-16 | [17 empty] | NO |  | 1b9 (Ewald: HV cable problem bad contact afterwards noise was gone)
76 | BJBHT.A10R7 | BLECF_ | 381 | 0x | 017D | 8 | RR77 | BY02 | 7 | CR | 3 | 6 | E | 3-4 | SR7 | BY05 | C | 8 | 3-4 | BY02 | R | 3 | 7 | 4 | 1-8 | [17 empty] | NO |  | 166
77 | BJBHT.A11R7 | BLECF_ | 428 | 0x | 01AC | 8 | RR77 | BY02 | 7 | CR | 3 | 7 | E | 5-6 | SR7 | BY05 | C | 8 | 5-6 | BY02 | E | 4 | 14 | 10 | 9-16 | [17 empty] | NO |  | broken fiber surface
78 | BJBHT.B11R7 | BLECF_ | 441 | 0x | 01B9 | 7 | RR77 | BY02 | 7 | CR | 3 | 8 | F | 1-2 | SR7 | BY05 | C | 7 | 1-2 | BY02 | R | 3 | 7 | 4 | 9-16 | [7 empty] | 159 | [8 empty] | NES18 | NO |  | Reset needed
79 | BJBHT.A12R7 | BLECF_ | 283 | 0x | 011B | 6 | Arc | BYPLM.A12R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | [7 empty] | 11b | [9 empty] | NO |  | DAC Reset with HV not executed ( reset of card needed and need to be checked again)
80 | BJBHT.A13R7 | BLECF_ | 716 | 0x | 02CC | 6 | Arc | BYPLM.A13R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | [17 empty] | NO
81 | BJBHT.A14R7 | BLECF_ | 44 | 0x | 002C | 6 | Arc | BYPLM.A14R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | [17 empty] | NO
82 | BJBHT.A15R7 | BLECF_ | 264 | 0x | 0108 | 6 | Arc | BYPLM.A15R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | [17 empty] | NO |  | reset needed 4 channels blocked
83 | BJBHT.A16R7 | BLECF_ | 413 | 0x | 019D | 6 | Arc | BYPLM.A16R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | [17 empty] | NO
84 | BJBHT.A17R7 | BLECF_ | 467 | 0x | 01D3 | 6 | Arc | BYPLM.A17R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | [17 empty] | NO |  | DAC Reset with HV not executed ( reset of card needed and need to be checked again)
85 | BJBHT.A18R7 | BLECF_ | 263 | 0x | 0107 | 6 | Arc | BYPLM.A18R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | [17 empty] | NO
86 | BJBHT.A19R7 | BLECF_ | 291 | 0x | 0123 | 6 | Arc | BYPLM.A19R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | [17 empty] | NO
87 | BJBHT.A20R7 | BLECF_ | 236 | 0x | 00EC | 6 | Arc | BYPLM.A20R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | [17 empty] | NO |  | DAC Reset with HV not executed ( reset of card needed and need to be checked again)
88 | BJBHT.A21R7 | BLECF_ | 158 | 0x | 009E | 6 | Arc | BYPLM.A21R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | [17 empty] | NO
89 | BJBHT.A22R7 | BLECF_ | 256 | 0x | 0100 | 6 | Arc | BYPLM.A22R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | [17 empty] | NO
90 | BJBHT.A23R7 | BLECF_ | 65 | 0x | 0041 | 6 | Arc | BYPLM.A23R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | [17 empty] | NO
91 | BJBHT.A24R7 | BLECF_ | 152 | 0x | 0098 | 6 | Arc | BYPLM.A24R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | [17 empty] | NO
92 | BJBHT.A25R7 | BLECF_ | 220 | 0x | 00DC | 6 | Arc | BYPLM.A25R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | [17 empty] | NO
93 | BJBHT.A26R7 | BLECF_ | 133 | 0x | 0085 | 6 | Arc | BYPLM.A26R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | [17 empty] | NO
94 | BJBHT.A27R7 | BLECF_ | 138 | 0x | 008A | 6 | Arc | BYPLM.A27R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | [17 empty] | NO
95 | BJBHT.A28R7 | BLECF_ | 238 | 0x | 00EE | 6 | Arc | BYPLM.A28R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | [17 empty] | NO
96 | BJBHT.A29R7 | BLECF_ | 216 | 0x | 00D8 | 6 | Arc | BYPLM.A29R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | [17 empty] | NO
97 | BJBHT.A30R7 | BLECF_ | 174 | 0x | 00AE | 6 | Arc | BYPLM.A30R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | [17 empty] | NO
98 | BJBHT.A31R7 | BLECF_ | 132 | 0x | 0084 | 6 | Arc | BYPLM.A31R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | [17 empty] | NO
99 | BJBHT.A32R7 | BLECF_ | 397 | 0x | 018D | 6 | Arc | BYPLM.A32R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | [17 empty] | NO
100 | BJBHT.A33R7 | BLECF_ | 2 | 0x | 0002 | 6 | Arc | BYPLM.A33R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | [17 empty] | NO
101 | BJBHT.A34R7 | BLECF_ | 486 | 0x | 01E6 | 6 | Arc | BYPLM.A34R7 | 7 | R |  |  |  | 7-9 | SR7 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | [17 empty] | NO |  | reset of card needed --> call for verification --> 220V main socket was dismounted --> working since 17.2.15 10:38 (Joel installed the socket)
| Door | [44 empty] | 1701206 | RR73 | BJBAP.C5L7=R74
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k | [9 empty] | 1701207 | RR73 | BJBAP.E5L7=R74
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current) | [9 empty] | 1701208 | RR77 | BJBAP.C5R7=R771
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing | [18 empty] | 1701209 | RR77 | BJBAP.B11R7=R78
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed | [10 empty] | 1714366 | RR77 | BJBAP.C9R7=R78
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
[44 empty] | R7 | BLMEI.07R7.B1E10_TCLA.B7R7.B1 | BLMEI.07R7.B1E10_TCLA.B7R7.B1:LOSS_RS09
[44 empty] | R7 | BLMEL.06R7.B2I10_TCHSS.6R7.B2 | BLMEL.06R7.B2I10_TCHSS.6R7.B2:LOSS_RS09
[44 empty] | R7 | BLMTI.07R7.B1E10_TCLA.A7R7.B1 | BLMTI.07R7.B1E10_TCLA.A7R7.B1:LOSS_RS09
[44 empty] | R7 | BLMTL.06R7.B1E10_TCLA.B6R7.B1 | BLMTL.06R7.B1E10_TCLA.B6R7.B1:LOSS_RS09
[44 empty] | L7 | BLMQI.32L7.B1E30_MQ | BLMQI.32L7.B1E30_MQ:LOSS_RS09
[44 empty] | L7 | BLMQI.32L7.B1E30_MQ | BLMQI.32L7.B1E30_MQ:LOSS_RS09
[43 empty] | BJBAP.B7R7_8 | R7 | BLMEI.07R7.B1E10_TCLA.B7R7.B1 | BLMEI.07R7.B1E10_TCLA.B7R7.B1:LOSS_RS09
[43 empty] | BJSAP.F6R7_2 | R7 | BLMEL.06R7.B2I10_TCHSS.6R7.B2 | BLMEL.06R7.B2I10_TCHSS.6R7.B2:LOSS_RS09
[43 empty] | BJBAP.B7R7_7 | R7 | BLMTI.07R7.B1E10_TCLA.A7R7.B1 | BLMTI.07R7.B1E10_TCLA.A7R7.B1:LOSS_RS09
[43 empty] | BJSAP.F6R7_1 | R7 | BLMTL.06R7.B1E10_TCLA.B6R7.B1 | BLMTL.06R7.B1E10_TCLA.B6R7.B1:LOSS_RS09
Note on D62: Ewald Effinger: was 412
Note on D67: Ewald Effinger: was 334
Note on D68: Ewald Effinger: was 031
Note on D69: Ewald Effinger: was 350
Note on D70: Ewald Effinger: was 081
Note on D71: Ewald Effinger: was 394
Note on AM74: Ewald Effinger: level was on, no o.k
Note on AN74: Ewald Effinger: level was on, no o.k
Note on AJ57: Ewald Effinger: evel was on, no o.k
Note on AK58: Ewald Effinger: evel was on, no o.k
Note on AL58: Ewald Effinger: evel was on, no o.k
Note on AL59: Ewald Effinger: evel was on, no o.k
Note on AN64: Ewald Effinger: evel was on, no o.k
Note on AI67: Ewald Effinger: evel was on, no o.k
Note on AJ67: Ewald Effinger: evel was on, no o.k
Note on AL67: Ewald Effinger: evel was on, no o.k
Note on AN67: Ewald Effinger: evel was on, no o.k
Note on AS69: Ewald Effinger: Problem with HV
Note on AS72: Ewald Effinger: HV connector was bad connected!
Note on AD83: Ewald Effinger: LF
Note on AS77: Ewald Effinger: Problem with HV
Note on AS81: Ewald Effinger: Problem with HV
Note on AD72: Ewald Effinger: LF
Note on AI81: Ewald Effinger: Noise on all channels)
Note on AC47: Ewald Effinger: LF
Note on AD47: Ewald Effinger: LF
Note on AT84: Ewald Effinger: NES 18 ! To be checked if it is connected!
Note on AT72: Ewald Effinger: Fiber connection to be checked! (Ewald: Adrien cleaned the fibers and checked if the GOH was proper connected)
Note on AT77: Ewald Effinger: HV not o.k. (Ewald: HV Cable was wrong connected
Note on D20: Ewald Effinger: BLECF_0571 removed since ch7 had an issue. New number BLECF_0208
Note on AO20: Ewald Effinger: BLECF_0571, CFC-error was on , Level was o.k channel was 0
Note on AP23: Ewald Effinger: Channel is o.k, Level is off , CFC error on ---> TC needs to be checked
Note on AP24: Ewald Effinger: Channel is o.k, Level is off , CFC error on ---> TC needs to be checked

## Point 8
| Tunnel - position | [14 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | NES18 | 1024 | HV | Comments
Pos. | position |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L8 |  | BLECF_ | 229 | 0x | 00E5 | 6 | Arc | BYPLM.A33L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
2 | BJBHT.A32L8 |  | BLECF_ | 237 | 0x | 00ED | 6 | Arc | BYPLM.A32L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
3 | BJBHT.A31L8 |  | BLECF_ | 292 | 0x | 0124 | 6 | Arc | BYPLM.A31L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
4 | BJBHT.A30L8 |  | BLECF_ | 257 | 0x | 0101 | 6 | Arc | BYPLM.A30L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
5 | BJBHT.A29L8 |  | BLECF_ | 215 | 0x | 00D7 | 6 | Arc | BYPLM.A29L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
6 | BJBHT.A28L8 |  | BLECF_ | 213 | 0x | 00D5 | 6 | Arc | BYPLM.A28L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
7 | BJBHT.A27L8 |  | BLECF_ | 275 | 0x | 0113 | 6 | Arc | BYPLM.A27L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
8 | BJBHT.A26L8 |  | BLECF_ | 277 | 0x | 0115 | 6 | Arc | BYPLM.A26L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
9 | BJBHT.A25L8 |  | BLECF_ | 247 | 0x | 00F7 | 6 | Arc | BYPLM.A25L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
10 | BJBHT.A24L8 |  | BLECF_ | 137 | 0x | 0089 | 6 | Arc | BYPLM.A24L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
11 | BJBHT.A23L8 |  | BLECF_ | 145 | 0x | 0091 | 6 | Arc | BYPLM.A23L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
12 | BJBHT.A22L8 |  | BLECF_ | 129 | 0x | 0081 | 6 | Arc | BYPLM.A22L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
13 | BJBHT.A21L8 |  | BLECF_ | 112 | 0x | 0070 | 6 | Arc | BYPLM.A21L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
14 | BJBHT.A20L8 |  | BLECF_ | 100 | 0x | 0064 | 6 | Arc | BYPLM.A20L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
15 | BJBHT.A19L8 |  | BLECF_ | 114 | 0x | 0072 | 6 | Arc | BYPLM.A19L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
16 | BJBHT.A18L8 |  | BLECF_ | 293 | 0x | 0125 | 6 | Arc | BYPLM.A18L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO |  | Counter reset 3.12.14 (to be checked after a view days)
17 | BJBHT.A17L8 |  | BLECF_ | 150 | 0x | 0096 | 6 | Arc | BYPLM.A17L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
18 | BJBHT.A16L8 |  | BLECF_ | 130 | 0x | 0082 | 6 | Arc | BYPLM.A16L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
19 | BJBHT.A15L8 |  | BLECF_ | 101 | 0x | 0065 | 6 | Arc | BYPLM.A15L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
20 | BJBHT.A14L8 |  | BLECF_ | 105 | 0x | 0069 | 6 | Arc | BYPLM.A14L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
21 | BJBHT.A13L8 |  | BLECF_ | 99 | 0x | 0063 | 6 | Arc | BYPLM.A13L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
22 | BJBHT.A12L8 |  | BLECF_ | 712 | 0x | 02C8 | 6 | Arc | BYPLM.A12L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO |  | reset needed
23 | BJBHT.C11L8 |  | BLECF_ | 95 | 0x | 005F | 6 | UA83 | BY05 | 8 | CL | 2 | 9 | C | 13-14 | SR8 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
24 | BJBHT.B11L8 |  | BLECF_ | 92 | 0x | 005C | 6 | UA83 | BY05 | 8 | CL | 2 | 8 | C | 11-12 | SR8 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
25 | BJBHT.A11L8 | yes | BLECF_ | 86 | 0x | 0056 | 6 | UA83 | BY05 | 8 | CL | 2 | 7 | C | 9-10 | SR8 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 8 | 5 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
26 | BJBHT.A10L8 |  | BLECF_ | 67 | 0x | 0043 | 6 | UA83 | BY05 | 8 | CL | 2 | 6 | C | 7-8 | SR8 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
27 | BJBHT.A9L8 |  | BLECF_ | 33 | 0x | 0021 | 6 | UA83 | BY05 | 8 | CL | 2 | 5 | C | 5-6 | SR8 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
28 | BJBHT.C8L8 |  | BLECF_ | 224 | 0x | 00E0 | 6 | UA83 | BY05 | 8 | CL | 2 | 4 | C | 3-4 | SR8 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
29 | BJBHT.B8L8 |  | BLECF_ | 205 | 0x | 00CD | 6 | UA83 | BY05 | 8 | CL | 2 | 3 | C | 1-2 | SR8 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
30 | BJBHT.A8L8 | yes | BLECF_ | 151 | 0x | 0097 | 6 | UA83 | BY05 | 8 | CL | 2 | 2 | B | 23-24 | SR8 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 7 | 4 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
31 | BJBHT.A7L8 |  | BLECF_ | 743 | 0x | 02E7 | 6 | UA83 | BY05 | 8 | CL | 2 | 1 | B | 21-22 | SR8 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO |  | Ch6 : BLMQI.07L8.B1E30_MQM no modulation --> doen o.k. now!!!!
32 | BJBHT.B6L8 | yes | BLECF_ | 239 | 0x | 00EF | 7 | UA83 | BY05 | 8 | CL | 1 | 10 | B | 19-20 | SR8 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO |  | Ch8 : BLMTI.06L8.B2I11_TCLIB.6L8.B2 no modulation ---> IC was removed in the tunnel, must be removed in the data base.
33 | BJBHT.A6L8 | yes | BLECF_ | 231 | 0x | 00E7 | 1 | UA83 | BY05 | 8 | CL | 1 | 9 | B | 17-18 | SR8 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 7 | 4 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
34 | BJBHT.A5L8 |  | BLECF_ | 240 | 0x | 00F0 | 6 | UA83 | BY05 | 8 | CL | 1 | 8 | B | 15-16 | SR8 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [9 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
35 | BJBHT.C4L8 |  | BLECF_ | 246 | 0x | 00F6 | 8 | UA83 | BY05 | 8 | CL | 1 | 7 | B | 13-14 | SR8 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO |  | CFC Reset needed --> call for verification!
36 | BJBHT.B4L8 | yes | BLECF_ | 248 | 0x | 00F8 | 1 | UA83 | BY05 | 8 | CL | 1 | 6 | B | 11-12 | SR8 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
37 | BJBHT.A4L8 | yes | BLECF_ | 243 | 0x | 00F3 | 2 | UA83 | BY05 | 8 | CL | 1 | 5 | B | 9-10 | SR8 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
38 | BJBHT.A3L8 |  | BLECF_ | 250 | 0x | 00FA | 8 | UA83 | BY05 | 8 | CL | 1 | 4 | B | 7-8 | SR8 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
39 | BJBHT.A2L8 |  | BLECF_ | 261 | 0x | 0105 | 8 | UA83 | BY05 | 8 | CL | 1 | 3 | B | 5-6 | SR8 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
40 | BJBHT.B1L8 |  | BLECF_ | 269 | 0x | 010D | 5 | UA83 | BY05 | 8 | CL | 1 | 2 | B | 3-4 | SR8 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO
41 | BJBHT.A1L8 | yes | BLECF_ | 180 | 0x | 00B4 | 1 | UA83 | BY05 | 8 | CL | 1 | 1 | B | 1-2 | SR8 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | OK | OK | OK | OK | OK | OK | OK | [10 empty] | NO |  | Connector not porperly connected
42 | BJBHT.A1R8 | yes | BLECF_ | 78 | 0x | 004E | 1 | UA87 | BY06 | 8 | CR | 1 | 1 | B | 1-2 | SR8 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
43 | BJBHT.B1R8 |  | BLECF_ | 75 | 0x | 004B | 5 | UA87 | BY06 | 8 | CR | 1 | 2 | B | 3-4 | SR8 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
44 | BJBHT.A2R8 |  | BLECF_ | 6 | 0x | 0006 | 8 | UA87 | BY06 | 8 | CR | 1 | 3 | B | 5-6 | SR8 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
45 | BJBHT.A3R8 |  | BLECF_ | 528 | 0x | 0210 | 4 | UA87 | BY06 | 8 | CR | 1 | 4 | B | 7-8 | SR8 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO |  | CFC Reset needed --> call for verification!
46 | BJBHT.A4R8 | yes | BLECF_ | 71 | 0x | 0047 | 7 | UA87 | BY06 | 8 | CR | 1 | 5 | B | 9-10 | SR8 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | NO |  | CH2 : BLMEI.04R8.I2E20 (no modulation signal) CH4 : BLMEI.04R8.B2E10_TCTVB.4R8
47 | BJBHT.B4R8 | yes | BLECF_ | 287 | 0x | 011F | 4 | UA87 | BY06 | 8 | CR | 1 | 6 | B | 11-12 | SR8 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO |  | Noise on all channels
48 | BJBHT.C4R8 | yes | BLECF_ | 278 | 0x | 0116 | 1 | UA87 | BY06 | 8 | CR | 1 | 7 | B | 13-14 | SR8 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
49 | BJBHT.D4R8 | yes | BLECF_ | 209 | 0x | 00D1 | 8 | UA87 | BY06 | 8 | CR | 1 | 8 | B | 15-16 | SR8 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
50 | BJBHT.A5R8 |  | BLECF_ | 233 | 0x | 00E9 | 6 | UA87 | BY06 | 8 | CR | 1 | 9 | B | 17-18 | SR8 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
51 | BJBHT.A6R8 | yes | BLECF_ | 267 | 0x | 010B | 6 | UA87 | BY06 | 8 | CR | 1 | 10 | B | 19-20 | SR8 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
52 | BJBHT.B6R8 | yes | BLECF_ | 29 | 0x | 001D | 6 | UA87 | BY06 | 8 | CR | 2 | 1 | B | 21-22 | SR8 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
53 | BJBHT.C6R8 | yes | BLECF_ | 27 | 0x | 001B | 6 | UA87 | BY06 | 8 | CR | 2 | 2 | B | 23-24 | SR8 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | NO |  | NES 18 ! To be checked if it is connected!
54 | BJBHT.A7R8 | yes | BLECF_ | 26 | 0x | 001A | 6 | UA87 | BY06 | 8 | CR | 2 | 3 | C | 1-2 | SR8 | BY05 | C | 10 | 1-2 | BY02 | R | 3 | 5 | 2 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | NO |  | CH5: BLMQI.07R8.B2E10_MQM NES 18 ! To be checked if it is connected!
55 | BJBHT.A8R8 |  | BLECF_ | 24 | 0x | 0018 | 6 | UA87 | BY06 | 8 | CR | 2 | 4 | C | 3-4 | SR8 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
56 | BJBHT.B8R8 | yes | BLECF_ | 25 | 0x | 0019 | 6 | UA87 | BY06 | 8 | CR | 2 | 5 | C | 5-6 | SR8 | BY05 | C | 10 | 5-6 | BY02 | C | 2 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
57 | BJBHT.C8R8 | yes | BLECF_ | 127 | 0x | 007F | 6 | UA87 | BY06 | 8 | CR | 2 | 6 | C | 7-8 | SR8 | BY05 | C | 10 | 7-8 | BY02 | R | 3 | 6 | 3 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
58 | BJBHT.A9R8 |  | BLECF_ | 80 | 0x | 0050 | 6 | UA87 | BY06 | 8 | CR | 2 | 7 | C | 9-10 | SR8 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 6 | 3 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
59 | BJBHT.A10R8 |  | BLECF_ | 353 | 0x | 0161 | 6 | UA87 | BY06 | 8 | CR | 2 | 8 | C | 11-12 | SR8 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 7 | 4 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
60 | BJBHT.A11R8 |  | BLECF_ | 39 | 0x | 0027 | 6 | UA87 | BY06 | 8 | CR | 2 | 9 | C | 13-14 | SR8 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
61 | BJBHT.B11R8 | yes | BLECF_ | 113 | 0x | 0071 | 6 | UA87 | BY06 | 8 | CR | 2 | 10 | C | 15-16 | SR8 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
62 | BJBHT.C11R8 |  | BLECF_ | 300 | 0x | 012C | 6 | UA87 | BY06 | 8 | CR | 3 | 1 | C | 17-18 | SR8 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 7 | 4 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [9 empty] | NES18 | NO
63 | BJBHT.A12R8 |  | BLECF_ | 272 | 0x | 0110 | 6 | Arc | BYPLM.A12R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
64 | BJBHT.A13R8 |  | BLECF_ | 51 | 0x | 0033 | 6 | Arc | BYPLM.A13R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
65 | BJBHT.A14R8 |  | BLECF_ | 54 | 0x | 0036 | 6 | Arc | BYPLM.A14R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
66 | BJBHT.A15R8 |  | BLECF_ | 30 | 0x | 001E | 6 | Arc | BYPLM.A15R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
67 | BJBHT.A16R8 |  | BLECF_ | 142 | 0x | 008E | 6 | Arc | BYPLM.A16R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 00 | [9 empty] | NO
68 | BJBHT.A17R8 |  | BLECF_ | 61 | 0x | 003D | 6 | Arc | BYPLM.A17R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
69 | BJBHT.A18R8 |  | BLECF_ | 148 | 0x | 0094 | 6 | Arc | BYPLM.A18R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | 00 | [9 empty] | NO
70 | BJBHT.A19R8 |  | BLECF_ | 15 | 0x | 000F | 6 | Arc | BYPLM.A19R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
71 | BJBHT.A20R8 |  | BLECF_ | 294 | 0x | 0126 | 6 | Arc | BYPLM.A20R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
72 | BJBHT.A21R8 |  | BLECF_ | 286 | 0x | 011E | 6 | Arc | BYPLM.A21R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
73 | BJBHT.A22R8 |  | BLECF_ | 206 | 0x | 00CE | 6 | Arc | BYPLM.A22R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
74 | BJBHT.A23R8 |  | BLECF_ | 270 | 0x | 010E | 6 | Arc | BYPLM.A23R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
75 | BJBHT.A24R8 |  | BLECF_ | 131 | 0x | 0083 | 6 | Arc | BYPLM.A24R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
76 | BJBHT.A25R8 |  | BLECF_ | 227 | 0x | 00E3 | 6 | Arc | BYPLM.A25R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
77 | BJBHT.A26R8 |  | BLECF_ | 253 | 0x | 00FD | 6 | Arc | BYPLM.A26R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
78 | BJBHT.A27R8 |  | BLECF_ | 62 | 0x | 003E | 6 | Arc | BYPLM.A27R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
79 | BJBHT.A28R8 |  | BLECF_ | 153 | 0x | 0099 | 6 | Arc | BYPLM.A28R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
80 | BJBHT.A29R8 |  | BLECF_ | 106 | 0x | 006A | 6 | Arc | BYPLM.A29R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
81 | BJBHT.A30R8 |  | BLECF_ | 252 | 0x | 00FC | 6 | Arc | BYPLM.A30R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
82 | BJBHT.A31R8 |  | BLECF_ | 104 | 0x | 0068 | 6 | Arc | BYPLM.A31R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
83 | BJBHT.A32R8 |  | BLECF_ | 271 | 0x | 010F | 6 | Arc | BYPLM.A32R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
84 | BJBHT.A33R8 |  | BLECF_ | 225 | 0x | 00E1 | 6 | Arc | BYPLM.A33R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
85 | BJBHT.A34R8 |  | BLECF_ | 297 | 0x | 0129 | 6 | Arc | BYPLM.A34R8 | 8 | R |  |  |  | 7-9 | SR8 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | O.K | O.K | O.K | O.K | O.K | O.K | O.K | [10 empty] | NO
| Door | [46 empty] | 1800433 | UA87 | BJBAP.A4R8=RA87
[27 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k | [10 empty] | 1800474 | UA87 | BJBAP.C6R8=RA87
[27 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current) | [10 empty] | 1802020 | UA87 | BJBAP.A7R8=RH87
[7 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing | [19 empty] | 1802043 | UA87 | BJBAP.C11R8=R88
[7 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed | [11 empty] | 1802051 | UA83 | BJBAP.A5L8=RA83
[7 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[27 empty] | O.K. | No problems observed
[45 empty] | L8 | BLMQI.07L8.B1E30_MQM | BLMQI.07L8.B1E30_MQM:LOSS_RS09
[45 empty] | L8 | BLMTI.06L8.B2I11_TCLIB.6L8.B2 | BLMTI.06L8.B2I11_TCLIB.6L8.B2:LOSS_RS09
Note on P52: effinger: Fibre 2 : C22
Note on U52: effinger: Fibre 2: C13 10
Note on AE84: Ewald Effinger: E288 5/9/14 15:40 E288 8/9/14 17:00 o.k.
Note on AE70: Ewald Effinger: 8D11 5/9/14 15:40 8D11 8/9/14 17:30 o.k
Note on E87: Ewald Effinger: Number correct!
Note on E86: Ewald Effinger: Number correct!
Note on AU68: Ewald Effinger: In the new NES18 cable was a pin swap. Now it is o.k! 1. Check connector of the CFC 2. disconnecting of cable
Note on AK56: Ewald Effinger: was blogged to negative
Note on AT56: Ewald Effinger: HV was not o.k
Note on AE20: Ewald Effinger: ↑ increased slowly counter reset on the 3.2.14 11:30
Note on E36: Ewald Effinger: was BLECF_0221
Note on E46: Ewald Effinger: Was BLECF_0265
Note on AE92: Ewald Effinger: Fibers on the wrong input 3-4 instead of 1-2
Note on AT46: Ewald Effinger: HV missing
Note on AE38: Ewald Effinger: LF ↑↑↑ , fibers and CFC cleaned! O.k
Note on AE28: Ewald Effinger: LF ↑↑↑ , connector was badly plugged!
Note on AC36: Ewald Effinger: LF ↑↑↑, GOH checked
Note on AF36: Ewald Effinger: LF ↑↑↑, GOH checked
Note on AB20: Ewald Effinger: ↑ increased slowly
Note on AF20: Ewald Effinger: ↑ increased slowly
Note on AU67: Ewald Effinger: Channel 2 have peaks of 400 BLMbits => peaks disappeared still some noise but this can come from the cable length
Note on AU66: Ewald Effinger: Wrong number -> cheking of card number in the tunnel
Note on AU56: Ewald Effinger: Channel 2 is blogged to negative! 1. Check connector of the CFC 2. disconnecting of cable
Note on AU50: Ewald Effinger: Wrong number -> cheking of card number in the tunnel
Note on AU92: Ewald Effinger: DAC Reset with HV not executed ( reset of card needed and need to be checked again)
Note on AU87: Ewald Effinger: Numbers inverted! 0099 -> Checking of fiber in the surface -> Fibers mixed up in the BLETC
Note on AU86: Ewald Effinger: Numbers inverted! 006A -> Checking of fiber in the surface -> Fibers mixed up in the BLETC
Note on AU70: Ewald Effinger: Powersupply exchanged! The transmission was alternating between on/off! O.k. for now!
Note on AU72: Ewald Effinger: Powersupply exchanged! The transmission was alternating between on/off! O.k. for now!
Note on AU77: Ewald Effinger: Powersupply exchanged! The transmission was alternating between on/off! O.k. for now!
Note on AU80: Ewald Effinger: Powersupply exchanged! The transmission was alternating between on/off! O.k. for now!
Note on AU81: Ewald Effinger: Powersupply exchanged! The transmission was alternating between on/off! O.k. for now!

## Vertical Slice
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | other
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L2 | BLECF_ | 399 | 0x | 018F | 6 | Arc | BYPLM.A33L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16
2 | BJBHT.A32L2 | BLECF_ | 348 | 0x | 015C | 6 | Arc | BYPLM.A32L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8
3 | BJBHT.A31L2 | BLECF_ | 497 | 0x | 01F1 | 6 | Arc | BYPLM.A31L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16
4 | BJBHT.A30L2 | BLECF_ | 432 | 0x | 01B0 | 6 | Arc | BYPLM.A30L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8
5 | BJBHT.A29L2 | BLECF_ | 520 | 0x | 0208 | 6 | Arc | BYPLM.A29L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16
6 | BJBHT.A28L2 | BLECF_ | 439 | 0x | 01B7 | 6 | Arc | BYPLM.A28L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8
7 | BJBHT.A27L2 | BLECF_ | 342 | 0x | 0156 | 6 | Arc | BYPLM.A27L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16
8 | BJBHT.A26L2 | BLECF_ | 460 | 0x | 01CC | 6 | Arc | BYPLM.A26L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8
9 | BJBHT.A25L2 | BLECF_ | 388 | 0x | 0184 | 6 | Arc | BYPLM.A25L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16
10 | BJBHT.A24L2 | BLECF_ | 440 | 0x | 01B8 | 6 | Arc | BYPLM.A24L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8
11 | BJBHT.A23L2 | BLECF_ | 468 | 0x | 01D4 | 6 | Arc | BYPLM.A23L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16
12 | BJBHT.A22L2 | BLECF_ | 456 | 0x | 01C8 | 6 | Arc | BYPLM.A22L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8
13 | BJBHT.A21L2 | BLECF_ | 111 | 0x | 006F | 6 | Arc | BYPLM.A21L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16
14 | BJBHT.A20L2 | BLECF_ | 11 | 0x | 000B | 6 | Arc | BYPLM.A20L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8
15 | BJBHT.A19L2 | BLECF_ | 715 | 0x | 02CB | 6 | Arc | BYPLM.A19L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16
16 | BJBHT.A18L2 | BLECF_ | 366 | 0x | 016E | 6 | Arc | BYPLM.A18L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8
17 | BJBHT.A17L2 | BLECF_ | 427 | 0x | 01AB | 6 | Arc | BYPLM.A17L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16
18 | BJBHT.A16L2 | BLECF_ | 385 | 0x | 0181 | 6 | Arc | BYPLM.A16L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8
19 | BJBHT.A15L2 | BLECF_ | 444 | 0x | 01BC | 6 | Arc | BYPLM.A15L2 | 2 | L |  |  |  | 5-6 | SR2 | BY05 | F | 4 | 5-6 | BY02 | L | 1 | 9 | 6 | 9-16
20 | BJBHT.A14L2 | BLECF_ | 380 | 0x | 017C | 6 | Arc | BYPLM.A14L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8
21 | BJBHT.A13L2 | BLECF_ | 347 | 0x | 015B | 6 | Arc | BYPLM.A13L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16
22 | BJBHT.A12L2 | BLECF_ | 200 | 0x | 00C8 | 8 | Arc | BYPLM.A12L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8
23 | BJBHT.C11L2 | BLECF_ | 583 | 0x | 0247 | 6 | UA23 | BY02 | 2 | CL | 3 | 2 | C | 19-20 | SR2 | BY05 | C | 4 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16
24 | BJBHT.B11L2 | BLECF_ | 343 | 0x | 0157 | 6 | UA23 | BY02 | 2 | CL | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 10 | 7 | 9-16
25 | BJBHT.A11L2 | BLECF_ | 378 | 0x | 017A | 6 | UA23 | BY02 | 2 | CL | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 1-8
26 | BJBHT.B10L2 | BLECF_ | 134 | 0x | 0086 | 6 | UA23 | BY02 | 2 | CL | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 1-8
27 | BJBHT.A10L2 | BLECF_ | 50 | 0x | 0032 | 7 | UA23 | BY02 | 2 | CL | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 9-16
28 | BJBHT.A9L2 | BLECF_ | 517 | 0x | 0205 | 6 | UA23 | BY02 | 2 | CL | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 9-16
29 | BJBHT.C8L2 | BLECF_ | 357 | 0x | 0165 | 6 | UA23 | BY02 | 2 | CL | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 6 | 3 | 1-8
30 | BJBHT.B8L2 | BLECF_ | 333 | 0x | 014D | 6 | UA23 | BY02 | 2 | CL | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 9 | 6 | 1-8
31 | BJBHT.A8L2 | BLECF_ | 491 | 0x | 01EB | 6 | UA23 | BY02 | 2 | CL | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 9-16
32 | BJBHT.A7L2 | BLECF_ | 362 | 0x | 016A | 6 | UA23 | BY02 | 2 | CL | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 5 | 2 | 9-16
33 | BJBHT.C6L2 | BLECF_ | 189 | 0x | 00BD | 6 | UA23 | BY02 | 2 | CL | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 1-8
34 | BJBHT.B6L2 | BLECF_ | 344 | 0x | 0158 | 6 | UA23 | BY02 | 2 | CL | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8
35 | BJBHT.A6L2 | BLECF_ | 474 | 0x | 01DA | 6 | UA23 | BY02 | 2 | CL | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16
36 | BJBHT.A5L2 | BLECF_ | 328 | 0x | 0148 | 6 | UA23 | BY02 | 2 | CL | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16
37 | BJBHT.D4L2 | BLECF_ | 371 | 0x | 0173 | 8 | UA23 | BY02 | 2 | CL | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8
X | BJBHT.TEST | BLECF_ | 208 | 0x | 00D0 | 2 | UA23 | BY02 | 2 | CL | 3 | 10 | C | 23-24 | SR2 | BY05 | C | 4 | 11-12 | BY02 | V | 1 | 6 | 3 | 1-8 | [16 empty] | dismounted
38 | BJBHT.C4L2 | BLECF_ | 181 | 0x | 00B5 | 1 | UA23 | BY02 | 2 | CL | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8
39 | BJBHT.B4L2 | BLECF_ | 320 | 0x | 0140 | 4 | UA23 | BY02 | 2 | CL | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16
40 | BJBHT.A4L2 | BLECF_ | 516 | 0x | 0204 | 7 | UA23 | BY02 | 2 | CL | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8
41 | BJBHT.A3L2 | BLECF_ | 324 | 0x | 0144 | 4 | UA23 | BY02 | 2 | CL | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16
42 | BJBHT.A2L2 | BLECF_ | 188 | 0x | 00BC | 8 | UA23 | BY02 | 2 | CL | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8
43 | BJBHT.B1L2 | BLECF_ | 510 | 0x | 01FE | 5 | UA23 | BY02 | 2 | CL | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16
44 | BJBHT.A1L2 | BLECF_ | 476 | 0x | 01DC | 1 | UA23 | BY02 | 2 | CL | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8
45 | BJBHT.A1R2 | BLECF_ | 706 | 0x | 02C2 | 1 | UA27 | BY03 | 2 | CR | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8
46 | BJBHT.B1R2 | BLECF_ | 606 | 0x | 025E | 5 | UA27 | BY03 | 2 | CR | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16
47 | BJBHT.A2R2 | BLECF_ | 675 | 0x | 02A3 | 8 | UA27 | BY03 | 2 | CR | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8
48 | BJBHT.A3R2 | BLECF_ | 641 | 0x | 0281 | 4 | UA27 | BY03 | 2 | CR | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16
49 | BJBHT.A4R2 | BLECF_ | 719 | 0x | 02CF | 5 | UA27 | BY03 | 2 | CR | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8
50 | BJBHT.B4R2 | BLECF_ | 738 | 0x | 02E2 | 2 | UA27 | BY03 | 2 | CR | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16
51 | BJBHT.C4R2 | BLECF_ | 652 | 0x | 028C | 1 | UA27 | BY03 | 2 | CR | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8
52 | BJBHT.D4R2 | BLECF_ | 665 | 0x | 0299 | 8 | UA27 | BY03 | 2 | CR | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8
53 | BJBHT.A5R2 | BLECF_ | 740 | 0x | 02E4 | 6 | UA27 | BY03 | 2 | CR | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16
54 | BJBHT.A6R2 | BLECF_ | 666 | 0x | 029A | 1 | UA27 | BY03 | 2 | CR | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16
55 | BJBHT.B6R2 | BLECF_ | 722 | 0x | 02D2 | 7 | UA27 | BY03 | 2 | CR | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 5 | 2 | 1-8
56 | BJBHT.A7R2 | BLECF_ | 658 | 0x | 0292 | 6 | UA27 | BY03 | 2 | CR | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16
57 | BJBHT.A8R2 | BLECF_ | 626 | 0x | 0272 | 6 | UA27 | BY03 | 2 | CR | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8
58 | BJBHT.B8R2 | BLECF_ | 406 | 0x | 0196 | 6 | UA27 | BY03 | 2 | CR | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16
59 | BJBHT.C8R2 | BLECF_ | 4 | 0x | 0004 | 6 | UA27 | BY03 | 2 | CR | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 1-8
60 | BJBHT.A9R2 | BLECF_ | 279 | 0x | 0117 | 6 | UA27 | BY03 | 2 | CR | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 10 | 7-8 | BY02 | R | 3 | 6 | 3 | 9-16
61 | BJBHT.A10R2 | BLECF_ | 97 | 0x | 0061 | 6 | UA27 | BY03 | 2 | CR | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 18 | 14 | 1-8
62 | BJBHT.B10R2 | BLECF_ | 721 | 0x | 02D1 | 6 | UA27 | BY03 | 2 | CR | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 7 | 4 | 1-8
63 | BJBHT.A11R2 | BLECF_ | 128 | 0x | 0080 | 6 | UA27 | BY03 | 2 | CR | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16
64 | BJBHT.B11R2 | BLECF_ | 627 | 0x | 0273 | 6 | UA27 | BY03 | 2 | CR | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8
65 | BJBHT.C11R2 | BLECF_ | 689 | 0x | 02B1 | 6 | UA27 | BY03 | 2 | CR | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 7 | 4 | 9-16
66 | BJBHT.A12R2 | BLECF_ | 681 | 0x | 02A9 | 8 | Arc | BYPLM.A12R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8
67 | BJBHT.A13R2 | BLECF_ | 687 | 0x | 02AF | 6 | Arc | BYPLM.A13R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16
68 | BJBHT.A14R2 | BLECF_ | 684 | 0x | 02AC | 6 | Arc | BYPLM.A14R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8
69 | BJBHT.A15R2 | BLECF_ | 484 | 0x | 01E4 | 6 | Arc | BYPLM.A15R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16
70 | BJBHT.A16R2 | BLECF_ | 686 | 0x | 02AE | 6 | Arc | BYPLM.A16R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8
71 | BJBHT.A17R2 | BLECF_ | 143 | 0x | 008F | 6 | Arc | BYPLM.A17R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16
72 | BJBHT.A18R2 | BLECF_ | 463 | 0x | 01CF | 6 | Arc | BYPLM.A18R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8
73 | BJBHT.A19R2 | BLECF_ | 379 | 0x | 017B | 6 | Arc | BYPLM.A19R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16
74 | BJBHT.A20R2 | BLECF_ | 382 | 0x | 017E | 6 | Arc | BYPLM.A20R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8
75 | BJBHT.A21R2 | BLECF_ | 451 | 0x | 01C3 | 6 | Arc | BYPLM.A21R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16
76 | BJBHT.A22R2 | BLECF_ | 404 | 0x | 0194 | 6 | Arc | BYPLM.A22R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8
77 | BJBHT.A23R2 | BLECF_ | 84 | 0x | 0054 | 6 | Arc | BYPLM.A23R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16
78 | BJBHT.A24R2 | BLECF_ | 418 | 0x | 01A2 | 6 | Arc | BYPLM.A24R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8
79 | BJBHT.A25R2 | BLECF_ | 390 | 0x | 0186 | 6 | Arc | BYPLM.A25R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16
80 | BJBHT.A26R2 | BLECF_ | 214 | 0x | 00D6 | 6 | Arc | BYPLM.A26R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8
81 | BJBHT.A27R2 | BLECF_ | 694 | 0x | 02B6 | 6 | Arc | BYPLM.A27R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16
82 | BJBHT.A28R2 | BLECF_ | 685 | 0x | 02AD | 6 | Arc | BYPLM.A28R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8
83 | BJBHT.A29R2 | BLECF_ | 725 | 0x | 02D5 | 6 | Arc | BYPLM.A29R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16
84 | BJBHT.A30R2 | BLECF_ | 195 | 0x | 00C3 | 6 | Arc | BYPLM.A30R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8
85 | BJBHT.A31R2 | BLECF_ | 690 | 0x | 02B2 | 6 | Arc | BYPLM.A31R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16
86 | BJBHT.A32R2 | BLECF_ | 360 | 0x | 0168 | 6 | Arc | BYPLM.A32R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8
87 | BJBHT.A33R2 | BLECF_ | 23 | 0x | 0017 | 6 | Arc | BYPLM.A33R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16
88 | BJBHT.A34R2 | BLECF_ | 724 | 0x | 02D4 | 6 | Arc | BYPLM.A34R2 | 2 | R |  |  |  | 7-9 | SR2 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on T23: effinger: Fibre 7 & 8 broken, spare fibre used
Note on F82: effinger: Checked-----> BLECF_0382 installed! Wrong number 0x017E (382) instead of 0x0188 (392) ; to be check in the tunnel; Card number checked: --> BLECF_0382 installed; sent number 382

## Overview
267 | BLECF_ | 457
551 | BLECF_ | #REF!
578 | BLECF_ | 757
435 | BLECF_ | 674
574 | BLECF_ | 753
596 | #REF! | 750
594 | BLECF_ | 749
492 | BLECF_ | 748
519 | BLECF_ | 744
502 | BLECF_ | 743
619 | BLECF_ | 742
422 | BLECF_ | 741
622 | BLECF_ | 739
615 | BLECF_ | 737
625 | BLECF_ | 736
632 | BLECF_ | 735
628 | BLECF_ | 50
616 | BLECF_ | 733
618 | BLECF_ | 732
633 | BLECF_ | 731
449 | BLECF_ | 729
613 | BLECF_ | 728
412 | BLECF_ | 727
597 | BLECF_ | 726
404 | BLECF_ | 725
409 | BLECF_ | 724
418 | BLECF_ | 717
448 | BLECF_ | 711
599 | BLECF_ | 710
64 | BLECF_ | 708
413 | BLECF_ | 705
630 | BLECF_ | 704
431 | BLECF_ | 703
415 | BLECF_ | 702
629 | BLECF_ | 701
612 | BLECF_ | 699
614 | BLECF_ | 698
601 | BLECF_ | 697
620 | BLECF_ | 695
402 | BLECF_ | 694
631 | BLECF_ | 693
623 | BLECF_ | 691
406 | BLECF_ | 690
621 | BLECF_ | 688
388 | BLECF_ | 687
391 | BLECF_ | 686
403 | BLECF_ | 685
389 | BLECF_ | 684
626 | BLECF_ | 683
651 | BLECF_ | 682
387 | BLECF_ | 681
559 | BLECF_ | 680
636 | #REF! | 679
638 | #REF! | 678
451 | BLECF_ | 677
285 | BLECF_ | 676
434 | BLECF_ | 574
609 | #REF! | 673
437 | BLECF_ | 571
595 | #REF! | 671
510 | BLECF_ | 670
426 | BLECF_ | 668
262 | BLECF_ | 667
444 | BLECF_ | 175
503 | BLECF_ | 664
598 | BLECF_ | 663
494 | BLECF_ | 662
634 | BLECF_ | 661
606 | BLECF_ | 660
512 | BLECF_ | 659
432 | BLECF_ | 657
603 | BLECF_ | 656
500 | #REF! | 655
433 | BLECF_ | 654
286 | BLECF_ | 653
516 | BLECF_ | 651
554 | BLECF_ | 650
640 | BLECF_ | 649
602 | #REF! | 648
294 | BLECF_ | 647
414 | BLECF_ | 646
421 | BLECF_ | 645
533 | BLECF_ | 57
513 | BLECF_ | 643
509 | BLECF_ | 642
572 | BLECF_ | 640
280 | BLECF_ | 251
440 | BLECF_ | 637
284 | BLECF_ | 636
279 | BLECF_ | 635
617 | BLECF_ | 634
661 | BLECF_ | 633
582 | BLECF_ | 632
501 | BLECF_ | 631
425 | BLECF_ | 720
442 | BLECF_ | 629
563 | BLECF_ | 625
417 | BLECF_ | 623
428 | BLECF_ | 622
610 | BLECF_ | 621
424 | BLECF_ | 620
410 | BLECF_ | 619
604 | BLECF_ | 383
624 | BLECF_ | 617
493 | BLECF_ | 616
507 | BLECF_ | 615
423 | BLECF_ | 614
499 | BLECF_ | 613
1 | BLECF_ | 612
469 | BLECF_ | 527
2 | BLECF_ | 610
436 | BLECF_ | 609
416 | BLECF_ | 607
608 | #REF! | 605
427 | BLECF_ | 604
611 | BLECF_ | 603
447 | BLECF_ | 602
411 | BLECF_ | 601
430 | BLECF_ | 600
511 | BLECF_ | 599
600 | BLECF_ | 598
273 | BLECF_ | 597
605 | BLECF_ | 596
276 | BLECF_ | 401
263 | BLECF_ | 594
277 | BLECF_ | 593
278 | BLECF_ | 592
508 | BLECF_ | 591
656 | BLECF_ | 590
565 | BLECF_ | 589
268 | BLECF_ | 588
644 | BLECF_ | 587
591 | BLECF_ | 586
537 | BLECF_ | 585
524 | BLECF_ | 584
549 | BLECF_ | 582
575 | BLECF_ | 581
120 | BLECF_ | 580
592 | BLECF_ | 579
514 | BLECF_ | 578
445 | BLECF_ | 577
439 | BLECF_ | 576
446 | BLECF_ | 575
438 | BLECF_ | 103
497 | BLECF_ | 572
3 | BLECF_ | 208
301 | BLECF_ | 570
4 | BLECF_ | 569
531 | BLECF_ | 568
297 | BLECF_ | 567
5 | BLECF_ | 566
6 | BLECF_ | 565
7 | BLECF_ | 564
8 | BLECF_ | 563
466 | BLECF_ | 562
9 | BLECF_ | 561
648 | BLECF_ | 638
639 | BLECF_ | 559
535 | BLECF_ | 558
10 | BLECF_ | 557
523 | BLECF_ | 235
11 | BLECF_ | 555
650 | BLECF_ | 360
505 | BLECF_ | 552
544 | #REF! | 551
265 | BLECF_ | 550
12 | BLECF_ | 549
13 | BLECF_ | 548
543 | BLECF_ | 547
566 | BLECF_ | 546
14 | BLECF_ | 545
637 | BLECF_ | 544
450 | #REF! | 543
15 | BLECF_ | 542
456 | #REF! | 541
16 | BLECF_ | 540
299 | BLECF_ | 539
17 | BLECF_ | 538
540 | BLECF_ | 537
443 | BLECF_ | 536
18 | BLECF_ | 534
19 | BLECF_ | 533
20 | BLECF_ | 532
530 | BLECF_ | 531
272 | BLECF_ | 530
21 | BLECF_ | 611
581 | BLECF_ | 525
476 | BLECF_ | 523
308 | BLECF_ | 522
585 | BLECF_ | 521
326 | BLECF_ | 520
483 | BLECF_ | 519
266 | BLECF_ | 518
349 | BLECF_ | 517
361 | BLECF_ | 516
305 | BLECF_ | 515
472 | BLECF_ | 514
642 | BLECF_ | 513
292 | BLECF_ | 512
23 | BLECF_ | 511
364 | BLECF_ | 510
525 | BLECF_ | 509
24 | BLECF_ | 508
564 | BLECF_ | 506
686 | BLECF_ | 505
470 | BLECF_ | 504
461 | #REF! | 502
283 | BLECF_ | 501
312 | BLECF_ | 500
25 | BLECF_ | #REF!
471 | BLECF_ | 499
324 | BLECF_ | 497
474 | BLECF_ | 496
318 | BLECF_ | 495
344 | BLECF_ | 583
473 | BLECF_ | 493
460 | BLECF_ | 709
352 | BLECF_ | 491
627 | BLECF_ | 490
485 | BLECF_ | 489
487 | BLECF_ | 488
454 | BLECF_ | 487
26 | BLECF_ | 486
479 | BLECF_ | 485
390 | BLECF_ | 484
680 | BLECF_ | 482
306 | BLECF_ | 481
465 | BLECF_ | 186
313 | BLECF_ | 479
314 | BLECF_ | 478
573 | BLECF_ | 477
365 | BLECF_ | 476
568 | BLECF_ | 475
356 | BLECF_ | 474
567 | BLECF_ | 473
468 | BLECF_ | 68
504 | BLECF_ | 471
683 | BLECF_ | 470
635 | BLECF_ | 469
332 | BLECF_ | 468
27 | BLECF_ | 467
29 | BLECF_ | 466
652 | #REF! | 265
681 | BLECF_ | 464
393 | BLECF_ | 463
641 | BLECF_ | 462
28 | BLECF_ | 461
329 | BLECF_ | 460
467 | BLECF_ | 459
333 | BLECF_ | 456
664 | BLECF_ | 455
30 | BLECF_ | 454
489 | BLECF_ | 453
491 | BLECF_ | 452
396 | BLECF_ | 451
31 | BLECF_ | 449
484 | BLECF_ | 448
557 | BLECF_ | 447
669 | BLECF_ | 446
659 | BLECF_ | 445
340 | BLECF_ | 444
663 | BLECF_ | 443
289 | BLECF_ | 442
690 | BLECF_ | 441
331 | BLECF_ | 440
327 | BLECF_ | 439
542 | BLECF_ | 438
288 | BLECF_ | 437
32 | BLECF_ | 436
546 | #REF! | 121
647 | BLECF_ | 434
679 | BLECF_ | 433
325 | BLECF_ | 432
666 | #REF! | 431
33 | BLECF_ | 334
34 | BLECF_ | 429
689 | BLECF_ | 428
338 | BLECF_ | 427
35 | BLECF_ | 426
36 | BLECF_ | #REF!
560 | BLECF_ | 424
310 | BLECF_ | 423
684 | BLECF_ | 700
670 | BLECF_ | 421
667 | BLECF_ | 420
561 | BLECF_ | 419
399 | BLECF_ | 418
320 | BLECF_ | 417
671 | BLECF_ | 415
555 | BLECF_ | 414
37 | BLECF_ | 413
668 | BLECF_ | 713
643 | BLECF_ | 411
464 | BLECF_ | 177
264 | BLECF_ | 409
515 | BLECF_ | 408
488 | BLECF_ | 407
541 | BLECF_ | 405
397 | BLECF_ | 404
534 | BLECF_ | 403
38 | BLECF_ | #REF!
486 | BLECF_ | 492
39 | BLECF_ | #REF!
322 | BLECF_ | 399
274 | BLECF_ | 398
40 | BLECF_ | 397
506 | BLECF_ | 395
677 | BLECF_ | 31
455 | BLECF_ | 393
41 | BLECF_ | 392
395 | BLECF_ | 382
42 | BLECF_ | 391
400 | BLECF_ | 390
550 | BLECF_ | 389
330 | BLECF_ | 388
478 | BLECF_ | 387
517 | BLECF_ | 386
339 | BLECF_ | 385
319 | BLECF_ | 384
290 | BLECF_ | 355
688 | BLECF_ | 381
341 | BLECF_ | 380
394 | BLECF_ | 379
346 | BLECF_ | 378
281 | BLECF_ | 377
287 | BLECF_ | 376
463 | BLECF_ | 375
275 | BLECF_ | 374
309 | BLECF_ | 373
658 | BLECF_ | 372
358 | BLECF_ | 371
587 | BLECF_ | 370
660 | BLECF_ | 369
672 | BLECF_ | 368
459 | BLECF_ | 367
337 | BLECF_ | 366
22 | BLECF_ | 234
348 | BLECF_ | 363
353 | BLECF_ | 362
300 | BLECF_ | 361
407 | BLECF_ | 82
682 | BLECF_ | 359
685 | BLECF_ | 358
350 | BLECF_ | 357
586 | #REF! | 18
649 | BLECF_ | 364
311 | BLECF_ | 352
665 | BLECF_ | 351
675 | BLECF_ | 228
298 | BLECF_ | 349
323 | BLECF_ | 348
342 | BLECF_ | 347
570 | BLECF_ | 346
687 | BLECF_ | 345
355 | BLECF_ | 344
345 | BLECF_ | 343
328 | BLECF_ | 342
457 | BLECF_ | 341
539 | BLECF_ | 340
458 | BLECF_ | 339
296 | BLECF_ | 338
482 | BLECF_ | 337
43 | BLECF_ | 336
552 | BLECF_ | 335
673 | BLECF_ | 192
351 | BLECF_ | 333
282 | BLECF_ | 331
553 | BLECF_ | 330
317 | BLECF_ | 329
357 | BLECF_ | 328
453 | BLECF_ | 326
293 | BLECF_ | 325
362 | BLECF_ | 324
291 | BLECF_ | 323
269 | BLECF_ | 322
321 | BLECF_ | 321
360 | BLECF_ | 320
419 | BLECF_ | 319
498 | BLECF_ | 21
307 | BLECF_ | 81
522 | BLECF_ | 316
538 | BLECF_ | 315
475 | BLECF_ | 314
584 | BLECF_ | 313
590 | BLECF_ | 288
303 | BLECF_ | 311
316 | BLECF_ | 310
480 | BLECF_ | 309
662 | BLECF_ | 308
556 | BLECF_ | 307
302 | BLECF_ | 306
429 | BLECF_ | 305
315 | BLECF_ | 317
304 | BLECF_ | 303
133 | BLECF_ | 301
44 | BLECF_ | 300
45 | BLECF_ | 299
46 | BLECF_ | #REF!
47 | BLECF_ | 297
48 | BLECF_ | 295
49 | BLECF_ | 294
50 | BLECF_ | 293
51 | BLECF_ | 292
52 | BLECF_ | 291
53 | BLECF_ | #REF!
54 | BLECF_ | 289
55 | BLECF_ | #REF!
56 | BLECF_ | 287
57 | BLECF_ | 286
58 | BLECF_ | 285
59 | BLECF_ | 284
60 | BLECF_ | 212
61 | BLECF_ | 282
62 | BLECF_ | 281
63 | BLECF_ | 280
65 | BLECF_ | 278
66 | BLECF_ | 277
67 | BLECF_ | 276
68 | BLECF_ | 275
69 | BLECF_ | 274
655 | #REF! | 273
70 | BLECF_ | 272
71 | BLECF_ | 271
72 | BLECF_ | 270
73 | BLECF_ | 269
74 | BLECF_ | 268
75 | BLECF_ | 267
76 | BLECF_ | 266
77 | BLECF_ | 180
78 | BLECF_ | 264
79 | BLECF_ | 263
80 | BLECF_ | 262
81 | BLECF_ | 261
82 | BLECF_ | #REF!
83 | BLECF_ | #REF!
84 | BLECF_ | 257
85 | BLECF_ | 256
86 | BLECF_ | 255
87 | BLECF_ | 253
88 | BLECF_ | 252
89 | BLECF_ | #REF!
90 | BLECF_ | 250
91 | BLECF_ | 249
92 | BLECF_ | 248
93 | BLECF_ | 247
94 | BLECF_ | 246
95 | BLECF_ | #REF!
96 | BLECF_ | #REF!
97 | BLECF_ | 243
98 | BLECF_ | #REF!
99 | BLECF_ | 241
100 | BLECF_ | 240
101 | BLECF_ | 239
102 | BLECF_ | 238
103 | BLECF_ | 237
104 | BLECF_ | 236
105 | BLECF_ | #REF!
106 | BLECF_ | 233
107 | BLECF_ | 232
108 | BLECF_ | 231
109 | BLECF_ | 730
110 | BLECF_ | 229
111 | BLECF_ | 528
112 | BLECF_ | 227
558 | BLECF_ | 226
113 | BLECF_ | 225
114 | BLECF_ | 224
115 | BLECF_ | 223
116 | BLECF_ | 743
117 | BLECF_ | 220
118 | BLECF_ | 219
119 | BLECF_ | 218
529 | BLECF_ | 217
121 | BLECF_ | 216
122 | BLECF_ | 215
401 | BLECF_ | 214
123 | BLECF_ | 213
441 | BLECF_ | 211
124 | BLECF_ | 210
125 | BLECF_ | 209
126 | BLECF_ | 62
127 | BLECF_ | 207
128 | BLECF_ | 206
129 | BLECF_ | 205
130 | BLECF_ | 204
131 | BLECF_ | 202
132 | BLECF_ | 201
343 | BLECF_ | 200
536 | BLECF_ | 304
134 | BLECF_ | 198
135 | BLECF_ | 197
569 | BLECF_ | 196
405 | BLECF_ | 195
270 | BLECF_ | 194
593 | #REF! | 193
548 | BLECF_ | 230
646 | BLECF_ | 191
579 | BLECF_ | 190
354 | BLECF_ | 189
363 | BLECF_ | 188
136 | BLECF_ | 187
520 | #REF! | 624
477 | BLECF_ | 185
481 | BLECF_ | 184
462 | BLECF_ | 183
452 | #REF! | 182
359 | BLECF_ | 181
137 | BLECF_ | 154
138 | BLECF_ | 179
139 | BLECF_ | 178
140 | BLECF_ | 177
141 | BLECF_ | 176
142 | BLECF_ | 174
143 | BLECF_ | 173
144 | BLECF_ | 172
145 | BLECF_ | #REF!
146 | BLECF_ | 170
147 | BLECF_ | #REF!
148 | BLECF_ | 168
149 | BLECF_ | 167
150 | BLECF_ | 166
151 | BLECF_ | 165
152 | BLECF_ | 164
153 | BLECF_ | 163
154 | BLECF_ | 162
155 | BLECF_ | 161
156 | BLECF_ | 283
157 | BLECF_ | 159
158 | BLECF_ | 158
518 | BLECF_ | 157
159 | BLECF_ | 156
160 | BLECF_ | 153
161 | BLECF_ | 152
162 | BLECF_ | 151
163 | BLECF_ | 150
607 | BLECF_ | 149
164 | BLECF_ | 148
165 | BLECF_ | 145
166 | BLECF_ | 144
392 | BLECF_ | 143
167 | BLECF_ | 142
168 | BLECF_ | 141
169 | BLECF_ | 140
170 | BLECF_ | #REF!
171 | BLECF_ | 138
172 | BLECF_ | 137
347 | BLECF_ | 98
173 | BLECF_ | 133
174 | BLECF_ | 132
175 | BLECF_ | 131
176 | BLECF_ | 130
177 | BLECF_ | 129
178 | BLECF_ | 127
179 | BLECF_ | 126
180 | BLECF_ | 125
577 | BLECF_ | 124
181 | BLECF_ | 123
182 | BLECF_ | 122
589 | BLECF_ | 72
183 | BLECF_ | 39
184 | BLECF_ | 716
645 | #REF! | 118
185 | BLECF_ | 117
186 | BLECF_ | 116
187 | BLECF_ | 115
188 | BLECF_ | 114
189 | BLECF_ | 113
190 | BLECF_ | 112
334 | BLECF_ | 111
191 | BLECF_ | #REF!
490 | BLECF_ | 109
192 | BLECF_ | 108
193 | BLECF_ | 107
194 | BLECF_ | 106
195 | BLECF_ | 105
196 | BLECF_ | 104
583 | BLECF_ | 535
197 | BLECF_ | 102
198 | BLECF_ | 101
199 | BLECF_ | 100
200 | BLECF_ | 99
201 | BLECF_ | 712
653 | #REF! | 96
202 | BLECF_ | 95
654 | BLECF_ | 94
203 | BLECF_ | 93
204 | BLECF_ | 92
205 | BLECF_ | 91
206 | BLECF_ | 88
207 | BLECF_ | 87
208 | BLECF_ | 86
295 | BLECF_ | 85
398 | BLECF_ | 55
545 | BLECF_ | 83
209 | BLECF_ | 480
676 | BLECF_ | 245
210 | BLECF_ | 80
211 | BLECF_ | 79
212 | BLECF_ | 78
213 | BLECF_ | 77
496 | BLECF_ | 76
214 | BLECF_ | 75
215 | BLECF_ | 74
216 | BLECF_ | 844
217 | BLECF_ | 71
218 | BLECF_ | 6
219 | BLECF_ | 69
220 | BLECF_ | 472
221 | BLECF_ | 67
222 | BLECF_ | 66
223 | BLECF_ | 65
224 | BLECF_ | 63
225 | BLECF_ | 61
420 | BLECF_ | 745
226 | BLECF_ | 59
227 | BLECF_ | 58
532 | BLECF_ | 57
228 | BLECF_ | 56
229 | BLECF_ | 54
230 | BLECF_ | #REF!
231 | BLECF_ | 64
232 | BLECF_ | 51
233 | BLECF_ | 49
571 | BLECF_ | 48
234 | BLECF_ | 47
235 | BLECF_ | #REF!
236 | BLECF_ | 45
237 | BLECF_ | 44
238 | BLECF_ | 43
562 | BLECF_ | 34
239 | BLECF_ | #REF!
547 | BLECF_ | 40
240 | BLECF_ | 23
495 | BLECF_ | 38
241 | BLECF_ | #REF!
242 | BLECF_ | 36
243 | BLECF_ | 35
244 | BLECF_ | 400
245 | BLECF_ | 33
674 | BLECF_ | 221
246 | BLECF_ | 30
247 | BLECF_ | 29
248 | BLECF_ | 242
249 | BLECF_ | 27
250 | BLECF_ | 26
251 | BLECF_ | 25
252 | BLECF_ | 24
408 | BLECF_ | 751
253 | BLECF_ | 22
254 | BLECF_ | 754
255 | BLECF_ | 20
256 | BLECF_ | 19
257 | BLECF_ | 353
258 | BLECF_ | 17
259 | BLECF_ | 15
576 | BLECF_ | 14
678 | BLECF_ | 13
260 | BLECF_ | 12
335 | BLECF_ | 11
527 | BLECF_ | 10
580 | BLECF_ | 9
526 | BLECF_ | 7
521 | BLECF_ | 5
588 | BLECF_ | 3
261 | BLECF_ | 2
657 | BLECF_ | 244
271 | BLECF_ | 88
336 | BLECF_ | 715
366 | BLECF_ | 706
367 | BLECF_ | 606
368 | BLECF_ | 675
369 | BLECF_ | 641
370 | BLECF_ | 719
371 | BLECF_ | 738
372 | BLECF_ | 652
373 | BLECF_ | 665
374 | BLECF_ | 740
375 | BLECF_ | 666
376 | BLECF_ | 722
377 | BLECF_ | 658
378 | BLECF_ | 626
379 | BLECF_ | 406
380 | BLECF_ | 4
381 | BLECF_ | 279
382 | BLECF_ | 97
383 | BLECF_ | 721
384 | BLECF_ | 128
385 | BLECF_ | 627
386 | BLECF_ | 689
528 | BLECF_ | 119