| LHC- - - 1999-09-22 |
| --- |
| MPS Commissioning Procedure |
| The commissioning of the LHC machine protection system MPS aspects of the Beam Loss Monitor system commissioning |
| Abstract This document describes the set of tests which will be carried-out to validate for operation the machine protection aspects of the LHC Beam Loss Monitor system. The area concerned by these tests extends over the whole LHC machine for each of the two LHC beams. These tests include the Hardware Commissioning, the machine check-out and the tests with beam. |
| Prepared by : Bernd Dehning Laurette Ponce | Checked by : Ralph Assmann, Roger Bailey Bernd Dehning, Brennan Goddard, Eva Barbara Holzer, Verena Kain, Mike Lamont, Blanca Perea Solano Stefano Redaelli, Rudiger Schmidt, Benjamin Todd, Jorg Wenninger, Markus Zerlauth | Approved by : Rüdiger Schmidt |
| History of Changes |
| Rev. No. | Date | Pages | Description of Changes |
| 0.1 | 2007-05-16 | 12 | First draft for circulation. |

Table of Contents

1.	Introduction	4

2.	Scope	4

3.	Purpose	4

4.	The layout	4

5.	Tests performed during the hardware commissioning	5

5.1	Individual BLM system tests	5

6.	Link to other equipement	7

6.1	IntErfaceS with the Beam Interlock system	7

6.2	interface with the beam dumping system	7

7.	System tests during the machine checkout	7

7.1	Conditions required to perform tests	8

7.2	Description of the tests	8

7.3	Status of the system after the system tests	8

8.	Tests with beam.	8

8.1	Pilot of 1(1010 p+ at 450 Gev	8

8.2	43 bunches of 4(1010 p+ at 450 Gev	9

8.3	156 bunches of 9(1010 p+ at 450 Gev	9

8.4	pilot of 1(1010 p+ at 7 tev	9

8.5	43 bunches of 4(1010 p+ at 7 tev	10

8.6	156 bunches of 9(1010 p+ at 7 tev	10

8.7	936 bunches of 4(1010 p+ at 7 tev	10

8.8	936 bunches of 9(1010 p+ at 7 tev	10

8.9	Half nominal: 2808 bunches of 5(1010 p+ at 7 tev	10

8.10	Nominal: 2808 bunches of 11(1010 p+ at 7 tev	10

8.11	Stages depending on optics requiring test	10

8.12	Stages depending on crossing at IP requiring test	10

# Introduction

The beam loss monitor system protects the equipment against damages by the energy deposition of the beam. It is the only system which allows to disable the beam permit in the loss duration range between 4 and 100 turns. For longer durations there is also the quench protection system able to initiate an abort trigger.

# Scope

This document covers the tests which will be carried out to validate the operation of the MPS relevant components of the Beam Loss Monitor (BLM) system and the relation to the machine protection system (MPS) for LHC beam 1 and 2. The area concerned by these tests is the whole LHC ring and the beam dump lines TD62 and TD68. The equipments concerned are the beam loss monitor electronics BLE (BLEB, BLECF, BLERX, BLETC, BLETX), and the monitors in the ring BLM (BLMEI ,BLMES, BLMQI) and in the injection lines BLMD.

# Purpose

This document describes the procedures which will be applied for these tests and their sequence.

Each test in the test list has in front one of the following letters, defining at which interval or at which occasion the described test needs to be repeated (in the column labelled Repetition):

| N | Not to be repeated |
| --- | --- |
| S | To be repeated after every Shutdown |
| P | Periodical repetition required, like 1 x per month; details to be defined in text |
| O | To be repeated when LHC optics is changed |
| X | To be repeated when crossing scheme is changed |

This document is meant to be the reference document for the checklist which will be used during the commissioning of the MPS. Results of the tests will be documented in the MTF database.

# The layout

The BLM system is comprised of some 4000 monitors (3700 Ionisation Chambers and 280 Secondary Emission detector) distributed around the machine. The principle of the system is to measure outside the cryostat or the vacuum chamber the secondary shower particles created by protons impacting on the vacuum chamber.  When the signal is over a pre-defined threshold, the BLM system triggers a beam dump to protect against damage of equipment and prevent quenches of magnet coils.

The derivation of the thresholds is based on a combination of several simulations (loss maps, secondary shower development, quench level, detector response) and verification measurements. These dependencies are schematically depicted in Figure 1.

Figure 1: Schematically view of the signal flow of the beam loss system including the beam permit signal.

To ensure the specified reliability level of the system, some of the tests presented here have to be done regularly (see G. Guaglio, PH-D thesis, CERN-LIB), namely before every fill.

# Tests performed during the hardware commissioning

This part describes the tests which are validated during the equipment’s hardware commissioning period and have to be validated also for the machine protection system commissioning.

## Individual BLM system tests

The LHC beam loss monitors are foreseen to be hardware commissioned during the electronics installation phases. Independent tests of each monitor and each channel will be made by AB/BI/BLM to validate the connectivity and the topology of the system. All these tests are referenced in the MTF database and described in the EDMS document LHC-BLM-TP-0001.00.

### Conditions required to perform tests

Interconnection of cryostats performed in the sector

Installation of front-end mini-racks, crates, fan Unit supervised by AB/BI/BL.

Optical fibre terminations supervised by AB/BI/BL

BLM detector installation supervised by AB/BI/BL

‘daisy-chain’ high-voltage and signal cables supervised by AB/BI/BL

Installation of BLM electronics supervised by AB/BI/BL

### Conditions during the tests

Test conditions are listed for each test separately in chapter 5.1.3

- 240 V power in tunnel (240 V)

- 240 V power in tunnel + optical fibre connection + Ethernet =  env I

- 240 V power in tunnel + optical fibre connection + Ethernet + closing the tunnel for source test = env II

- 240 V power in tunnel + optical fibre connection + Ethernet + connection to the beam interlock system = env II

- data base LSA + software infrastructure = env III

### Description of the tests

This test is described in more details in the EDMS document (LHC-BLM-TP-0001.00). The procedure is identical and applied to all detectors and all locations.

|  | Rep. | Action | Test condition | Group(s) Responsible |
| --- | --- | --- | --- | --- |
| 1 | N | Tunnel Cables Test and Functional Test (MTF test 10-BLM Property6: LC(offset) w/cable&M-r): • Cable identification check (TS/EL sticker + AB/BI specific sticker) • Cable continuity check; particularly verify if there are no conductors swapped • Test of insulation, correct interconnection and current limits • The crate shall be equipped with a test card to record the measurements. • Measure the leakage current (offset) with the cable and the monitor connected. The measured current on all channels shall be below 0.5 pA. Functional tests: Test of high tension distribution, insulation and current limits (MTF test ????20-BLM) | 240 V | AB/BI |
| 2 | S | Test of insulation, correct interconnections and current limits (MTF test 10-BLM channel check (leakage, bias current) P7: Signal via source w/cable). Current source test: Check connection of the chamber up to the tunnel card | 240 V | AB/BI |
| 3 | P | High Voltage modulation test (MTF test BLM 40): Check the connectivity up to the surface card before every fill | env I | AB/BI |
| 4 | P | monitoring 10pA (MTF 30-BLM Acquisition Chain Test via test signal, P2:10pA/84 sec run.sum): continuous check of the tunnel electronics | env I | AB/BI |
| 5 | P | Double optical line comparison: continuous check of the connectivity up to the surface card. | env I | AB/BI |
| 6 | S | Acquisition Chain Test via test signal (MTF test 30-BLM, P1:100pA/84 sec run.sum) • Performed in sequence with the HV modulation test • The bias current is increased by 100 pA and the 84 second running sum is measured and checked. Bias current reduced to previous value. BLM interlock tests (MTF 30-BLM): Initiate the automatic procedure using a portable PC remotely connected to the control network initiate acquisition of the 10 pA test and HV modulation automatic analyse and display of the results | env I | AB/BI AB/CO |
| 7 | S | Functional test of full acquisition chain with Radioactive Source (MTF 50-BLM Acquisition Chain test via RA source): Create a signal on the ionisation chamber with the RA source Check the presence of signal in the corresponding DAB card | env II | AB/BI |
| 8 | P | Interlock test (MTF test 60-BLM) Threshold Channel and matrix Test (MTF test 60-BLM Property1: Threshold/Ch Matrix Test) Disabling of user beam permit (MTF test 60-BLM Property2: TTable User Beam Permit T) Beam permit transmission from all threshold comparators to the last combiner card (MTF test 60-BLM Property3: UBP Check BLETC to BLECS) Beam permit transmission from all threshold comparators to the last combiner card (MTF test 60-BLM Property4: UBP Check BLECS to CIBU) | env III | AB/BI |
| 9 | ? | EMC test (MTF test 70-BLM) | ? | AB/BI |
| 10 | ? | Beam energy reception test (MTF test 80-BLM) | ? | AB/BI |

### Status of the system after tests

After these tests, the hardware is fully installed, cables are connected to the monitors and to electronics; fibre link and digital acquisition are operational. The logic of the interlocks has been checked and the system is ready to be tested with the beam interlock system.

# Link to other equipement

The interfaces listed in this paragraph concern only the ones in relation with the Machine Protection System; it does not describe any procedures to test the interfaces to protect individual equipment but only summarise the logic applied in the design.

## IntErfaceS with the Beam Interlock system

### SIGNALS between BLM SYSTEM and Beam INTERLOCK SYSTEM

The BLM system produces a beam permit from each monitor which are concentrated in 2 user beam permits (one for mask able and one for unmask able channels) and sent to 2 CIBU (mask able and unmask able) per octant.

The “Beam Info” signal given by the BIS is used to ensure that there is no beam before starting the periodic tests.

The beam energy is used to change the beam permit thresholds accordingly. The signal is received by the timing system via the CRTV VME module.

### Conditions and sequence for a beam dump

The beam user permit is changed to FALSE if:

- One of the 12 running sums for each channel is over the corresponding threshold.

- The comparison of the thresholds table loaded on the FPGA with the table stored in the database failed.

- Failure of one of the periodic test (loss of communication) on any of the channels.

- Test mode of the BLMs.

Figure 2: Schematic view of beam permit, beam info and beam energy flow.

## interface with the beam dumping system

Direct interlocked monitors are the last protection in case of failure of the standard BLM system or of the BIS. There are two monitors per beam, one on the TCDQ and one on the TCSG. The threshold-signal comparison for the beam dump is done by comparing the voltage which is proportional to the losses with a reference voltage. The reference voltage adjustment can only be done with an access to the tunnel.

# System tests during the machine checkout

When the individual system tests, described in the previous sections, have been successfully completed, the integral system should be tested from the CERN Control Centre (CCC), simulating as much as possible future operations with beam.

## Conditions required to perform tests

Successful completion of the individual tests of BLM system sector by sector.

Logging system available (measurement and logging)

Setting data base (LSA) available

BIS operational

LBDS operational

Beam energy distribution operational

## Description of the tests

|  |  | Action | Group(s) Responsible |
| --- | --- | --- | --- |
| 1 | S | Test of beam permit by disabling of signal initiated by MPS for each octant: Generate interlock Measure the delay between detection of “loss over threshold” and the beam dump request | AB/CO AB/BT AB/BI |
| 2 | S | Test of the test sequence by the sequencer: Verify the time allocated in the sequence Change threshold value with expert application in BLETC (do not update data base) and observe the detection of the disagreement between BLETC value and data base | AB/CO AB/BI AB/OP |
| 3 | S | Test the change of the threshold value with the beam energy signal | AB/BI |
| 4 | S | Test of the threshold comparison : Verify the triggering of the test before every fill Verify the triggering of the test after every change of the thresholds table | AB/CO AB/BI |

## Status of the system after the system tests

After these tests, the complete acquisition system is commissioned for each octant.

# Tests with beam.

It is foreseen to make two different tests with beam, one to check the correct functioning of the beam permit disabling by depositing energy in the beam loss detectors and a second to check the correct setting of the threshold to prevent quenches. Below is the general description of the tests given, a more detailed description will be later appended to the document.

## Pilot of 5(109 p+ at 450 Gev

|  |  | Action | Group(s) Responsible |
| --- | --- | --- | --- |
| 1 | S | Test MPS functionnality of the BLM (without quenching): Decrease thresholds to very low value Create local bump to force a beam dump request with very low level of losses Measure delay between signal over threshold and the beam dump : the whole chain from the chamber to the LBDS, via the BIC, is tested | AB/CO AB/BI AB/BT |

## 43 bunches of 4(1010 p+ at 450 Gev (below damage value)

### Conditions required to perform tests

Stable and known optics : needed to determine the impact position and the lost proton density

Orbit corrected to 3 mm

BPM data logging available

### Conditions during the tests

Start with the pilot intensity and if not quenching, increasing step by step (to be defined) the intensity up to the quench level.

### DESCRIPTION of the tests

The test has to be done for one beam only.

|  |  | Action | Group(s) Responsible |
| --- | --- | --- | --- |
| 1 | N | Provoked quench (transient losses) Steer the beam in the chosen magnet Check level of losses and BLM readings at the quench level Repeat for MQM, MQTL, MQ and MB | AB/CO AB/BI |
| 2 | N | Provoked quench (steady state losses) Create a local bump |  |
| 4 | N | Test the direct interlocked BLM in IR6 |  |

# Text boxes

Beam_Info

LBDS

User_Permit

BIS

Beam Energy

From timing system

CTRV

BLM