{"text_format":"5","workbook":"xls","n_sheets":9,"sheets":["14","28","56","84","112","182","252","526","Comparator"],"props":{"creator":"gguaglio","last_modified_by":"gguaglio","company":"CERN"},"not_rendered":{"charts":9,"external_links":1},"unopened":[{"sheet":"Comparator","index":8,"reason":"not_a_worksheet"}],"read":"whole","wrapped":"gz","date_system":"1900","formulas":"cached_only"}

{"sheet":"14","index":0,"n_rows":31,"n_cols":8,"rows_kept":31,"flags":["formulas_cached_only"]}
["Energy   :    60.12 [MeV]"]
[" Run time :       53 [sec]"]
[" Fluence  : 1.042e+10 [p/cm2]"]
[" Dose     : 1.427e+03 [rad]"]
[]
[" Flux     : 1.965e+08 [p/cm2/sec]"]
[" Doserate : 2.692e+01 [rad/sec]"]
[null,null,"before"," during"," after","delta after %","delta during %","delta during"]
[null,"average",{"v":5.123779924666667,"d":"5.12E+00"},{"v":5.31751863625,"d":"5.32E+00"},{"v":5.123431404000001,"d":"5.12E+00"},{"v":-0.006802022565177701,"d":"-6.80E-03"},{"v":3.781167701029564,"d":"3.78"},{"v":0.19373871158333333,"d":"1.94E-01"}]
[null,"stdev",0.0002844910401291463,0.004385894776629373,0.001186518076284454,null,{"v":1441.6635879423027,"d":"1441.66"},0.00439511184653799]
[]
[null,"CH 01"]
[null,"2003-10-05","00:40:00"]
[null,"Time in seconds from 00:00, 1 Jan 1904"]
[null,"Time","Amps","Hz","Sigma [1/s]","Periods [s]"]
[37751.02831601851,3135364846.337,{"v":9.999837e-10,"d":"1.00E-09"},{"v":5.123483726,"d":"5.12E+00"}]
[37751.02854748842,3135364866.336,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.124051048,"d":"5.12E+00"}]
[37751.028663275465,3135364876.34,{"v":9.999824e-10,"d":"1.00E-09"},{"v":5.123805,"d":"5.12E+00"}]
[37751.02877895833,3135364886.335,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.124069342,"d":"5.12E+00"}]
[37751.028894745374,3135364896.339,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.279424706,"d":"5.28E+00"}]
[37751.02901042824,3135364906.334,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.318713404,"d":"5.32E+00"}]
[37751.02913201389,3135364916.839,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.311020931,"d":"5.31E+00"}]
[37751.02925348379,3135364927.334,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.320169006,"d":"5.32E+00"}]
[37751.02936927083,3135364937.338,{"v":9.999852e-10,"d":"1.00E-09"},{"v":5.320171204,"d":"5.32E+00"}]
[37751.029484942126,3135364947.332,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.236994295,"d":"5.24E+00"}]
[37751.02960074074,3135364957.337,{"v":9.999845e-10,"d":"1.00E-09"},{"v":5.123001272,"d":"5.12E+00"}]
[37751.029716412035,3135364967.331,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.123322859,"d":"5.12E+00"}]
[37751.02983221065,3135364977.336,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.124668465,"d":"5.12E+00"}]
[37751.030057881944,3135364996.834,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.122302888,"d":"5.12E+00"}]
[37751.030098449075,3135365000.339,{"v":9.999828e-10,"d":"1.00E-09"}]
[37751.03013890046,3135365003.834,{"v":9.999822e-10,"d":"1.00E-09"}]

{"sheet":"28","index":1,"n_rows":31,"n_cols":8,"rows_kept":31,"flags":["formulas_cached_only"]}
["Energy   :    60.12 [MeV]"]
[" Run time :       56 [sec]"]
[" Fluence  : 1.034e+10 [p/cm2]"]
[" Dose     : 1.416e+03 [rad]"]
[]
[" Flux     : 1.847e+08 [p/cm2/sec]"]
[" Doserate : 2.529e+01 [rad/sec]"]
[null,null,"before"," during"," after","delta after %","delta during %","delta during"]
[null,"average",{"v":5.123624902666666,"d":"5.12E+00"},{"v":5.306221954000001,"d":"5.31E+00"},{"v":5.123403848,"d":"5.12E+00"},{"v":-0.0043144193977107115,"d":"-4.31E-03"},{"v":3.5638255103003225,"d":"3.56"},{"v":0.18259705133333437,"d":"1.83E-01"}]
[null,"stdev",0.00030882205869252636,0.009001569573706741,0.0007975623699827126,null,{"v":2814.8078384740666,"d":"2814.81"},0.009006865484407883]
[]
[null,"CH 01"]
[null,"2003-10-05","00:43:00"]
[null,"Time in seconds from 00:00, 1 Jan 1904"]
[null,"Time","Amps","Hz","Sigma [1/s]","Periods [s]"]
[37751.03035309028,3135365022.34,{"v":9.999818e-10,"d":"1.00E-09"},{"v":5.123980744,"d":"5.12E+00"}]
[37751.03046877315,3135365032.335,{"v":9.999824e-10,"d":"1.00E-09"},{"v":5.123467076,"d":"5.12E+00"}]
[37751.03058456018,3135365042.339,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.123426888,"d":"5.12E+00"}]
[37751.03070023148,3135365052.333,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.207848127,"d":"5.21E+00"}]
[37751.03081603009,3135365062.338,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.30289717,"d":"5.30E+00"}]
[37751.03093170139,3135365072.332,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.297620615,"d":"5.30E+00"}]
[37751.0310475,3135365082.337,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.315576432,"d":"5.32E+00"}]
[37751.031163287036,3135365092.341,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.305468815,"d":"5.31E+00"}]
[37751.031284756944,3135365102.836,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.301379804,"d":"5.30E+00"}]
[37751.03140634259,3135365113.341,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.122210249,"d":"5.12E+00"}]
[37751.03152202546,3135365123.336,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.123096051,"d":"5.12E+00"}]
[37751.03174769676,3135365142.834,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.122985932,"d":"5.12E+00"}]
[37751.03186928241,3135365153.339,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.12459604,"d":"5.12E+00"}]
[37751.03197915509,3135365162.832,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.122937369,"d":"5.12E+00"}]
[37751.03204290509,3135365168.34,{"v":9.999828e-10,"d":"1.00E-09"}]
[37751.032083356484,3135365171.835,{"v":9.999831e-10,"d":"1.00E-09"}]

{"sheet":"56","index":2,"n_rows":38,"n_cols":8,"rows_kept":38,"flags":["formulas_cached_only"]}
["Energy   :    60.12 [MeV]"]
[" Run time :      133 [sec]"]
[" Fluence  : 2.023e+10 [p/cm2]"]
[" Dose     : 2.771e+03 [rad]"]
[]
[" Flux     : 1.521e+08 [p/cm2/sec]"]
[" Doserate : 2.084e+01 [rad/sec]"]
[null,null,"before"," during"," after","delta after %","delta during %","delta during"]
[null,"average",{"v":5.124906576666667,"d":"5.12E+00"},{"v":5.284811093545454,"d":"5.28E+00"},{"v":5.123806162399999,"d":"5.12E+00"},{"v":-0.021471889296043045,"d":"-2.15E-02"},{"v":3.1201450111660867,"d":"3.12"},{"v":0.15990451687878782,"d":"1.60E-01"}]
[null,"stdev",0.0005042714427499753,0.008023957008213794,0.0009000835260761865,null,{"v":1491.1979794961703,"d":"1491.20"},0.008039787046784038]
[]
[null,"CH 01"]
[null,"2003-10-05","00:46:00"]
[null,"Time in seconds from 00:00, 1 Jan 1904"]
[null,"Time","Amps","Hz","Sigma [1/s]","Periods [s]"]
[37751.03226278935,3135365187.338,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.12462561,"d":"5.12E+00"}]
[37751.032378460644,3135365197.332,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.124605378,"d":"5.12E+00"}]
[37751.03249424769,3135365207.336,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.125488742,"d":"5.13E+00"}]
[37751.0326100463,3135365217.341,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.124758667,"d":"5.12E+00"}]
[37751.03272571759,3135365227.335,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.254765086,"d":"5.25E+00"}]
[37751.03284730324,3135365237.84,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.282549444,"d":"5.28E+00"}]
[37751.03296877315,3135365248.335,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.273679933,"d":"5.27E+00"}]
[37751.0330903588,3135365258.84,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.2822539,"d":"5.28E+00"}]
[37751.03321184027,3135365269.336,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.282533888,"d":"5.28E+00"}]
[37751.03343751157,3135365288.834,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.277788505,"d":"5.28E+00"}]
[37751.033559097224,3135365299.339,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.275504156,"d":"5.28E+00"}]
[37751.03367476852,3135365309.333,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.285658918,"d":"5.29E+00"}]
[37751.03379055555,3135365319.337,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.287274599,"d":"5.29E+00"}]
[37751.03390635417,3135365329.342,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.292746805,"d":"5.29E+00"}]
[37751.03402202546,3135365339.336,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.292879422,"d":"5.29E+00"}]
[37751.03413782408,3135365349.341,{"v":9.999823e-10,"d":"1.00E-09"},{"v":5.300052459,"d":"5.30E+00"}]
[37751.03425349537,3135365359.335,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.194257389,"d":"5.19E+00"}]
[37751.034484965276,3135365379.334,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.122838731,"d":"5.12E+00"}]
[37751.03460075231,3135365389.338,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.124848365,"d":"5.12E+00"}]
[37751.03471642361,3135365399.332,{"v":9.999822e-10,"d":"1.00E-09"},{"v":5.122980947,"d":"5.12E+00"}]
[37751.03483222222,3135365409.337,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.124539105,"d":"5.12E+00"}]
[37751.03494800926,3135365419.341,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.123823664,"d":"5.12E+00"}]
[37751.03498846065,3135365422.836,{"v":9.999828e-10,"d":"1.00E-09"}]

{"sheet":"84","index":3,"n_rows":38,"n_cols":8,"rows_kept":38,"flags":["formulas_cached_only"]}
["Energy   :    60.12 [MeV]"]
[" Run time :      135 [sec]"]
[" Fluence  : 2.016e+10 [p/cm2]"]
[" Dose     : 2.762e+03 [rad]"]
[]
[" Flux     : 1.494e+08 [p/cm2/sec]"]
[" Doserate : 2.046e+01 [rad/sec]"]
[null,null,"before"," during"," after","delta after %","delta during %","delta during"]
[null,"average",{"v":5.123236381666666,"d":"5.12E+00"},{"v":5.277200264666667,"d":"5.28E+00"},{"v":5.123467328749999,"d":"5.12E+00"},{"v":0.004507835792244563,"d":"4.51E-03"},{"v":3.0052074807821905,"d":"3.01"},{"v":0.1539638830000012,"d":"1.54E-01"}]
[null,"stdev",0.0005636420065687638,0.003526122614072227,0.0015640592850145229,null,{"v":525.5961360186598,"d":"525.60"},0.003570886864777491]
[]
[null,"CH 01"]
[null,"2003-10-05","00:50:00"]
[null,"Time in seconds from 00:00, 1 Jan 1904"]
[null,"Time","Amps","Hz","Sigma [1/s]","Periods [s]"]
[37751.03522572917,3135365443.336,{"v":9.999824e-10,"d":"1.00E-09"},{"v":5.123351574,"d":"5.12E+00"}]
[37751.035445601854,3135365462.333,{"v":9.999823e-10,"d":"1.00E-09"},{"v":5.122624042,"d":"5.12E+00"}]
[37751.03556138889,3135365472.337,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.123733529,"d":"5.12E+00"}]
[37751.035677071755,3135365482.332,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.217076686,"d":"5.22E+00"}]
[37751.0357986574,3135365492.837,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.279541766,"d":"5.28E+00"}]
[37751.03592012731,3135365503.332,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.279769738,"d":"5.28E+00"}]
[37751.03604171296,3135365513.837,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.281601945,"d":"5.28E+00"}]
[37751.036163182864,3135365524.332,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.272857646,"d":"5.27E+00"}]
[37751.03627898148,3135365534.337,{"v":9.999823e-10,"d":"1.00E-09"},{"v":5.270898445,"d":"5.27E+00"}]
[37751.03639476852,3135365544.341,{"v":9.999823e-10,"d":"1.00E-09"},{"v":5.272213262,"d":"5.27E+00"}]
[37751.036510439815,3135365554.335,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.27928248,"d":"5.28E+00"}]
[37751.03663202546,3135365564.84,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.277944743,"d":"5.28E+00"}]
[37751.036753506945,3135365575.336,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.276465876,"d":"5.28E+00"}]
[37751.036874976846,3135365585.831,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.276783601,"d":"5.28E+00"}]
[37751.0369965625,3135365596.336,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.281177988,"d":"5.28E+00"}]
[37751.03711803241,3135365606.831,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.277865686,"d":"5.28E+00"}]
[37751.037239618054,3135365617.336,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.210786939,"d":"5.21E+00"}]
[37751.03745949074,3135365636.333,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.122626425,"d":"5.12E+00"}]
[37751.03756949074,3135365645.837,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.125638908,"d":"5.13E+00"}]
[37751.03769096064,3135365656.332,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.123519842,"d":"5.12E+00"}]
[37751.037812546296,3135365666.837,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.12208414,"d":"5.12E+00"}]
[37751.037876180555,3135365672.335,{"v":9.999832e-10,"d":"1.00E-09"}]
[37751.037916747686,3135365675.84,{"v":9.99983e-10,"d":"1.00E-09"}]

{"sheet":"112","index":4,"n_rows":37,"n_cols":8,"rows_kept":37,"flags":["formulas_cached_only"]}
["Energy   :    60.12 [MeV]"]
[" Run time :      123 [sec]"]
[" Fluence  : 2.015e+10 [p/cm2]"]
[" Dose     : 2.761e+03 [rad]"]
[]
[" Flux     : 1.639e+08 [p/cm2/sec]"]
[" Doserate : 2.245e+01 [rad/sec]"]
[null,null,"before"," during"," after","delta after %","delta during %","delta during"]
[null,"average",{"v":5.1233690557500005,"d":"5.12E+00"},{"v":5.2897994836999995,"d":"5.29E+00"},{"v":5.1235561480000005,"d":"5.12E+00"},{"v":0.0036517425928916936,"d":"3.65E-03"},{"v":3.2484567506065654,"d":"3.25"},{"v":0.16643042794999907,"d":"1.66E-01"}]
[null,"stdev",0.0013192422532215453,0.0033797819902243145,0.000831458052566345,null,{"v":156.19115685318602,"d":"156.19"},0.00362812987972174]
[]
[null,"CH 01"]
[null,"2003-10-05","00:54:00"]
[null,"Time in seconds from 00:00, 1 Jan 1904"]
[null,"Time","Amps","Hz","Sigma [1/s]","Periods [s]"]
[37751.03806140046,3135365688.338,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.122621728,"d":"5.12E+00"}]
[37751.03817707176,3135365698.332,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.125138324,"d":"5.13E+00"}]
[37751.03829287037,3135365708.337,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.123568793,"d":"5.12E+00"}]
[37751.038408541666,3135365718.331,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.122147378,"d":"5.12E+00"}]
[37751.03852434028,3135365728.336,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.215111875,"d":"5.22E+00"}]
[37751.038645810186,3135365738.831,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.294460558,"d":"5.29E+00"}]
[37751.03876739583,3135365749.336,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.285409796,"d":"5.29E+00"}]
[37751.03888886574,3135365759.831,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.287861855,"d":"5.29E+00"}]
[37751.039010451386,3135365770.336,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.292116307,"d":"5.29E+00"}]
[37751.039131921294,3135365780.831,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.291004612,"d":"5.29E+00"}]
[37751.03925350694,3135365791.336,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.287735003,"d":"5.29E+00"}]
[37751.03937497685,3135365801.831,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.293222017,"d":"5.29E+00"}]
[37751.0394965625,3135365812.336,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.287101297,"d":"5.29E+00"}]
[37751.0396180324,3135365822.831,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.285810207,"d":"5.29E+00"}]
[37751.03985530092,3135365843.331,{"v":9.999824e-10,"d":"1.00E-09"},{"v":5.293273185,"d":"5.29E+00"}]
[37751.039971087965,3135365853.335,{"v":9.999824e-10,"d":"1.00E-09"},{"v":5.170327991,"d":"5.17E+00"}]
[37751.040086886576,3135365863.34,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.124475121,"d":"5.12E+00"}]
[37751.04020255787,3135365873.334,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.123380012,"d":"5.12E+00"}]
[37751.0403183449,3135365883.338,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.123863894,"d":"5.12E+00"}]
[37751.04043402778,3135365893.333,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.122505565,"d":"5.12E+00"}]
[37751.0405150463,3135365900.333,{"v":9.99983e-10,"d":"1.00E-09"}]
[37751.04055561342,3135365903.838,{"v":9.999825e-10,"d":"1.00E-09"}]

{"sheet":"182","index":5,"n_rows":37,"n_cols":24,"rows_kept":37,"flags":["formulas_cached_only"]}
["Energy   :    60.12 [MeV]"]
[" Run time :      243 [sec]"]
[" Fluence  : 5.040e+10 [p/cm2]"]
[" Dose     : 6.904e+03 [rad]"]
[]
[" Flux     : 2.074e+08 [p/cm2/sec]"]
[" Doserate : 2.841e+01 [rad/sec]"]
[null,null,"before"," during"," after","delta after %","delta during %","delta during",{"$nulls":10},"before"," during"," after","delta after %","delta during %","delta during"]
[null,"average",{"v":5.127109862666667,"d":"5.13E+00"},{"v":5.323235900777778,"d":"5.32E+00"},{"v":5.121352562999999,"d":"5.12E+00"},{"v":-0.11229132631991323,"d":"-1.12E-01"},{"v":3.8252747330267622,"d":"3.83"},{"v":0.19612603811111118,"d":"1.96E-01"},{"$nulls":9},"average",{"v":5.028612418000001,"d":"5.03E+00"},{"v":6.431486242777777,"d":"6.43E+00"},{"v":5.03431242525,"d":"5.03E+00"},{"v":0.11335149294058322,"d":"1.13E-01"},{"v":27.89783161168211,"d":"27.90"},{"v":1.4028738247777763,"d":"1.40E+00"}]
[null,"stdev",0.006312490206900431,0.006413579160226647,0.0010759768202149467,null,{"v":1.6014116460048022,"d":"1.60"},0.008998973844650699,{"$nulls":9},"stdev",0.0011359918913630743,0.05581159620923732,0.0033155634970135272,null,{"v":4813.027692677375,"d":"4813.03"},0.05582315602866069]
[]
[null,"CH 01",{"$nulls":15},"CH 03"]
[null,"2003-10-05","00:58:00",{"$nulls":14},"2003-10-05","00:58:00"]
[null,"Time in seconds from 00:00, 1 Jan 1904",{"$nulls":15},"Time in seconds from 00:00, 1 Jan 1904"]
[null,"Time","Amps","Hz","Sigma [1/s]","Periods [s]",{"$nulls":11},"Time","Amps","Hz","Sigma [1/s]","Periods [s]"]
[37751.0409259375,3135365935.834,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.134385188,"d":"5.13E+00"},{"$nulls":12},37751.04104752315,3135365946.339,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.02904841,"d":"5.03E+00"}]
[37751.041163194444,3135365956.333,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.123859193,"d":"5.12E+00"},{"$nulls":12},37751.041278993056,3135365966.338,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.027323016,"d":"5.03E+00"}]
[37751.041504664354,3135365985.836,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.123085207,"d":"5.12E+00"},{"$nulls":12},37751.041736134255,3135366005.835,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.029465828,"d":"5.03E+00"}]
[37751.04185771991,3135366016.34,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.123235348,"d":"5.12E+00"},{"$nulls":12},37751.04197918982,3135366026.835,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.031369105,"d":"5.03E+00"}]
[37751.04221064815,3135366046.833,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.308151057,"d":"5.31E+00"},{"$nulls":12},37751.04232644675,3135366056.838,{"v":9.999828e-10,"d":"1.00E-09"},{"v":6.364712503,"d":"6.36E+00"}]
[37751.04244791667,3135366067.333,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.318773495,"d":"5.32E+00"},{"$nulls":12},37751.042569502315,3135366077.838,{"v":9.999824e-10,"d":"1.00E-09"},{"v":6.397571599,"d":"6.40E+00"}]
[37751.04269097222,3135366088.333,{"v":9.999839e-10,"d":"1.00E-09"},{"v":5.316727496,"d":"5.32E+00"},{"$nulls":12},37751.04280677083,3135366098.338,{"v":9.99983e-10,"d":"1.00E-09"},{"v":6.321810486,"d":"6.32E+00"}]
[37751.042922442124,3135366108.332,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.312129212,"d":"5.31E+00"},{"$nulls":12},37751.04303822917,3135366118.336,{"v":9.999828e-10,"d":"1.00E-09"},{"v":6.371757218,"d":"6.37E+00"}]
[37751.04325811342,3135366137.334,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.324247445,"d":"5.32E+00"},{"$nulls":12},37751.04337390046,3135366147.338,{"v":9.999825e-10,"d":"1.00E-09"},{"v":6.471188324,"d":"6.47E+00"}]
[37751.043489571755,3135366157.332,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.322784342,"d":"5.32E+00"},{"$nulls":12},37751.04360537037,3135366167.337,{"v":9.99983e-10,"d":"1.00E-09"},{"v":6.484717733,"d":"6.48E+00"}]
[37751.04372115741,3135366177.341,{"v":9.999851e-10,"d":"1.00E-09"},{"v":5.330953841,"d":"5.33E+00"},{"$nulls":12},37751.043836840276,3135366187.336,{"v":9.999829e-10,"d":"1.00E-09"},{"v":6.451359305,"d":"6.45E+00"}]
[37751.04406829861,3135366207.334,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.328327202,"d":"5.33E+00"},{"$nulls":12},37751.04418409723,3135366217.339,{"v":9.999828e-10,"d":"1.00E-09"},{"v":6.473294477,"d":"6.47E+00"}]
[37751.04429976852,3135366227.333,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.33065671,"d":"5.33E+00"},{"$nulls":12},37751.04441555555,3135366237.337,{"v":9.999824e-10,"d":"1.00E-09"},{"v":6.470572195,"d":"6.47E+00"}]
[37751.044531238425,3135366247.332,{"v":9.999844e-10,"d":"1.00E-09"},{"v":5.324523364,"d":"5.32E+00"},{"$nulls":12},37751.04464702546,3135366257.336,{"v":9.999834e-10,"d":"1.00E-09"},{"v":6.441104848,"d":"6.44E+00"}]
[37751.04476282407,3135366267.341,{"v":9.999848e-10,"d":"1.00E-09"},{"v":5.323814841,"d":"5.32E+00"},{"$nulls":12},37751.04487849537,3135366277.335,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.464070632,"d":"5.46E+00"}]
[37751.04499428241,3135366287.339,{"v":9.999822e-10,"d":"1.00E-09"},{"v":5.119938696,"d":"5.12E+00"},{"$nulls":12},37751.04510996528,3135366297.334,{"v":9.999824e-10,"d":"1.00E-09"},{"v":5.030641388,"d":"5.03E+00"}]
[37751.045225752314,3135366307.338,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.121111556,"d":"5.12E+00"},{"$nulls":12},37751.04534722222,3135366317.833,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.030402702,"d":"5.03E+00"}]
[37751.04546880787,3135366328.338,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.121113026,"d":"5.12E+00"},{"$nulls":12},37751.045590277776,3135366338.833,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.033498282,"d":"5.03E+00"}]
[37751.04571186342,3135366349.338,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.1203196,"d":"5.12E+00"},{"$nulls":12},37751.045833344906,3135366359.834,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.038387931,"d":"5.04E+00"}]
[37751.04595493056,3135366370.339,{"v":9.999851e-10,"d":"1.00E-09"},{"v":5.12286607,"d":"5.12E+00"},{"$nulls":12},37751.04607640046,3135366380.834,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.034960786,"d":"5.03E+00"}]
[37751.04616900463,3135366388.835,{"v":9.999828e-10,"d":"1.00E-09"},{"$nulls":13},37751.04620957176,3135366392.34,{"v":9.99983e-10,"d":"1.00E-09"}]
[37751.046250023144,3135366395.835,{"v":9.999825e-10,"d":"1.00E-09"}]

{"sheet":"252","index":6,"n_rows":29,"n_cols":32,"rows_kept":29,"flags":["formulas_cached_only"]}
["Energy   :    60.12 [MeV]"]
[" Run time :      250 [sec]"]
[" Fluence  : 5.030e+10 [p/cm2]"]
[" Dose     : 6.891e+03 [rad]"]
[]
[" Flux     : 2.012e+08 [p/cm2/sec]"]
[" Doserate : 2.756e+01 [rad/sec]"]
[null,null,"before"," during"," after","delta after %","delta during %","delta during",null,null,"before"," during"," after","delta after %","delta during %","delta during",null,null,"before"," during"," after","delta after %","delta during %","delta during",null,null,"before"," during"," after","delta after %","delta during %","delta during"]
[null,"average",{"v":5.122550945333333,"d":"5.12E+00"},{"v":5.316776041000001,"d":"5.32E+00"},{"v":5.121333879,"d":"5.12E+00"},{"v":-0.023758989345767922,"d":"-2.38E-02"},{"v":3.791569820181251,"d":"3.79"},{"v":0.1942250956666678,"d":"1.94E-01"},null,"average",{"v":5.1115469393333335,"d":"5.11E+00"},{"v":5.19293403225,"d":"5.19E+00"},{"v":5.10398528,"d":"5.10E+00"},{"v":-0.1479328943484126,"d":"-1.48E-01"},{"v":1.5922203959508385,"d":"1.59"},{"v":0.08138709291666668,"d":"8.14E-02"},null,"average",{"v":5.0307710173333335,"d":"5.03E+00"},{"v":6.3621640645,"d":"6.36E+00"},{"v":5.032464771,"d":"5.03E+00"},{"v":0.03366787438408014,"d":"3.37E-02"},{"v":26.464990010068057,"d":"26.46"},{"v":1.3313930471666664,"d":"1.33E+00"},null,"average",{"v":5.082685480333333,"d":"5.08E+00"},{"v":5.329225496000001,"d":"5.33E+00"},{"v":5.076845001,"d":"5.08E+00"},{"v":-0.11490932019956324,"d":"-1.15E-01"},{"v":4.8505857114436335,"d":"4.85"},{"v":0.2465400156666684,"d":"2.47E-01"}]
[null,"stdev",0.001086826294117148,0.009474873433046885,0.0020746332105906337,null,{"v":771.7928048238404,"d":"771.79"},0.009537002567150859,null,"stdev",0.004305769157652955,0.0066127465633661325,0.0006919633956909159,null,{"v":53.57875262803673,"d":"53.58"},0.00789101166836712,null,"stdev",0.0017168114158474352,0.059582693477966985,0.002656366067933394,null,{"v":3370.5438773283304,"d":"3370.54"},0.05960742238620081,null,"stdev",0.0011639953972651224,0.011769814621568607,0.0028013203177664186,null,{"v":911.156457252537,"d":"911.16"},0.011827232199925077]
[]
[null,"CH 01",{"$nulls":7},"CH 02",{"$nulls":7},"CH 03",{"$nulls":7},"CH 04"]
[null,"2003-10-05","01:06:00",{"$nulls":6},"2003-10-05","01:06:00",{"$nulls":6},"2003-10-05","01:06:00",{"$nulls":6},"2003-10-05","01:06:00"]
[null,"Time in seconds from 00:00, 1 Jan 1904",{"$nulls":7},"Time in seconds from 00:00, 1 Jan 1904",{"$nulls":7},"Time in seconds from 00:00, 1 Jan 1904",{"$nulls":7},"Time in seconds from 00:00, 1 Jan 1904"]
[null,"Time","Amps","Hz","Sigma [1/s]","Periods [s]",{"$nulls":3},"Time","Amps","Hz","Sigma [1/s]","Periods [s]",{"$nulls":3},"Time","Amps","Hz","Sigma [1/s]","Periods [s]",{"$nulls":3},"Time","Amps","Hz","Sigma [1/s]","Periods [s]"]
[37751.046556712965,3135366422.333,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.122188515,"d":"5.12E+00"},{"$nulls":4},37751.04666671296,3135366431.837,{"v":9.999823e-10,"d":"1.00E-09"},{"v":5.109862856,"d":"5.11E+00"},{"$nulls":4},37751.046788182866,3135366442.332,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.028889509,"d":"5.03E+00"},{"$nulls":4},37751.04690398148,3135366452.337,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.081353937,"d":"5.08E+00"}]
[37751.047118055554,3135366470.833,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.121691644,"d":"5.12E+00"},{"$nulls":4},37751.0472396412,3135366481.338,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.116440221,"d":"5.12E+00"},{"$nulls":4},37751.047361111116,3135366491.833,{"v":9.999824e-10,"d":"1.00E-09"},{"v":5.031171044,"d":"5.03E+00"},{"$nulls":4},37751.047482696755,3135366502.338,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.083509775,"d":"5.08E+00"}]
[37751.04760417824,3135366512.834,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.123772677,"d":"5.12E+00"},{"$nulls":4},37751.04772576389,3135366523.339,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.108337741,"d":"5.11E+00"},{"$nulls":4},37751.0478472338,3135366533.834,{"v":9.999824e-10,"d":"1.00E-09"},{"v":5.032252499,"d":"5.03E+00"},{"$nulls":4},37751.04796881945,3135366544.339,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.083192729,"d":"5.08E+00"}]
[37751.04808449074,3135366554.333,{"v":9.999838e-10,"d":"1.00E-09"},{"v":5.339109046,"d":"5.34E+00"},{"$nulls":4},37751.04820028935,3135366564.338,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.201927286,"d":"5.20E+00"},{"$nulls":4},37751.048315960645,3135366574.332,{"v":9.999832e-10,"d":"1.00E-09"},{"v":6.4728149,"d":"6.47E+00"},{"$nulls":4},37751.04843174768,3135366584.336,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.346797471,"d":"5.35E+00"}]
[37751.0485475463,3135366594.341,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.321510492,"d":"5.32E+00"},{"$nulls":4},37751.04866321759,3135366604.335,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.1961538,"d":"5.20E+00"},{"$nulls":4},37751.04878480324,3135366614.84,{"v":9.999828e-10,"d":"1.00E-09"},{"v":6.402830935,"d":"6.40E+00"},{"$nulls":4},37751.04890627315,3135366625.335,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.333518083,"d":"5.33E+00"}]
[37751.04913774306,3135366645.334,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.302590698,"d":"5.30E+00"},{"$nulls":4},37751.04925353009,3135366655.338,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.195266519,"d":"5.20E+00"},{"$nulls":4},37751.04936921297,3135366665.333,{"v":9.999824e-10,"d":"1.00E-09"},{"v":6.338491817,"d":"6.34E+00"},{"$nulls":4},37751.049484999996,3135366675.337,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.330131799,"d":"5.33E+00"}]
[37751.04960068287,3135366685.332,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.320786988,"d":"5.32E+00"},{"$nulls":4},37751.049716469905,3135366695.336,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.197227703,"d":"5.20E+00"},{"$nulls":4},37751.04983225695,3135366705.34,{"v":9.999823e-10,"d":"1.00E-09"},{"v":6.418055082,"d":"6.42E+00"},{"$nulls":4},37751.04994793981,3135366715.335,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.340461218,"d":"5.34E+00"}]
[37751.050063726856,3135366725.339,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.322215986,"d":"5.32E+00"},{"$nulls":4},37751.05017940972,3135366735.334,{"v":9.999824e-10,"d":"1.00E-09"},{"v":5.183088107,"d":"5.18E+00"},{"$nulls":4},37751.05029554398,3135366745.368,{"v":9.999831e-10,"d":"1.00E-09"},{"v":6.289278424,"d":"6.29E+00"},{"$nulls":4},37751.050410868054,3135366755.332,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.312790884,"d":"5.31E+00"}]
[37751.050526666666,3135366765.337,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.29999695,"d":"5.30E+00"},{"$nulls":4},37751.050752337964,3135366784.835,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.189221619,"d":"5.19E+00"},{"$nulls":4},37751.050983796296,3135366804.833,{"v":9.99982e-10,"d":"1.00E-09"},{"v":5.177971327,"d":"5.18E+00"},{"$nulls":4},37751.05110539352,3135366815.339,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.079196666,"d":"5.08E+00"}]
[37751.0512152662,3135366824.832,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.118938299,"d":"5.12E+00"},{"$nulls":4},37751.05133685185,3135366835.337,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.103186435,"d":"5.10E+00"},{"$nulls":4},37751.05145832176,3135366845.832,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.035298059,"d":"5.04E+00"},{"$nulls":4},37751.05157991898,3135366856.338,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.079976638,"d":"5.08E+00"}]
[37751.05170138889,3135366866.833,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.122532212,"d":"5.12E+00"},{"$nulls":4},37751.051822974536,3135366877.338,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.104370619,"d":"5.10E+00"},{"$nulls":4},37751.05194444444,3135366887.833,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.032065791,"d":"5.03E+00"},{"$nulls":4},37751.05206603009,3135366898.338,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.075980633,"d":"5.08E+00"}]
[37751.05218170139,3135366908.332,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.122531126,"d":"5.12E+00"},{"$nulls":4},37751.0522975,3135366918.337,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.104398786,"d":"5.10E+00"},{"$nulls":4},37751.05252895833,3135366938.335,{"v":9.999823e-10,"d":"1.00E-09"},{"v":5.030030463,"d":"5.03E+00"},{"$nulls":4},37751.05275462963,3135366957.833,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.074577732,"d":"5.07E+00"}]
[37751.05280099537,3135366961.839,{"v":9.999828e-10,"d":"1.00E-09"},{"$nulls":5},37751.05284144676,3135366965.334,{"v":9.99983e-10,"d":"1.00E-09"}]
[37751.05288201389,3135366968.839,{"v":9.999828e-10,"d":"1.00E-09"}]

{"sheet":"526","index":7,"n_rows":49,"n_cols":32,"rows_kept":49,"flags":["formulas_cached_only"]}
["Energy   :    60.12 [MeV]"]
[" Run time :     1033 [sec]"]
[" Fluence  : 2.002e+11 [p/cm2]"]
[" Dose     : 2.742e+04 [rad]"]
[]
[" Flux     : 1.938e+08 [p/cm2/sec]"]
[" Doserate : 2.654e+01 [rad/sec]"]
[null,null,"before"," during"," after","delta after %","delta during %","delta during",null,null,"before"," during"," after","delta after %","delta during %","delta during",null,null,"before"," during"," after","delta after %","delta during %","delta during",null,null,"before"," during"," after","delta after %","delta during %","delta during"]
[null,"average",{"v":5.122316655,"d":"5.12E+00"},{"v":5.317145976611111,"d":"5.32E+00"},{"v":5.18252781875,"d":"5.18E+00"},{"v":1.1754674262714149,"d":"1.18E+00"},{"v":3.803539193949179,"d":"3.80"},{"v":0.19482932161111144,"d":"1.95E-01"},null,"average",{"v":5.108242007666667,"d":"5.11E+00"},{"v":5.194193749235294,"d":"5.19E+00"},{"v":5.104795513,"d":"5.10E+00"},{"v":-0.06746929102995347,"d":"-6.75E-02"},{"v":1.6826090353516232,"d":"1.68"},{"v":0.08595174156862662,"d":"8.60E-02"},null,"average",{"v":5.040296041333333,"d":"5.04E+00"},{"v":6.3059579322222215,"d":"6.31E+00"},{"v":5.030326832750001,"d":"5.03E+00"},{"v":-0.19779013973739712,"d":"-1.98E-01"},{"v":25.11086413396617,"d":"25.11"},{"v":1.2656618908888886,"d":"1.27E+00"},null,"average",{"v":5.087696080000001,"d":"5.09E+00"},{"v":5.333599802666667,"d":"5.33E+00"},{"v":5.0576960125,"d":"5.06E+00"},{"v":-0.5896591900984904,"d":"-5.90E-01"},{"v":4.833302123397787,"d":"4.83"},{"v":0.24590372266666627,"d":"2.46E-01"}]
[null,"stdev",0.0025923529597754653,0.013156801432134758,0.11466757423425056,null,{"v":407.5235369675248,"d":"407.52"},0.013409762033409833,null,"stdev",0.005285626137140087,0.007299050430271346,0.0021110634842182894,null,{"v":38.09244620961159,"d":"38.09"},0.009011879983957995,null,"stdev",0.004649117211282236,0.10473988043442305,0.002808528664311868,null,{"v":2152.898253032764,"d":"2152.90"},0.10484301046927963,null,"stdev",0.0026576478748405676,0.05254025947664397,0.06282726638549524,null,{"v":1876.9458540400453,"d":"1876.95"},0.05260743253666464]
[]
[null,"CH 01",{"$nulls":7},"CH 02",{"$nulls":7},"CH 03",{"$nulls":7},"CH 04"]
[null,"2003-10-05","01:31:00",{"$nulls":6},"2003-10-05","01:31:00",{"$nulls":6},"2003-10-05","01:31:00",{"$nulls":6},"2003-10-05","01:31:00"]
[null,"Time in seconds from 00:00, 1 Jan 1904",{"$nulls":7},"Time in seconds from 00:00, 1 Jan 1904",{"$nulls":7},"Time in seconds from 00:00, 1 Jan 1904",{"$nulls":7},"Time in seconds from 00:00, 1 Jan 1904"]
[null,"Time","Amps","Hz","Sigma [1/s]","Periods [s]",{"$nulls":3},"Time","Amps","Hz","Sigma [1/s]","Periods [s]",{"$nulls":3},"Time","Amps","Hz","Sigma [1/s]","Periods [s]",{"$nulls":3},"Time","Amps","Hz","Sigma [1/s]","Periods [s]"]
[37751.06353587963,3135367889.333,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.124383081,"d":"5.12E+00"},{"$nulls":4},37751.06364587963,3135367898.837,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.111605191,"d":"5.11E+00"},{"$nulls":4},37751.06387733796,3135367918.835,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.035438104,"d":"5.04E+00"},{"$nulls":4},37751.06410880787,3135367938.834,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.085469898,"d":"5.09E+00"}]
[37751.06434027778,3135367958.833,{"v":9.999848e-10,"d":"1.00E-09"},{"v":5.119407887,"d":"5.12E+00"},{"$nulls":4},37751.064461863425,3135367969.338,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.10214968,"d":"5.10E+00"},{"$nulls":4},37751.06458333333,3135367979.833,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.044703551,"d":"5.04E+00"},{"$nulls":4},37751.06470491898,3135367990.338,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.090638427,"d":"5.09E+00"}]
[37751.064820590276,3135368000.332,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.123158997,"d":"5.12E+00"},{"$nulls":4},37751.06493059028,3135368009.836,{"v":9.999838e-10,"d":"1.00E-09"},{"v":5.110971152,"d":"5.11E+00"},{"$nulls":4},37751.065052060185,3135368020.331,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.040746469,"d":"5.04E+00"},{"$nulls":4},37751.06517364583,3135368030.836,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.086979915,"d":"5.09E+00"}]
[37751.06529511574,3135368041.331,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.123804388,"d":"5.12E+00"},{"$nulls":4},37751.06541091435,3135368051.336,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.110198626,"d":"5.11E+00"},{"$nulls":4},37751.065526701386,3135368061.34,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.041909666,"d":"5.04E+00"},{"$nulls":4},37751.06564238426,3135368071.335,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.081737364,"d":"5.08E+00"}]
[37751.065758171295,3135368081.339,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.123123133,"d":"5.12E+00"},{"$nulls":4},37751.06586805556,3135368090.833,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.110899004,"d":"5.11E+00"},{"$nulls":4},37751.0659896412,3135368101.338,{"v":9.99982e-10,"d":"1.00E-09"},{"v":5.037194559,"d":"5.04E+00"},{"$nulls":4},37751.06611111111,3135368111.833,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.08926278,"d":"5.09E+00"}]
[37751.06623269676,3135368122.338,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.338848073,"d":"5.34E+00"},{"$nulls":4},37751.066348368055,3135368132.332,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.207711753,"d":"5.21E+00"},{"$nulls":4},37751.066464166666,3135368142.337,{"v":9.999828e-10,"d":"1.00E-09"},{"v":6.476188836,"d":"6.48E+00"},{"$nulls":4},37751.066689837964,3135368161.835,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.345597346,"d":"5.35E+00"}]
[37751.06681142361,3135368172.34,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.324046853,"d":"5.32E+00"},{"$nulls":4},37751.06692709491,3135368182.334,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.19675246,"d":"5.20E+00"},{"$nulls":4},37751.06704289352,3135368192.339,{"v":9.999822e-10,"d":"1.00E-09"},{"v":6.368499128,"d":"6.37E+00"},{"$nulls":4},37751.067158564816,3135368202.333,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.337256083,"d":"5.34E+00"}]
[37751.06727435185,3135368212.337,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.323711501,"d":"5.32E+00"},{"$nulls":4},37751.06739003472,3135368222.332,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.202200572,"d":"5.20E+00"},{"$nulls":4},37751.06750582176,3135368232.336,{"v":9.999827e-10,"d":"1.00E-09"},{"v":6.455687581,"d":"6.46E+00"},{"$nulls":4},37751.06773149306,3135368251.834,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.343931871,"d":"5.34E+00"}]
[37751.06796296297,3135368271.833,{"v":9.999823e-10,"d":"1.00E-09"},{"v":5.322890662,"d":"5.32E+00"},{"$nulls":4},37751.068084548606,3135368282.338,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.19818698,"d":"5.20E+00"},{"$nulls":4},37751.0682002199,3135368292.332,{"v":9.999832e-10,"d":"1.00E-09"},{"v":6.418061752,"d":"6.42E+00"},{"$nulls":4},37751.068316018514,3135368302.337,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.346919384,"d":"5.35E+00"}]
[37751.06843748842,3135368312.832,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.325868314,"d":"5.33E+00"},{"$nulls":4},37751.06855907407,3135368323.337,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.195350979,"d":"5.20E+00"},{"$nulls":4},37751.068674745366,3135368333.331,{"v":9.999828e-10,"d":"1.00E-09"},{"v":6.369145503,"d":"6.37E+00"},{"$nulls":4},37751.06879054398,3135368343.336,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.320965134,"d":"5.32E+00"}]
[37751.069016215275,3135368362.834,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.319404836,"d":"5.32E+00"},{"$nulls":4},37751.06913780093,3135368373.339,{"v":9.999835e-10,"d":"1.00E-09"},{"v":5.204241035,"d":"5.20E+00"},{"$nulls":4},37751.069253472226,3135368383.333,{"v":9.999839e-10,"d":"1.00E-09"},{"v":6.416177871,"d":"6.42E+00"},{"$nulls":4},37751.06936927083,3135368393.338,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.337604236,"d":"5.34E+00"}]
[37751.06948494213,3135368403.332,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.327055335,"d":"5.33E+00"},{"$nulls":4},37751.06960072916,3135368413.336,{"v":9.999827e-10,"d":"1.00E-09"},{"v":5.203114479,"d":"5.20E+00"},{"$nulls":4},37751.06972219907,3135368423.831,{"v":9.999828e-10,"d":"1.00E-09"},{"v":6.391575114,"d":"6.39E+00"},{"$nulls":4},37751.06984379629,3135368434.337,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.33813386,"d":"5.34E+00"}]
[37751.06995946759,3135368444.331,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.352231362,"d":"5.35E+00"},{"$nulls":4},37751.070075254625,3135368454.335,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.200330706,"d":"5.20E+00"},{"$nulls":4},37751.07030092593,3135368473.833,{"v":9.999832e-10,"d":"1.00E-09"},{"v":6.380929554,"d":"6.38E+00"},{"$nulls":4},37751.07042251157,3135368484.338,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.336741618,"d":"5.34E+00"}]
[37751.07053819444,3135368494.333,{"v":9.999825e-10,"d":"1.00E-09"},{"v":5.313473305,"d":"5.31E+00"},{"$nulls":4},37751.07065978009,3135368504.838,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.197992334,"d":"5.20E+00"},{"$nulls":4},37751.070781250004,3135368515.333,{"v":9.999834e-10,"d":"1.00E-09"},{"v":6.407635526,"d":"6.41E+00"},{"$nulls":4},37751.07090283564,3135368525.838,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.347280245,"d":"5.35E+00"}]
[37751.07102430556,3135368536.333,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.329357575,"d":"5.33E+00"},{"$nulls":4},37751.0711458912,3135368546.838,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.196504291,"d":"5.20E+00"},{"$nulls":4},37751.07126736111,3135368557.333,{"v":9.999827e-10,"d":"1.00E-09"},{"v":6.376411193,"d":"6.38E+00"},{"$nulls":4},37751.07138315972,3135368567.338,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.335948732,"d":"5.34E+00"}]
[37751.071498831014,3135368577.332,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.303610113,"d":"5.30E+00"},{"$nulls":4},37751.071614629625,3135368587.337,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.183289638,"d":"5.18E+00"},{"$nulls":4},37751.07173030092,3135368597.331,{"v":9.999837e-10,"d":"1.00E-09"},{"v":6.119574061,"d":"6.12E+00"},{"$nulls":4},37751.071846087965,3135368607.335,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.530024592,"d":"5.53E+00"}]
[37751.072077557874,3135368627.334,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.300830402,"d":"5.30E+00"},{"$nulls":4},37751.0721933449,3135368637.338,{"v":9.999823e-10,"d":"1.00E-09"},{"v":5.187168889,"d":"5.19E+00"},{"$nulls":4},37751.07230902778,3135368647.333,{"v":9.999825e-10,"d":"1.00E-09"},{"v":6.204005995,"d":"6.20E+00"},{"$nulls":4},37751.07242481481,3135368657.337,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.314077725,"d":"5.31E+00"}]
[37751.07265048611,3135368676.835,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.308997364,"d":"5.31E+00"},{"$nulls":4},37751.07298615741,3135368705.837,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.189994812,"d":"5.19E+00"},{"$nulls":4},37751.073217627316,3135368725.836,{"v":9.999832e-10,"d":"1.00E-09"},{"v":6.279161404,"d":"6.28E+00"},{"$nulls":4},37751.07333921296,3135368736.341,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.303370695,"d":"5.30E+00"}]
[37751.07345488426,3135368746.335,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.308363362,"d":"5.31E+00"},{"$nulls":4},37751.073576469906,3135368756.84,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.18737878,"d":"5.19E+00"},{"$nulls":4},37751.07369793981,3135368767.335,{"v":9.999831e-10,"d":"1.00E-09"},{"v":6.211443549,"d":"6.21E+00"},{"$nulls":4},37751.073813738425,3135368777.34,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.31076858,"d":"5.31E+00"}]
[37751.07392940972,3135368787.334,{"v":9.999853e-10,"d":"1.00E-09"},{"v":5.317080894,"d":"5.32E+00"},{"$nulls":4},37751.074045208334,3135368797.339,{"v":9.999835e-10,"d":"1.00E-09"},{"$nulls":5},37751.07416667824,3135368807.834,{"v":9.999827e-10,"d":"1.00E-09"},{"v":6.25732413,"d":"6.26E+00"},{"$nulls":4},37751.07428826389,3135368818.339,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.307917089,"d":"5.31E+00"}]
[37751.074409733796,3135368828.834,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.309449045,"d":"5.31E+00"},{"$nulls":4},37751.07453131945,3135368839.339,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.190436693,"d":"5.19E+00"},{"$nulls":4},37751.07464699074,3135368849.333,{"v":9.999831e-10,"d":"1.00E-09"},{"v":6.208389137,"d":"6.21E+00"},{"$nulls":4},37751.07476278935,3135368859.338,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.303674972,"d":"5.30E+00"}]
[37751.07487846065,3135368869.332,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.320366452,"d":"5.32E+00"},{"$nulls":4},37751.075000046294,3135368879.837,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.201353532,"d":"5.20E+00"},{"$nulls":4},37751.0752315162,3135368899.836,{"v":9.999825e-10,"d":"1.00E-09"},{"v":6.277665001,"d":"6.28E+00"},{"$nulls":4},37751.07535298611,3135368910.331,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.30348569,"d":"5.30E+00"}]
[37751.075468773146,3135368920.335,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.306354111,"d":"5.31E+00"},{"$nulls":4},37751.07558457176,3135368930.34,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.183662349,"d":"5.18E+00"},{"$nulls":4},37751.075700243055,3135368940.334,{"v":9.999832e-10,"d":"1.00E-09"},{"v":6.23468983,"d":"6.23E+00"},{"$nulls":4},37751.075816041666,3135368950.339,{"v":9.999835e-10,"d":"1.00E-09"},{"v":5.301428455,"d":"5.30E+00"}]
[37751.075937511574,3135368960.834,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.295536093,"d":"5.30E+00"},{"$nulls":4},37751.07605909722,3135368971.339,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.183335208,"d":"5.18E+00"},{"$nulls":4},37751.07617476852,3135368981.333,{"v":9.999828e-10,"d":"1.00E-09"},{"v":6.130866451,"d":"6.13E+00"},{"$nulls":4},37751.07629055555,3135368991.337,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.285267487,"d":"5.29E+00"}]
[37751.07651623843,3135369010.836,{"v":9.999829e-10,"d":"1.00E-09"},{"v":5.304365332,"d":"5.30E+00"},{"$nulls":4},37751.07663782407,3135369021.341,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.175923213,"d":"5.18E+00"},{"$nulls":4},37751.07675349537,3135369031.335,{"v":9.999831e-10,"d":"1.00E-09"},{"v":6.112324185,"d":"6.11E+00"},{"$nulls":4},37751.07698496528,3135369051.334,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.275552269,"d":"5.28E+00"}]
[37751.077100752314,3135369061.338,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.287342897,"d":"5.29E+00"},{"$nulls":4},37751.07721643519,3135369071.333,{"v":9.999837e-10,"d":"1.00E-09"},{"v":5.172105004,"d":"5.17E+00"},{"$nulls":4},37751.07733222222,3135369081.337,{"v":9.999841e-10,"d":"1.00E-09"},{"v":6.024728962,"d":"6.02E+00"},{"$nulls":4},37751.07744800926,3135369091.341,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.273022008,"d":"5.27E+00"}]
[37751.07767368056,3135369110.839,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.280117924,"d":"5.28E+00"},{"$nulls":4},37751.077795150464,3135369121.334,{"v":9.999838e-10,"d":"1.00E-09"},{"v":5.170723023,"d":"5.17E+00"},{"$nulls":4},37751.07801503472,3135369140.332,{"v":9.999832e-10,"d":"1.00E-09"},{"v":6.128135135,"d":"6.13E+00"},{"$nulls":4},37751.078130821756,3135369150.336,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.225859197,"d":"5.23E+00"}]
[37751.07835649305,3135369169.834,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.115927933,"d":"5.12E+00"},{"$nulls":4},37751.07847807871,3135369180.339,{"v":9.999835e-10,"d":"1.00E-09"},{"v":5.106792735,"d":"5.11E+00"},{"$nulls":4},37751.07859376157,3135369190.334,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.033001262,"d":"5.03E+00"},{"$nulls":4},37751.07871534723,3135369200.839,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.075779939,"d":"5.08E+00"}]
[37751.07883681713,3135369211.334,{"v":9.999828e-10,"d":"1.00E-09"},{"v":5.119612978,"d":"5.12E+00"},{"$nulls":4},37751.07895840278,3135369221.839,{"v":9.999831e-10,"d":"1.00E-09"},{"v":5.107167335,"d":"5.11E+00"},{"$nulls":4},37751.07907987269,3135369232.334,{"v":9.999837e-10,"d":"1.00E-09"},{"v":5.033489785,"d":"5.03E+00"},{"$nulls":4},37751.079201458335,3135369242.839,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.028817867,"d":"5.03E+00"}]
[37751.07932292824,3135369253.334,{"v":9.999837e-10,"d":"1.00E-09"},{"v":5.119622012,"d":"5.12E+00"},{"$nulls":4},37751.079438726854,3135369263.339,{"v":9.999832e-10,"d":"1.00E-09"},{"v":5.105534333,"d":"5.11E+00"},{"$nulls":4},37751.07955439815,3135369273.333,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.0316593,"d":"5.03E+00"},{"$nulls":4},37751.07967018518,3135369283.337,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.028875295,"d":"5.03E+00"}]
[37751.07978586805,3135369293.332,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.13677554,"d":"5.14E+00"},{"$nulls":4},37751.07990165509,3135369303.336,{"v":9.999837e-10,"d":"1.00E-09"},{"v":5.102153839,"d":"5.10E+00"},{"$nulls":4},37751.080127326386,3135369322.834,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.027135893,"d":"5.03E+00"},{"$nulls":4},37751.08024891204,3135369333.339,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.021304982,"d":"5.02E+00"}]
[37751.080364594905,3135369343.334,{"v":9.999835e-10,"d":"1.00E-09"},{"v":5.354100745,"d":"5.35E+00"},{"$nulls":4},37751.08048618056,3135369353.839,{"v":9.999833e-10,"d":"1.00E-09"},{"v":5.104326545,"d":"5.10E+00"},{"$nulls":4},37751.08060765046,3135369364.334,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.029022353,"d":"5.03E+00"},{"$nulls":4},37751.08072923611,3135369374.839,{"v":9.99983e-10,"d":"1.00E-09"},{"v":5.151785906,"d":"5.15E+00"}]
[37751.08085070602,3135369385.334,{"v":9.999834e-10,"d":"1.00E-09"},{"v":5.118667256,"d":"5.12E+00"},{"$nulls":4},37751.08093172454,3135369392.334,{"v":9.999835e-10,"d":"1.00E-09"},{"$nulls":5},37751.08097229167,3135369395.839,{"v":9.99983e-10,"d":"1.00E-09"},{"$nulls":5},1462.0833352662012]
[37751.08101274305,3135369399.334,{"v":9.999835e-10,"d":"1.00E-09"},{"$nulls":5},1462.0833352662012,{"$nulls":7},1462.0833352662012,{"$nulls":7},1462.0833352662012]