## Slide 1: LHC Beam Loss Monitoring System LHC Beam Commissioning Workshop 19-20 January 2010
Christos Zamantzas for the BLM team
Bernd Dehning, Claudine Chery, Ewald Effinger, Jonathan Emery, Csaba Jr Hajdu, Eva Barbara Holzer, Christoph Kurfuerst, Aurelien Marsili, Annika Nordt, Mariusz Sapinski, Raymond Tissier, Giuseppe Guido Venturini

## Slide 2: Contents
Acquisition
Threshold Values
Processing
Applications & Databases
Reliability

## Slide 3: Acquisition
Section 1:

## Slide 4: Offset Level
18.12.2009 - 01.01.2010
18.12.2009
Offset [pA] | % of total | Comments
1-8 | 6.91% | Too low (problem has been identified)
10-20 | 77.71% | Very good
20-30 | 6.83% | Good
30-80 | 3.28% | High (DAC reset needed)
>80 | 5.00% | Problematic
=0 | 0.28% | no data (faulty cards/connections)
Relevant for:
Minimum detectable signal or Sensitivity of loss measurements
False triggers – LHC availability
Mean offset values for all IC monitors 1 hours of data – 18/12/2009

## Slide 5: Noise Level
Single channel frequency distribution:
Noise on short cable (RS01)
Noise on long cable (RS01)
Noise [BITS] | Monitors | % of total | Comments
1-30 | 1083 | 30.1% | very good
30-100 | 1646 | 45.7% | good
100-200 | 600 | 16.7% | ok
200-300 | 201 | 5.6% | candidates for problematic channels
> 300 | 54 | 1.5% | critical noise (channels to be repaired)
= 0 | 18 | 0.5% | no data from 3 cards
Noise in RS01 for all IC monitors
9 hours of data; maximum values – 13/01/2010
Several cables have been exchanged
some of them 800m long
Installation consists of 3 batches from different manufacturers
Expected noise reduction
Standard cables is factor 2
New cables is > factor 5 (single pair shielded)
Mid-term plan use new type of cables

## Slide 6: Noise Level
Secondary Emission Monitors (302)
Max. noise value frequency distribution:
Ionisation Chambers (3602)
Noise [BITS] | Monitors | % of total | Comments
1-30 | 1083 | 30.1% | very good
30-100 | 1646 | 45.7% | good
100-200 | 600 | 16.7% | ok
200-300 | 201 | 5.6% | candidates for problematic channels
> 300 | 54 | 1.5% | critical noise (channels to be repaired)
= 0 | 18 | 0.5% | no data from 3 cards
Noise in RS01 for all IC monitors
9 hours of data; maximum values – 13/01/2010
Installation consists of 3 batches from different manufacturers
Several cables have been exchanged
some of them 800m long
Expected noise reduction
Standard cables is factor 2
New cables is > factor 5 (single pair shielded)
Mid-term plan use new type of cables

## Slide 7: Ongoing Modifications
Some acquisition cards need to be exchanged
  Cards affected ~ 4%
  Too low offset current
Interconnection boxes
  Ionisation of the air inside the box measured
  Adding better isolation to all SEM
Global reset of cards not reliable
  Cards affected ~ 5%
  Expected to get some improvement with WorldFIP; uncertain on WorldFIP reliability under radiation
  Currently only point 6 has been connected.

## Slide 8: Over-Injection problem
BLMQI.03L2.B2I30_MQXA
BLMQI.03R8.B1I30_MQXA
Problem consists of losses at the LHC TDI and exceeding thresholds at the MQXA
  Some small difference between injected beams
  Installation has been verified to be correct.

## Slide 9: Over-Injection tests
Extensive tests done (with beam 2):
Over-injection, TDI open, 2e9 proton
Injection on closed TDI (+- 3mm), kicker off
Injection on closed TDI (+- 3mm) and TCTV (upper jaw -6 mm, down jaw -12 mm)‏
No significant difference observed
Most probably losses come from outside the cryostat

## Slide 10: Over-Injection Solutions
Spread the signal by hardware means
  installation of an additional capacitor
Add shielding
  most favourable
  needs simulations
Last year's test with threshold values set over the maximum allowed limit is not a safe solution.

## Slide 11: Threshold values
Section 2:

## Slide 12: Thresholds & Noise
Frequency distribution of the maximum noise values over 14 days for each channel
(starting from 18-12-2009)
All Ionisation Chamber monitors except on collimators (= 3432)
Threshold values are MASTER * 0.1 (i.e. default monitor coefficient)
All Ionisation Chamber monitors on collimators (= 160)
Threshold values are MASTER * 1.0 (i.e. maximum monitor coefficient)
No significant problems expected
Some channels to be repaired
Thresholds can be increased with no safety impact.
Significantly lower values
Thresholds cannot be increased using the coefficient.
Recalculation of values will be done.

## Slide 13: Threshold Values
No threshold values are missing
  There are some weak points
  Some precise simulations are missing
Tests for verification and fine tuning of the threshold values.
  transient losses with special nQPS
  steady-state losses with temperature sensors
More on the Chamonix talk.

## Slide 14: Processing
Section 3:

## Slide 15: System latency
Analysed several Beam Dumps in detail.
Difference in time between the bunch at the MKI and the break of the beam permit loop (by the BLMS) recorded at the BIC was always between 100 and 130 μs.
Time of flight:
  MKI => monitor = 10 μs
  monitor => acq. electronics (0.5 km cable) = 3 μs
  Acq. => processing electronics (1 km fibre) = 3 μs
Detection of change in frequency in the daisy-chain = 5 μs
Integration in the acquisition electronics = 40 μs.
Decision at the BLETC (for fast losses) is taken every 40 μs.
Processing of data to decide < 1 μs.

## Slide 16: System Self-Monitoring
Example: Continuous checks on the data transmission between the tunnel and the surface installation:
Link A | Link B | ERR A | ERR B | Ignore CFC | CHOISE | Comment
0 | 0 | 0 | 0 | x | LK_B | all ok (choice based on PCB properties)
0 | 0 | 0 | 1 | x | LK_A | LK_B has errors
0 | 0 | 1 | 0 | x | LK_B | LK_A has errors
0 | 0 | 1 | 1 | x | Dump & LK_B | LK_A and LK_B have errors
0 | 1 | 0 | x | x | LK_A | LK_B is down
0 | 1 | 1 | x | x | Dump & LK_A | LK_B is down and LK_A has error
1 | 0 | x | 0 | x | LK_B | LK_A is down
1 | 0 | x | 1 | x | Dump & LK_B | LK_A is down and LK_B has error
1 | 1 | x | x | 0 | Dump & LK_B | LK_A and LK_B are both down
1 | 1 | x | x | 1 | LK_B | Both links down but CFC not present
Activation of all the self-monitoring processes:
Acquisition operation state monitoring (8 cases)
Configuration cross-check (one case)
Data transmission between the tunnel and the surface installation (14 cases)
Inhibit Beam Dump requests from cards not connected to BIS (2 cases)
Failed cases induce a beam dump

## Slide 17: Applications & DB
Section 4:

## Slide 18: Storage of Parameters
Restructure of all databases:
  MTF – remove BLMS architecture; keep only ref. measurements and components track history [ongoing]
  Layout – become the entry point for updating connections or adding new installations [ongoing]
  LSA – reorganisation of tables to become more relational [done]
Internal LSA DB Constrains will be reviewed
Internal DB Check for disabled channels [in testing]
  Based on monitor criticality and adjacent disabled channels
  Each monitors is being tagged on its criticality
  Current version only informs on rules violations

## Slide 19: Data Concentration
Logging data
  Reason for losing packets still uncertain – investigation continues
  Modification of CMW settings and Concentrator have improved the situation (positive side effect: FECs have stopped rebooting)
Capture data
  Temporary solution with “synchronised GET” seems to work
  Will test more and return to 2048 samples
JMS broker – new configuration and tests are foreseen

## Slide 20: ALARMS
Warnings for the ALARMS consoles to be added:
System Check is ON-GOING
  concentrated from all checks
Timeout from last System Check - TEST MANDATORY
  MCS online, Connectivity, Beam Permit, ...
System Check has failed
  console should give some details on actions
System has requested a Beam Dump.

## Slide 21: Data Displays/Applications
Most applications ready.
  Fixed display,
  Capture display,
  Monitor factor editor, ...
Status application
  Main parts ready – still several parts missing
  Need usability update for OP
Management application
  Only GUI, missing code
Post-Mortem modules update
  Adding some more variables (mainly statuses).
  Adding more functions and better decoding of data.
Minor to no changes expected
Ongoing
Lower priority
Ongoing

## Slide 22: Reliability
Section 5:

## Slide 23: Automated checks
MCS Online check
  Verify consistency of parameters
Connectivity (HV Modulation) check
  Verify monitor connections
Internal Beam Permit check (BLETC -> BLECS)
  Verify the ability of each card to trigger a dump
Will activate in the firmware the enforcing of the agreed periodic run of those checks by the Sequencer.
External Beam Permit check (BLMS -> BIS)
  Verify connection to BIS
One minor bug pending
Finalising
Completed
Completed

## Slide 24: Connectivity check
View on the LHC BLM system for the connectivity check.
The high voltage power supply to the monitors is modulated with a 60mHz 30V sinusoidal signal.
A current on the monitors is generated by the HV modulation and detected by the normal acquisition system of the BLM.
The combiner card uses the running sum (RS_09) from the threshold comparators to determine the amplitude and phase of each monitor (256 per crate).
This results are send to the logging and compared to thresholds to permit or block the next injection if a non conformity is detected.
The signals of each monitors are stored (in the SRAMs) (see right plots) to be read on demand with the diagnostic application (under development).
The thresholds (called internal parameters) are unique for each monitor. They are calculated out of multiple measurements by a script gathering data in the logging.
View of the SRAM data containing the original (stairs) and filtered RS09 data of each channel for one crate.

## Slide 25: System Verification Tests
Automatic extraction from DB + human inspection (repeated daily or weekly)
  Frequency distribution of channel noise
  Optical links degradation
Automatic verification on the Vertical Slice Test system (before releasing new firmware)
  Exhaustive Threshold triggering tests (also used in LHC)
  Reception and Status tests
  Linearity, impulse, and some predefined patterns of input signals check
  …
In the future will run inside the DB

## Slide 26: Summary

## Slide 27: Summary
BLMS removes correctly the BEAM PERMIT signal if measurements are over threshold.
No reliability issues observed.
System is well understood since it has been up and running for more than a year.
Some availability issues (false dumps) at energies higher than the injection are to be expected if thresholds don’t change in some regions.
Continuous monitoring of noise is required.
Sequencer initiated tests will be enforced to be run regularly.
More tests to verify and adjust the threshold values are needed.
Investigation of spurious signals from the SEMs are ongoing and first corrections are being implemented.

## Slide 29: Reserved Slides

## Slide 30: System overview

## Slide 31: New noise reduction algorithm
Additional digital filter to reduce “noise” from the ADC circuit.
  Safe; operates only when losses are very small
  Verified to be correct; operational in 2009.
RS01
RS02
RS06
RS09
OLD
NEW
Rad Source Measurements with different firmware

## Slide 32: Variables Update
Displays should be driven as much as possible by the electronics.
  Avoid wrong information on the displays
  Candidates for inclusion:
    ChannelMask flag
    IsCableConnected flag
    IsConnectedtoBIS flag
Provide a bit encoded variable for the reason of the Beam Dump Request.
  Final details of request is being discussed

## Slide 33: Expert Applications
Management of parameters (Expert Threshold App)
  Disabling and masking of channels
  Configuration of mobile monitors
  Pass complete login info to DB
  Connect using LSA DB API
  Roll-back on history
Diagnostics of misbehaving channels (for when checks fail)
  expert tool initially but later will hand over to OP.
Next weeks
Next months

## Slide 34: Other Known Issues:
PM/XPOC data are some times wrongly decoded [Firmware]
PM data channels that have been over threshold remain in history [FESA]
MCS fails always the first time [CO and FESA]
Collimator beam based alignment will not be available for this year [Collimation and BLM]

## Slide 35: Beam Permit Lines Check
The internal beam permit check verifies the ability of each BLETC to remove the beam permit.
A timer ensure that this check is done at least every 24 hours (just before the connectivity check)
The external beam permit check verifies that the BIS (CIBUS interface) receives correctly the beam permit signal.
These lines are under the responsibility of the interlock system. A BLM interface is available to control these lines. Regular checks are foreseen.
BLECS
U
BLETC N°1
BLETC N°16
BLETC N°2
U
M
M
M
Last BLECS
U
BLETC N°1
BLETC N°16
BLETC N°2
U
M
M
U
M
CIBUS Unmaskable
U
Daisy chain from TC1 to BLECS
Daisy chain of BLM crates (up to 4) from BLECS1 to CIBUS
BLM Crate 1
CIBUS Maskable
BIS network