## Installation overview
BLECF installation (tunnel electronic) | [13 empty] | Tests
| ARC | [6 empty] | straight section | [5 empty] | otical link |  |  |  |  | HV levels test |  |  | rad.
| BLECF |  |  |  | Backplane | 220V | Completed | BLECF |  |  |  | 220V |  | connection | connection | connection | TC | CFC - |  | resistor | completely | source
| to install | installed | Rst. OK | tested | soldered | checked |  | to install | installed | Rst. OK | tested | checked | Completed | total installed | tunnel surface | Jumper |  | ID | scanned | replaced | tested | test
1 LEFT | 22 | 22 | 22 | 22 | 22 | 22 | OK | 19 | 19 | 19 | 19 | 19 | OK | 82 | 82 | 82 | 21 | 41
1 RIGHT | 23 | 23 | 23 | 23 | 23 | 23 | OK | 19 | 19 | 19 | 19 | 19 | OK | 84 | 84 | 84 | 21 | 42
2 LEFT | 22 | 22 | 22 | 22 | 22 | 22 | OK | 22 | 22 | 22 | 22 | 22 | OK | 88 | 86 +2 spares | 88 | 22 | 44
2 RIGHT | 23 | 23 | 23 | 23 | 23 | 23 | OK | 21 | 21 | 21 | 21 | 21 | OK | 88 | 88 | 88 | 22 | 44
3 LEFT | 22 | 22 | 22 | 22 | 22 | 22 | OK | 18 | 18 | 18 | 18 | 18 | OK | 80 | 80 | 80 | 20 | 40
3 RIGHT | 23 | 23 | 23 | 23 | 23 | 23 | OK | 18 | 18 | 18 | 18 | 18 | OK | 82 | 80 + 2 spares | 82 | 21 | 41
4 LEFT | 22 | 22 | 22 | 22 | 22 | 22 | OK | 8 | 8 | 8 | 8 | 8 | OK | 60 | 60 | 60 | 15 | 30
4 RIGHT | 23 | 23 | 23 | 23 | 23 | 23 | OK | 8 | 8 | 8 | 8 | 8 | OK | 62 | 62 | 62 | 14 | 31
5 LEFT | 22 | 22 | 22 | 22 | 22 | 22 | OK | 19 | 19 | 19 | 19 | 19 | OK | 82 | 81 | 82 |  | 41
5 RIGHT | 23 | 23 | 23 | 23 | 23 | 23 | OK | 19 | 19 | 19 | 19 | 19 | OK | 84 | 84 | 84 |  | 42
6 LEFT | 22 | 22 | 22 | 22 | 22 | 22 | OK | 21 | 21 | 21 | 21 | 21 | OK | 86 | 86 | 86 |  | 43
6 RIGHT | 23 | 23 | 23 | 23 | 23 | 23 | OK | 21 | 21 | 21 | 21 | 21 | OK | 88 | 88 | 88 |  | 44
7 LEFT | 22 | 22 | 22 | 22 | 22 | 22 | OK | 28 | 28 | 28 | 28 | 28 | OK | 100 | 100 | 99 |  | 50
7 RIGHT | 23 | 23 | 23 | 23 | 23 | 23 | OK | 28 | 28 | 28 | 28 | 28 | OK | 102 | 102 | 102 |  | 51
8 LEFT | 22 | 22 | 22 | 22 | 22 | 22 | OK | 19 | 19 | 19 | 19 | 19 | OK | 82 | 82 | 82 |  | 41
8 RIGHT | 23 | 23 | 23 | 23 | 23 | 23 | OK | 21 | 21 | 21 | 21 | 21 | OK | 88 | 88 | 88 |  | 44
Total | 360 | 360 | 360 | 360 | 360 | 360 |  | 309 | 309 | 309 | 309
in % |  | 100 | 100 | 100 | 100 | 100 |  |  | 100 | 100 | 100
Missing |  | 0 | 0 | 0 | 0 | 0 |  |  | 0 | 0 | 0
Sum BLECF complete installation |  |  |  |  | 669
Installed BLECF |  |  |  |  | 669 |  | in % | 100 | [11 empty] | updated 28.8.08
Missing BLECF complete installation |  |  |  |  | 0 |  | in % | 0
Note on F39: effinger: BJBHT.A32L7 220V missing --> Rack is connected to a closeby 220V power distribution
Note on F31: effinger: BYPLM.A17R5: Power supply 5V not working --> repaired, backplane was not soldered! --> Card is working and tested!

## Point 1
| Tunnel - position | [13 empty] | Surface position | [25 empty] | LEVEL | [8 empty] | Signal
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | Comment | ERRA | ERRB | ERRA | ERRB | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | HV | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Comment | Reset | Capa | 220V | Noise | HV | Source | other
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec. | 23.6.08 |  |  | 8.8.08 |  | 12/18.08.08 |  | 26-27.8.08 | [25 empty] | tested | tested |  | levels | tested | test | Comments
1 | BJBHT.A33L1 | BLECF_ | 285 | 0x | 011D | 6 | Arc | BYPLM.A33L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
2 | BJBHT.A32L1 | BLECF_ | 280 | 0x | 0118 | 6 | Arc | BYPLM.A32L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | OK | OK |  | ↑↑↑ | OK | ↑ | OK | ↑ | OK | N | N | OK | OK | OK | CX | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
3 | BJBHT.A31L1 | BLECF_ | 276 | 0x | 0114 | 6 | Arc | BYPLM.A31L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
4 | BJBHT.A30L1 | BLECF_ | 201 | 0x | 00C9 | 6 | Arc | BYPLM.A30L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
5 | BJBHT.A29L1 | BLECF_ | 141 | 0x | 008D | 6 | Arc | BYPLM.A29L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | 1-20 | [8 empty] | OK | OK | OK
6 | BJBHT.A28L1 | BLECF_ | 249 | 0x | 00F9 | 6 | Arc | BYPLM.A28L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | LF | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
7 | BJBHT.A27L1 | BLECF_ | 241 | 0x | 00F1 | 6 | Arc | BYPLM.A27L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-30 (80) | [8 empty] | OK | OK | OK
8 | BJBHT.A26L1 | BLECF_ | 282 | 0x | 011A | 6 | Arc | BYPLM.A26L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | OK | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | ↑ | OK | ↑ | 1X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | 1-30 | [8 empty] | OK | OK | OK
9 | BJBHT.A25L1 | BLECF_ | 218 | 0x | 00DA | 6 | Arc | BYPLM.A25L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-10 | [8 empty] | OK | OK | OK
10 | BJBHT.A24L1 | BLECF_ | 204 | 0x | 00CC | 6 | Arc | BYPLM.A24L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | LF | LF |  | LF | LF | OK | ↑ | OK | OK | N | N | OK | OK | OK | 2-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-10 | [8 empty] | OK | OK | OK
11 | BJBHT.A23L1 | BLECF_ | 212 | 0x | 00D4 | 6 | Arc | BYPLM.A23L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-10 | [8 empty] | OK | OK | OK
12 | BJBHT.A22L1 | BLECF_ | 284 | 0x | 011C | 6 | Arc | BYPLM.A22L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | OK | OK |  | ↑↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-10 (100) | [8 empty] | OK | OK | OK
13 | BJBHT.A21L1 | BLECF_ | 22 | 0x | 0016 | 6 | Arc | BYPLM.A21L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-60 | [8 empty] | OK | OK | OK
14 | BJBHT.A20L1 | BLECF_ | 21 | 0x | 0015 | 6 | Arc | BYPLM.A20L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | OK | OK |  | ↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-60 (160) | [8 empty] | OK | OK | OK
15 | BJBHT.A19L1 | BLECF_ | 35 | 0x | 0023 | 6 | Arc | BYPLM.A19L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | 1-30 (90) | [8 empty] | OK | OK | OK
16 | BJBHT.A18L1 | BLECF_ | 34 | 0x | 0022 | 6 | Arc | BYPLM.A18L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-30 (90) | [8 empty] | OK | OK | OK
17 | BJBHT.A17L1 | BLECF_ | 36 | 0x | 0024 | 6 | Arc | BYPLM.A17L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-30 | [8 empty] | OK | OK | OK
18 | BJBHT.A16L1 | BLECF_ | 52 | 0x | 0034 | 6 | Arc | BYPLM.A16L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | OK | OK |  | ↑↑↑ | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-30 | [8 empty] | OK | OK | OK
19 | BJBHT.A15L1 | BLECF_ | 59 | 0x | 003B | 6 | Arc | BYPLM.A15L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | 1-30 | [8 empty] | OK | OK | OK
20 | BJBHT.A14L1 | BLECF_ | 58 | 0x | 003A | 6 | Arc | BYPLM.A14L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | LF | LF |  | ↑↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-30 | [8 empty] | OK | OK | OK
21 | BJBHT.A13L1 | BLECF_ | 77 | 0x | 004D | 8 | Arc | BYPLM.A13L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-30 | [8 empty] | OK | OK | OK
22 | BJBHT.A12L1 | BLECF_ | 79 | 0x | 004F | 6 | Arc | BYPLM.A12L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-30 | [8 empty] | OK | OK | OK
23 | BJBHT.B11L1 | BLECF_ | 667 | 0x | 029B | 8 | RR13 | BY02 | 1 | CL | 2 | 9 | B | 13-14 | SR1 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | LF | LF |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 2-120 | [8 empty] | OK | OK | OK
24 | BJBHT.A11L1 | BLECF_ | 594 | 0x | 0252 | 8 | RR13 | BY02 | 1 | CL | 2 | 8 | B | 11-12 | SR1 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | no TC | no TC |  | ↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X-53 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
25 | BJBHT.B10L1 | BLECF_ | 409 | 0x | 0199 | 6 | RR13 | BY02 | 1 | CL | 2 | 7 | B | 9-10 | SR1 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 7 | 4 | 1-8 | OK | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 2-120 | [8 empty] | OK | OK | OK
26 | BJBHT.A10L1 | BLECF_ | 550 | 0x | 0226 | 2 | RR13 | BY02 | 1 | CL | 2 | 6 | B | 7-8 | SR1 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X (ch1,2 50) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
27 | BJBHT.A9L1 | BLECF_ | 518 | 0x | 0206 | 6 | RR13 | BY02 | 1 | CL | 2 | 5 | B | 5-6 | SR1 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 2-90 | [8 empty] | OK | OK | OK
28 | BJBHT.B8L1 | BLECF_ | 457 | 0x | 01C9 | 6 | RR13 | BY02 | 1 | CL | 2 | 4 | B | 3-4 | SR1 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8 | OK | OK |  | ↑↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 2-90 | [8 empty] | OK | OK | OK
29 | BJBHT.A8L1 | BLECF_ | 588 | 0x | 024C | 6 | RR13 | BY02 | 1 | CL | 2 | 3 | B | 1-2 | SR1 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | OK | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
30 | BJBHT.B7L1 | BLECF_ | 322 | 0x | 0142 | 7 | RR13 | BY02 | 1 | CL | 2 | 2 | A | 23-24 | SR1 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 2-40 | [8 empty] | OK | OK | OK
31 | BJBHT.A7L1 | BLECF_ | 194 | 0x | 00C2 | 1 | RR13 | BY02 | 1 | CL | 2 | 1 | A | 21-22 | SR1 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 7 | 4 | 9-16 | no TC | no TC |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X (ch1 50) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
32 | BJBHT.A6L1 | BLECF_ | 638 | 0x | 027E | 6 | RR13 | BY02 | 1 | CL | 1 | 10 | A | 19-20 | SR1 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | LF | LF |  | OK | OK | ↑ | OK | ↑ | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 2-40 | [8 empty] | OK | OK | OK
33 | BJBHT.B5L1 | BLECF_ | 530 | 0x | 0212 | 7 | RR13 | BY02 | 1 | CL | 1 | 9 | A | 17-18 | SR1 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | OK | OK |  | ↑↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 2-100 | [8 empty] | OK | OK | OK
34 | BJBHT.A5L1 | BLECF_ | 597 | 0x | 0255 | 1 | RR13 | BY02 | 1 | CL | 1 | 8 | A | 15-16 | SR1 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8 | no TC | no TC |  | ↑↑ | OK | ↑ | OK | ↑ | OK | N | N | OK | OK | OK | 0X (ch1 27) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
35 | BJBHT.C4L1 | BLECF_ | 398 | 0x | 018E | 6 | RR13 | BY02 | 1 | CL | 1 | 7 | A | 13-14 | SR1 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | OK | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 2-100 | [8 empty] | OK | OK | OK
36 | BJBHT.B4L1 | BLECF_ | 374 | 0x | 0176 | 4 | RR13 | BY02 | 1 | CL | 1 | 6 | A | 11-12 | SR1 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X (ch1-4) 0x (ch5-8) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
37 | BJBHT.A4L1 | BLECF_ | 595 | 0x | 0253 | 4 | RR13 | BY02 | 1 | CL | 1 | 5 | A | 9-10 | SR1 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | LF | LF |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 5X (ch 1-4) 0X (ch5-8) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
38 | BJBHT.A3L1 | BLECF_ | 593 | 0x | 0251 | 5 | RR13 | BY02 | 1 | CL | 1 | 4 | A | 7-8 | SR1 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | no TC | no TC |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X(ch1-5) 0X(ch6-8) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
39 | BJBHT.A2L1 | BLECF_ | 592 | 0x | 0250 | 8 | RR13 | BY02 | 1 | CL | 1 | 3 | A | 5-6 | SR1 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | no TC | no TC |  | ↑↑↑ | OK | ↑ | OK | ↑ | OK | N | N | ↑ | OK | OK | 4X-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [10 empty] | OK | OK
40 | BJBHT.B1L1 | BLECF_ | 635 | 0x | 027B | 5 | RR13 | BY02 | 1 | CL | 1 | 2 | A | 3-4 | SR1 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | no TC | no TC |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X(ch1-5) 1X(ch6-8) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
41 | BJBHT.A1L1 | BLECF_ | 639 | 0x | 027F | 1 | RR13 | BY02 | 1 | CL | 1 | 1 | A | 1-2 | SR1 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | no TC | no TC |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X (ch1 34) | ok | ok | bl | ok | ok | bl | bl | ok | ok | [9 empty] | OK | OK | OK
42 | BJBHT.A1R1 | BLECF_ | 377 | 0x | 0179 | 1 | RR17 | BY02 | 1 | CR | 1 | 1 | A | 1-2 | SR1 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X (ch4,6 7X) ch1 27 | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | sem, Modulation small but ok | OK | OK | OK |  |  |  | GND problem , HV status not o.k
43 | BJBHT.B1R1 | BLECF_ | 331 | 0x | 014B | 5 | RR17 | BY02 | 1 | CR | 1 | 2 | A | 3-4 | SR1 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X (ch5 26) | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK |  |  |  | small noise, TEST_CFC
44 | BJBHT.A2R1 | BLECF_ | 501 | 0x | 01F5 | 8 | RR17 | BY02 | 1 | CR | 1 | 3 | A | 5-6 | SR1 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 9X-BX (ch1 4F) | ok | bl | ok | ok | ok | bl | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | GND problem , LEVEL is blinking
45 | BJBHT.A3R1 | BLECF_ | 636 | 0x | 027C | 5 | RR17 | BY02 | 1 | CR | 1 | 4 | A | 7-8 | SR1 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK |  |  |  | 500-600bit noise
46 | BJBHT.A4R1 | BLECF_ | 676 | 0x | 02A4 | 4 | RR17 | BY02 | 1 | CR | 1 | 5 | A | 9-10 | SR1 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 3X (ch5-8 0X) | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
47 | BJBHT.B4R1 | BLECF_ | 653 | 0x | 028D | 4 | RR17 | BY02 | 1 | CR | 1 | 6 | A | 11-12 | SR1 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X (ch5-8 1X) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
48 | BJBHT.C4R1 | BLECF_ | 376 | 0x | 0178 | 6 | RR17 | BY02 | 1 | CR | 1 | 7 | A | 13-14 | SR1 | BY05 | C | 11 | 1-2 | BY02 | R | 3 | 4 | 1 | 1-8 | OK | OK |  | ↑↑↑ | OK | ↑↑ | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 3-90 | [8 empty] | OK | OK | OK |  |  |  | 50-100bit noise
49 | BJBHT.A5R1 | BLECF_ | 437 | 0x | 01B5 | 1 | RR17 | BY02 | 1 | CR | 1 | 8 | A | 15-16 | SR1 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | ↑ C | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X (ch1 26) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
50 | BJBHT.B5R1 | BLECF_ | 442 | 0x | 01BA | 7 | RR17 | BY02 | 1 | CR | 1 | 9 | A | 17-18 | SR1 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 3-90 | [8 empty] | OK | OK | OK |  |  |  | 200-300 bit noise
51 | BJBHT.A6R1 | BLECF_ | 383 | 0x | 017F | 6 | RR17 | BY02 | 1 | CR | 1 | 10 | A | 19-20 | SR1 | BY05 | C | 11 | 7-8 | BY02 | R | 3 | 5 | 2 | 1-8 | ↑↑↑ | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-40 | [8 empty] | OK | OK | OK
52 | BJBHT.A7R1 | BLECF_ | 323 | 0x | 0143 | 1 | RR17 | BY02 | 1 | CR | 2 | 1 | A | 21-22 | SR1 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 16 | 12 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X (ch1 48) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
53 | BJBHT.B7R1 | BLECF_ | 512 | 0x | 0200 | 7 | RR17 | BY02 | 1 | CR | 2 | 2 | A | 23-24 | SR1 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-40 | [8 empty] | OK | OK | OK |  |  |  | 50 bit noise
54 | BJBHT.A8R1 | BLECF_ | 325 | 0x | 0145 | 6 | RR17 | BY02 | 1 | CR | 2 | 3 | B | 1-2 | SR1 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | ↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X (ch7-8 0X) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
55 | BJBHT.B8R1 | BLECF_ | 647 | 0x | 0287 | 6 | RR17 | BY02 | 1 | CR | 2 | 4 | B | 3-4 | SR1 | BY05 | C | 10 | 3-4 | BY02 | R | 3 | 6 | 3 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 3-100 | [8 empty] | OK | OK | OK |  |  |  | 50 bit noise
56 | BJBHT.A9R1 | BLECF_ | 85 | 0x | 0055 | 6 | RR17 | BY02 | 1 | CR | 2 | 5 | B | 5-6 | SR1 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | 3-100 | [8 empty] | OK | OK | OK |  |  |  | CFC-TEST-ON
57 | BJBHT.A10R1 | BLECF_ | 338 | 0x | 0152 | 2 | RR17 | BY02 | 1 | CR | 2 | 6 | B | 7-8 | SR1 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X (ch1-2 4D) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
58 | BJBHT.B10R1 | BLECF_ | 567 | 0x | 0237 | 6 | RR17 | BY02 | 1 | CR | 2 | 7 | B | 9-10 | SR1 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 7 | 4 | 1-8 | ↑ 83 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 3-100 | [8 empty] | OK | OK | OK |  |  |  | 400-500 bit
59 | BJBHT.A11R1 | BLECF_ | 349 | 0x | 015D | 8 | RR17 | BY02 | 1 | CR | 2 | 8 | B | 11-12 | SR1 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
60 | BJBHT.B11R1 | BLECF_ | 327 | 0x | 0147 | 8 | RR17 | BY03 | 2 | CR | 2 | 9 | B | 13-14 | SR2 | BY06 | C | 9 | 1-2 | BY02 | R | 3 | 7 | 4 | 9-16 | OK | OK | Level blinking | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | FF (C1 4C) | ok | bl | bl | bl | bl | bl | bl | bl | ok | C1 40-100 rest 40-300 | [8 empty] | OK | OK | OK |  |  |  | GND problem , ch1 small noise, c2-8 blocked
61 | BJBHT.A12R1 | BLECF_ | 539 | 0x | 021B | 6 | Arc | BYPLM.A12R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | TC err | TC err |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
62 | BJBHT.A13R1 | BLECF_ | 361 | 0x | 0169 | 8 | Arc | BYPLM.A13R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | TC err | TC err |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
63 | BJBHT.A14R1 | BLECF_ | 570 | 0x | 023A | 6 | Arc | BYPLM.A14R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | LF | LF |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
64 | BJBHT.A15R1 | BLECF_ | 306 | 0x | 0132 | 6 | Arc | BYPLM.A15R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | LF | LF |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
65 | BJBHT.A16R1 | BLECF_ | 311 | 0x | 0137 | 6 | Arc | BYPLM.A16R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | LF | LF |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
66 | BJBHT.A17R1 | BLECF_ | 303 | 0x | 012F | 6 | Arc | BYPLM.A17R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | LF | LF |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
67 | BJBHT.A18R1 | BLECF_ | 515 | 0x | 0203 | 6 | Arc | BYPLM.A18R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | LF | LF |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
68 | BJBHT.A19R1 | BLECF_ | 481 | 0x | 01E1 | 6 | Arc | BYPLM.A19R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | LF | LF |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
69 | BJBHT.A20R1 | BLECF_ | 317 | 0x | 013D | 6 | Arc | BYPLM.A20R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | LF | LF |  | ↑↑↑ | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
70 | BJBHT.A21R1 | BLECF_ | 522 | 0x | 020A | 6 | Arc | BYPLM.A21R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | LF | LF |  | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
71 | BJBHT.A22R1 | BLECF_ | 373 | 0x | 0175 | 6 | Arc | BYPLM.A22R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | ↑ 1A24 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
72 | BJBHT.A23R1 | BLECF_ | 423 | 0x | 01A7 | 6 | Arc | BYPLM.A23R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | LF | LF |  | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
73 | BJBHT.A24R1 | BLECF_ | 352 | 0x | 0160 | 6 | Arc | BYPLM.A24R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | OK | OK |  | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
74 | BJBHT.A25R1 | BLECF_ | 500 | 0x | 01F4 | 6 | Arc | BYPLM.A25R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
75 | BJBHT.A26R1 | BLECF_ | 479 | 0x | 01DF | 6 | Arc | BYPLM.A26R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
76 | BJBHT.A27R1 | BLECF_ | 478 | 0x | 01DE | 6 | Arc | BYPLM.A27R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | LF | LF |  | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-15 | [8 empty] | OK | OK | OK
77 | BJBHT.A28R1 | BLECF_ | 304 | 0x | 0130 | 6 | Arc | BYPLM.A28R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | OK | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
78 | BJBHT.A29R1 | BLECF_ | 310 | 0x | 0136 | 6 | Arc | BYPLM.A29R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | 1-20 | [8 empty] | OK | OK | OK
79 | BJBHT.A30R1 | BLECF_ | 329 | 0x | 0149 | 6 | Arc | BYPLM.A30R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | LF | LF |  | LF | LF | OK | OK | ↑ | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK |  |  |  | M5V missing? To be checked!
80 | BJBHT.A31R1 | BLECF_ | 495 | 0x | 01EF | 6 | Arc | BYPLM.A31R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20- | [8 empty] | OK | OK | OK
81 | BJBHT.A32R1 | BLECF_ | 384 | 0x | 0180 | 6 | Arc | BYPLM.A32R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | LF | LF |  | LF | LF | ↑ | OK | ↑ | OK | N | N | ↑ | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
82 | BJBHT.A33R1 | BLECF_ | 417 | 0x | 01A1 | 6 | Arc | BYPLM.A33R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
83 | BJBHT.A34R1 | BLECF_ | 321 | 0x | 0141 | 6 | Arc | BYPLM.A34R1 | 1 | R |  |  |  | 7-9 | SR1 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | ↑↑↑ | OK |  | OK | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 7x | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
last entry 18/6/08
[9 empty] | Transformer soldered | [16 empty] | ↑↑↑ | CRC Error consantly fast increasing
[26 empty] | ↑↑ | CRC Error sometimes fast increasing
[26 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number correct | [19 empty] | LF | Lost Frame
|  |  |  | 0x | 0000 | Card number needs to be checked | [19 empty] | no TC | No Threshold Comparator installed
|  |  |  | 0x | 0000 | Card number unknown need to be checked in the tunnel | [19 empty] | TC err | Threshold Comparator doesn't work correctly
Opitcal link is working with Test system
|  | BLECF_ | 656 | Card was exchanged, Data base need to be checked for the correct number
Note on F33: effinger: Wrong number: 0x0206 (518) and number need to be verified in the tunnel
Note on F37: effinger: number need to be verified in the tunnel
Note on D100: effinger: 565 original number chnaged to 656
Note on AC68: effinger: fibre wrong connected in the tunnel 5&6 instead of 7&8
Note on AC69: effinger: fibre wrong connected in the tunnel 5&6 instead of 7&8
Note on AC70: effinger: fibre wrong connected in the tunnel 5&6 instead of 7&8
Note on AC71: effinger: fibre wrong connected in the tunnel 5&6 instead of 7&8
Note on AC72: effinger: fibre wrong connected in the tunnel 5&6 instead of 7&8 and CFC was not working after reset o.k.
Note on AC73: effinger: fibre wrong connected in the tunnel 5&6 instead of 7&8
Note on AC74: effinger: fibre wrong connected in the tunnel 5&6 instead of 7&8 and CFC was not working after reset o.k.
Note on AC75: effinger: fibre wrong connected in the tunnel 5&6 instead of 7&8
Note on AC76: effinger: fibre wrong connected in the tunnel 5&6 instead of 7&8
Note on AA90: effinger: 242uW
Note on AB90: effinger: 208uW
Note on AA82: effinger: 234uW
Note on AB82: effinger: 296uW
Note on AA78: effinger: 80uW
Note on AB78: effinger: 270uW
Note on AA79: effinger: 300uW
Note on AB79: effinger: 204uW
Note on AA64: effinger: 252uW
Note on AB64: effinger: 184uW
Note on AA61: effinger: 174uW
Note on AB61: effinger: 237uW
Note on AA57: effinger: 170uW
Note on AB57: effinger: 236uW
Note on AC28: effinger: CFC test in the tunnel, all CFC were working, to cards A6L1 and B8L1 haven't been connected , all fibres were reconnected, CFC cards haven't been fixed with the screws -> done
Note on AA10: effinger: 215uW
Note on AB10: effinger: 333uW
Note on AA24: effinger: 302uW
Note on AB24: effinger: 293uW
Note on AC90: effinger: 12.08.08 Checked in the tunnel --> all o.k.
Note on AC88: effinger: 12.08.08 Fibre wrong connected 5&6 instead of 7&8, changed to 7&8 CFC o.k no LF
Note on AC86: effinger:12.08.08 Fibre wrong connected 5&6 instead of 7&8, one patch broken, changed to 7&8 CFC o.k no LF
Note on AC85: effinger: Card number rechecked in the tunnel : 310 installed
Note on AC83: effinger: 12.08.08 Fibre wrong connected 5&6 instead of 7&8, changed to 7&8 CFC o.k no LF
Note on AC80: effinger:12.08.08 CFC not working -> reset -> CFC o.k no LF
Note on AC79: effinger: 12.08.08 CFC o.k no LF in the tunnel
Note on AC77: effinger: 12.08.08 Fibre wrong connected 5&6 instead of 7&8, changed to 7&8 CFC o.k no LF
Note on T79: effinger: 13.08.08 need to check if the fibre are changed to 5-6 --> checked 18.8.08 connected to 7-8
Note on T14: effinger: 13.08.08 need to check if the fibre are changed to 5-6 --> 18.8.08 was connected to 5&6 reconnected to 7&8
Note on AE79: effinger: Fibre 3-4 haven't been plugged correctly inside the TC -> plugged correctly and no LF
Note on Z84: effinger: 13.08.08 Wrong labling on th the patch cord, A28R1 3&4 and A29R1 3&4 --> A28R1 connected to 1&2 and A29R1 to 3&4 --> numbers correct now!!
Note on Z85: effinger: 13.08.08 Wrong labling on th the patch cord, A28R1 3&4 and A29R1 3&4 --> A28R1 connected to 1&2 and A29R1 to 3&4 --> numbers correct now!!
Note on D15: effinger: Card replaced because channels 3 -6 blocked to negative: BLECF_0283 replaced by BLECF_0212
Note on AC15: effinger: Card replaced because channels 3 -6 blocked to negative: BLECF_0283 replaced by BLECF_0212
Note on AC14: effinger: 18.8.08 CFC was not working --> reset --> no LF LF
Note on AC12: effinger: 18.8.08 CFC was not working --> -5V broken (cable broken on the backplane) repaired --> no LF LF

## Point 2
| Tunnel - position | [13 empty] | Surface position | [25 empty] | LEVEL | [8 empty] | Signal
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | Comment | ERRA | ERRB | ERRA | ERRB | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | HV | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Comment | Reset | Capa | 220V | Noise | HV | Source | other
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec. | 23.06.08 |  |  |  |  | 11.08.09 |  | 26-27.8.08 | [25 empty] | tested | tested |  | levels | tested | test | Comments
1 | BJBHT.A33L2 | BLECF_ | 399 | 0x | 018F | 6 | Arc | BYPLM.A33L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [8 empty] | while checking the | OK | OK | OK
2 | BJBHT.A32L2 | BLECF_ | 348 | 0x | 015C | 6 | Arc | BYPLM.A32L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | ↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X (ch1 FF) | ok | ok | ok | ok | ok | ok | ok | ok | [10 empty] | OK | OK | OK |  |  |  | HV not o.k
3 | BJBHT.A31L2 | BLECF_ | 497 | 0x | 01F1 | 6 | Arc | BYPLM.A31L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | OK | OK |  | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [8 empty] | signal | OK | OK | OK
4 | BJBHT.A30L2 | BLECF_ | 432 | 0x | 01B0 | 6 | Arc | BYPLM.A30L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | OK | OK |  | OK | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
5 | BJBHT.A29L2 | BLECF_ | 520 | 0x | 0208 | 6 | Arc | BYPLM.A29L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | OK | OK |  | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [8 empty] | all level went on | OK | OK | OK
6 | BJBHT.A28L2 | BLECF_ | 439 | 0x | 01B7 | 6 | Arc | BYPLM.A28L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | OK | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | FF (ch7-8 0X) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
7 | BJBHT.A27L2 | BLECF_ | 342 | 0x | 0156 | 6 | Arc | BYPLM.A27L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [8 empty] | except card 15 1-8 | OK | OK | OK
8 | BJBHT.A26L2 | BLECF_ | 460 | 0x | 01CC | 6 | Arc | BYPLM.A26L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ↑↑ | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
9 | BJBHT.A25L2 | BLECF_ | 388 | 0x | 0184 | 6 | Arc | BYPLM.A25L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | FF (ch7-8 0X) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
10 | BJBHT.A24L2 | BLECF_ | 440 | 0x | 01B8 | 6 | Arc | BYPLM.A24L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ↑↑ | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
11 | BJBHT.A23L2 | BLECF_ | 468 | 0x | 01D4 | 6 | Arc | BYPLM.A23L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | FF (ch7-8 0X) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
12 | BJBHT.A22L2 | BLECF_ | 456 | 0x | 01C8 | 6 | Arc | BYPLM.A22L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | OK | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | noise seen while irradiation
13 | BJBHT.A21L2 | BLECF_ | 111 | 0x | 006F | 6 | Arc | BYPLM.A21L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | FF (ch7-8 0X) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
14 | BJBHT.A20L2 | BLECF_ | 11 | 0x | 000B | 6 | Arc | BYPLM.A20L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | ↑ 83 | OK |  | ↑ | OK | ↑ | OK | ↑ | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
15 | BJBHT.A19L2 | BLECF_ | 715 | 0x | 02CB | 6 | Arc | BYPLM.A19L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
16 | BJBHT.A18L2 | BLECF_ | 366 | 0x | 016E | 6 | Arc | BYPLM.A18L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ↑ 2AEA0 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
17 | BJBHT.A17L2 | BLECF_ | 427 | 0x | 01AB | 6 | Arc | BYPLM.A17L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | ↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
18 | BJBHT.A16L2 | BLECF_ | 385 | 0x | 0181 | 6 | Arc | BYPLM.A16L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | ↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
19 | BJBHT.A15L2 | BLECF_ | 444 | 0x | 01BC | 6 | Arc | BYPLM.A15L2 | 2 | L |  |  |  | 5-6 | SR2 | BY05 | F | 4 | 5-6 | BY02 | L | 1 | 9 | 6 | 9-16 | LF | LF |  | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
20 | BJBHT.A14L2 | BLECF_ | 380 | 0x | 017C | 6 | Arc | BYPLM.A14L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | OK | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
21 | BJBHT.A13L2 | BLECF_ | 347 | 0x | 015B | 6 | Arc | BYPLM.A13L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
22 | BJBHT.A12L2 | BLECF_ | 200 | 0x | 00C8 | 8 | Arc | BYPLM.A12L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ↑ | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
23 | BJBHT.C11L2 | BLECF_ | 494 | 0x | 01EE | 6 | UA23 | BY02 | 2 | CL | 3 | 2 | C | 19-20 | SR2 | BY05 | C | 4 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | in RSTGOH | 1-250 (0) | [8 empty] | OK | OK | OK |  |  |  | RST_DAC
24 | BJBHT.B11L2 | BLECF_ | 343 | 0x | 0157 | 6 | UA23 | BY02 | 2 | CL | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 10 | 7 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | [10 empty] | OK | OK | OK |  |  |  | HV not o.k
25 | BJBHT.A11L2 | BLECF_ | 378 | 0x | 017A | 6 | UA23 | BY02 | 2 | CL | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 1-8 | OK | OK |  | OK | OK | ↑ | OK | ↑ | OK | N | N | ↑ | OK | ↑ | FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
26 | BJBHT.B10L2 | BLECF_ | 134 | 0x | 0086 | 6 | UA23 | BY02 | 2 | CL | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 1-8 | ↑↑↑ | OK |  | ↑ | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | FF (ch1-2 1X) | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-350 (0) | [8 empty] | OK | OK | OK
27 | BJBHT.A10L2 | BLECF_ | 363 | 0x | 016B | 7 | UA23 | BY02 | 2 | CL | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
28 | BJBHT.A9L2 | BLECF_ | 517 | 0x | 0205 | 6 | UA23 | BY02 | 2 | CL | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | Ch10,12-14 20-100 rest 1-20 (0) | [8 empty] | OK | OK | OK
29 | BJBHT.C8L2 | BLECF_ | 357 | 0x | 0165 | 6 | UA23 | BY02 | 2 | CL | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 6 | 3 | 1-8 | OK | OK |  | OK | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 1X (ch5 2C) | ok | ok | ok | ok | ok | ok | ok | ok | ok | whi1-30 (0) | [8 empty] | OK | OK | OK
30 | BJBHT.B8L2 | BLECF_ | 333 | 0x | 014D | 6 | UA23 | BY02 | 2 | CL | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 9 | 6 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
31 | BJBHT.A8L2 | BLECF_ | 491 | 0x | 01EB | 6 | UA23 | BY02 | 2 | CL | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
32 | BJBHT.A7L2 | BLECF_ | 362 | 0x | 016A | 6 | UA23 | BY02 | 2 | CL | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | 10-90 | [8 empty] | OK | OK | OK
33 | BJBHT.C6L2 | BLECF_ | 189 | 0x | 00BD | 6 | UA23 | BY02 | 2 | CL | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 1-8 | ↑ 329 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-20 | [8 empty] | OK | OK | OK
34 | BJBHT.B6L2 | BLECF_ | 344 | 0x | 0158 | 6 | UA23 | BY02 | 2 | CL | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK
35 | BJBHT.A6L2 | BLECF_ | 474 | 0x | 01DA | 6 | UA23 | BY02 | 2 | CL | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
36 | BJBHT.A5L2 | BLECF_ | 328 | 0x | 0148 | 6 | UA23 | BY02 | 2 | CL | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-30 | [8 empty] | OK | OK | OK
37 | BJBHT.D4L2 | BLECF_ | 371 | 0x | 0173 | 8 | UA23 | BY02 | 2 | CL | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8 | ↑↑↑ | OK |  | ↑ | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | 1-30 | [8 empty] | OK | OK | OK
38 | BJBHT.C4L2 | BLECF_ | 181 | 0x | 00B5 | 1 | UA23 | BY02 | 2 | CL | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8 | OK | OK |  | OK | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 1X (ch1 27) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
39 | BJBHT.B4L2 | BLECF_ | 320 | 0x | 0140 | 4 | UA23 | BY02 | 2 | CL | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
40 | BJBHT.A4L2 | BLECF_ | 516 | 0x | 0204 | 7 | UA23 | BY02 | 2 | CL | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | OK | OK |  | OK | ↑ | ↑ | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
41 | BJBHT.A3L2 | BLECF_ | 324 | 0x | 0144 | 4 | UA23 | BY02 | 2 | CL | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
42 | BJBHT.A2L2 | BLECF_ | 188 | 0x | 00BC | 8 | UA23 | BY02 | 2 | CL | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX-DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
43 | BJBHT.B1L2 | BLECF_ | 510 | 0x | 01FE | 5 | UA23 | BY02 | 2 | CL | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | HV not o.k
44 | BJBHT.A1L2 | BLECF_ | 476 | 0x | 01DC | 1 | UA23 | BY02 | 2 | CL | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | OK | OK |  | OK | OK | ↑ | OK | ↑ | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | DAC_RST
45 | BJBHT.A1R2 | BLECF_ | 706 | 0x | 02C2 | 1 | UA27 | BY03 | 2 | CR | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | ↑↑↑ | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
46 | BJBHT.B1R2 | BLECF_ | 606 | 0x | 025E | 5 | UA27 | BY03 | 2 | CR | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-CX | ok | bl | ok | ok | ok | bl | bl | ok | in TCFC | [9 empty] | OK | OK | OK
47 | BJBHT.A2R2 | BLECF_ | 675 | 0x | 02A3 | 8 | UA27 | BY03 | 2 | CR | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | OK | OK |  | OK | OK |  |  | ↑ | OK | N | N | ↑ | OK | ↑ | FF | ok | bl | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
48 | BJBHT.A3R2 | BLECF_ | 641 | 0x | 0281 | 4 | UA27 | BY03 | 2 | CR | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X-20 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
49 | BJBHT.A4R2 | BLECF_ | 719 | 0x | 02CF | 5 | UA27 | BY03 | 2 | CR | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | ↑ 7DF | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X-22 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
50 | BJBHT.B4R2 | BLECF_ | 738 | 0x | 02E2 | 2 | UA27 | BY03 | 2 | CR | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-20 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
51 | BJBHT.C4R2 | BLECF_ | 652 | 0x | 028C | 1 | UA27 | BY03 | 2 | CR | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | ↑↑ | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
52 | BJBHT.D4R2 | BLECF_ | 665 | 0x | 0299 | 8 | UA27 | BY03 | 2 | CR | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
53 | BJBHT.A5R2 | BLECF_ | 740 | 0x | 02E4 | 6 | UA27 | BY03 | 2 | CR | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
54 | BJBHT.A6R2 | BLECF_ | 666 | 0x | 029A | 1 | UA27 | BY03 | 2 | CR | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
55 | BJBHT.B6R2 | BLECF_ | 722 | 0x | 02D2 | 7 | UA27 | BY03 | 2 | CR | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 5 | 2 | 1-8 | ↑↑↑ | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
56 | BJBHT.A7R2 | BLECF_ | 658 | 0x | 0292 | 6 | UA27 | BY03 | 2 | CR | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X (ch7 36) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
57 | BJBHT.A8R2 | BLECF_ | 626 | 0x | 0272 | 6 | UA27 | BY03 | 2 | CR | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | ↑↑↑ | OK | Check of TC | ↑↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-13 | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK
58 | BJBHT.B8R2 | BLECF_ | 406 | 0x | 0196 | 6 | UA27 | BY03 | 2 | CR | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16 | OK | ↑↑↑ | Check of TC | OK | ↑↑↑ |  |  | OK | OK | N | N | OK | OK | OK | 1X-21 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
59 | BJBHT.C8R2 | BLECF_ | 4 | 0x | 0004 | 6 | UA27 | BY03 | 2 | CR | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 1-8 | ↑↑ | OK | Check of TC | OK | LF |  |  | OK | OK | N | N | OK | OK | OK | EX-FX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
60 | BJBHT.A9R2 | BLECF_ | 279 | 0x | 0117 | 6 | UA27 | BY03 | 2 | CR | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 10 | 7-8 | BY02 | R | 3 | 6 | 3 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | Y | N | OK | OK | OK | [19 empty] | OK | OK | OK
61 | BJBHT.A10R2 | BLECF_ | 97 | 0x | 0061 | 6 | UA27 | BY03 | 2 | CR | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 18 | 14 | 1-8 | ↑↑↑ | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-20 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | GND problem ?!
62 | BJBHT.B10R2 | BLECF_ | 721 | 0x | 02D1 | 6 | UA27 | BY03 | 2 | CR | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 7 | 4 | 1-8 | OK | OK |  | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | CX | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
63 | BJBHT.A11R2 | BLECF_ | 128 | 0x | 0080 | 6 | UA27 | BY03 | 2 | CR | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X (ch4 23) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
64 | BJBHT.B11R2 | BLECF_ | 627 | 0x | 0273 | 6 | UA27 | BY03 | 2 | CR | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8 | ↑ | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
65 | BJBHT.C11R2 | BLECF_ | 689 | 0x | 02B1 | 6 | UA27 | BY03 | 2 | CR | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 7 | 4 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | RSTGOH | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
66 | BJBHT.A12R2 | BLECF_ | 681 | 0x | 02A9 | 8 | Arc | BYPLM.A12R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | ↑↑ | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
67 | BJBHT.A13R2 | BLECF_ | 687 | 0x | 02AF | 6 | Arc | BYPLM.A13R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
68 | BJBHT.A14R2 | BLECF_ | 684 | 0x | 02AC | 6 | Arc | BYPLM.A14R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | OK | OK |  | LF | LF |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | CH1 22.6pA DAC0
69 | BJBHT.A15R2 | BLECF_ | 484 | 0x | 01E4 | 6 | Arc | BYPLM.A15R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
70 | BJBHT.A16R2 | BLECF_ | 686 | 0x | 02AE | 6 | Arc | BYPLM.A16R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
71 | BJBHT.A17R2 | BLECF_ | 143 | 0x | 008F | 6 | Arc | BYPLM.A17R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
72 | BJBHT.A18R2 | BLECF_ | 463 | 0x | 01CF | 6 | Arc | BYPLM.A18R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | ↑ 10 | OK |  | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | CH3,4 24.8/24.3pA DAC0
73 | BJBHT.A19R2 | BLECF_ | 379 | 0x | 017B | 6 | Arc | BYPLM.A19R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
74 | BJBHT.A20R2 | BLECF_ | 382 | 0x | 017E | 6 | Arc | BYPLM.A20R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | OK | ↑ 30 |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK |  |  |  | Ch4 27,4pA DAC 24
75 | BJBHT.A21R2 | BLECF_ | 451 | 0x | 01C3 | 6 | Arc | BYPLM.A21R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
76 | BJBHT.A22R2 | BLECF_ | 404 | 0x | 0194 | 6 | Arc | BYPLM.A22R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | OK | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
77 | BJBHT.A23R2 | BLECF_ | 84 | 0x | 0054 | 6 | Arc | BYPLM.A23R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
78 | BJBHT.A24R2 | BLECF_ | 418 | 0x | 01A2 | 6 | Arc | BYPLM.A24R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | ↑↑↑ | OK |  | ↑ | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
79 | BJBHT.A25R2 | BLECF_ | 390 | 0x | 0186 | 6 | Arc | BYPLM.A25R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
80 | BJBHT.A26R2 | BLECF_ | 214 | 0x | 00D6 | 6 | Arc | BYPLM.A26R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
81 | BJBHT.A27R2 | BLECF_ | 694 | 0x | 02B6 | 6 | Arc | BYPLM.A27R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
82 | BJBHT.A28R2 | BLECF_ | 685 | 0x | 02AD | 6 | Arc | BYPLM.A28R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | ↑ 1 | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
83 | BJBHT.A29R2 | BLECF_ | 725 | 0x | 02D5 | 6 | Arc | BYPLM.A29R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0x-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
84 | BJBHT.A30R2 | BLECF_ | 195 | 0x | 00C3 | 6 | Arc | BYPLM.A30R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | OK | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
85 | BJBHT.A31R2 | BLECF_ | 690 | 0x | 02B2 | 6 | Arc | BYPLM.A31R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
86 | BJBHT.A32R2 | BLECF_ | 360 | 0x | 0168 | 6 | Arc | BYPLM.A32R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | ↑↑ | OK |  | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
87 | BJBHT.A33R2 | BLECF_ | 23 | 0x | 0017 | 6 | Arc | BYPLM.A33R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch 7 23.3pA DAC0
88 | BJBHT.A34R2 | BLECF_ | 724 | 0x | 02D4 | 6 | Arc | BYPLM.A34R2 | 2 | R |  |  |  | 7-9 | SR2 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
last entry 17/6/08
[9 empty] | Transformer soldered | [16 empty] | ↑↑↑ | CRC Error consantly fast increasing
[26 empty] | ↑↑ | CRC Error sometimes fast increasing
[26 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number correct | [19 empty] | LF | Lost Frame
|  |  |  | 0x | 0000 | Card number needs to be checked | [19 empty] | no TC | No Threshold Comparator installed
|  |  |  | 0x | 0000 | Card number unknown need to be checked in the tunnel | [20 empty] | Opitcal link is working with Test system
|  | BLECF_ | 656 | Card was exchanged, Data base need to be checked for the correct number
Note on D105: effinger: 565 original number chnaged to 656
Note on F81: effinger: Checked-----> BLECF_0382 installed! Wrong number 0x017E (382) instead of 0x0188 (392) ; to be check in the tunnel; Card number checked: --> BLECF_0382 installed; sent number 382
Note on AA63: effinger: ERROR C
Note on AB64: effinger: ERROR C ERROR CID
Note on AD18: effinger: 248uW also some FID errors
Note on AD91: effinger: 261uW also some FID errors
Note on AD63: effinger: 245uW ERROR C; Error FID
Note on AE64: effinger: 327uW ERROR C ERROR CID ERROR FID
Note on AC7: effinger: CFC-Card not working: ->reset -> o.k. no LF on both
Note on AC9: effinger: CFC-Card not working: ->reset -> o.k. no LF on both
Note on AC23: effinger: fibre connected to 5&6 -> reconnected to 7&8; CFC-Card working --> fibre 7 & 8 broken -> reconnected to 5 & 6
Note on AC65: effinger: CFC-Card working: -> o.k. no LF on both
Note on AC75: effinger: CFC - Card working; no LF ; one Patch cord broken --> no LF on the both
Note on AC81: effinger: Card number checked: --> BLECF_0382 installed; sent number 382
Note on AD7: effinger: 242uW
Note on AE7: effinger: 309uW
Note on AD9: effinger: 275uW
Note on AE9: effinger: 226uW
Note on AD10: effinger: 260uW
Note on AE10: effinger: 251uW
Note on AD12: effinger: 250uW
Note on AE12: effinger: 259uW
Note on AD14: effinger: 258uW
Note on AE14: effinger: 301uW
Note on AD16: effinger: 310uW
Note on AE16: effinger: 265uW
Note on AE18: effinger: 303uW
Note on AD23: effinger: 0W
Note on AE23: effinger: 0W
Note on AD31: effinger: 258uW
Note on AE31: effinger: 261uW
Note on AD42: effinger: 223uW
Note on AE42: effinger: 259uW
Note on AD45: effinger: 224uW
Note on AE45: effinger: 190uW
Note on AD51: effinger: 292uW
Note on AE51: effinger: 210uW
Note on AD55: effinger: 174uW
Note on AE55: effinger: 240uW
Note on AD61: effinger: 284uW
Note on AE61: effinger: 261uW
Note on AD65: effinger: 304uW
Note on AE65: effinger: 232uW
Note on AE63: effinger: 255uW
Note on AD64: effinger: 255uW
Note on AD67: effinger: 314uW
Note on AE67: effinger: 253uW
Note on AD70: effinger: 247uW
Note on AE70: effinger: 296uW
Note on AD73: effinger: 220uW
Note on AE73: effinger: 266uW
Note on AD75: effinger: 265uW
Note on AE75: effinger: 310uW
Note on AD83: effinger: 281uW
Note on AE83: effinger: 268uW
Note on AD85: effinger: 274uW
Note on AE85: effinger: 265uW
Note on AE91: effinger: 299uW
Note on AC63: effinger:
Note on T23: effinger: Fibre 7 & 8 broken, spare fibre used
Note on AF31: effinger: Also some FID errors

## Point 3
| Tunnel - position | [13 empty] | Surface position | [25 empty] | LEVEL | [8 empty] | Signal
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ` | ERRB | Comment | ERRA | ERRB | ERRA | ERRB | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | HV | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Comment | Reset | Capa | 220V | Noise | HV | Source | other
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec. | 23.06.08 |  |  | 25.7.08 |  | 1.8.08 |  | 26-27.8.08 | [25 empty] | tested | tested |  | levels | tested | test | Comments
1 | BJBHT.A33L3 | BLECF_ | 619 | 0x | 026B | 6 | Arc | BYPLM.A33L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
2 | BJBHT.A32L3 | BLECF_ | 601 | 0x | 0259 | 6 | Arc | BYPLM.A32L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | ↑ 2EA0 | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
3 | BJBHT.A31L3 | BLECF_ | 727 | 0x | 02D7 | 6 | Arc | BYPLM.A31L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
4 | BJBHT.A30L3 | BLECF_ | 705 | 0x | 02C1 | 6 | Arc | BYPLM.A30L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | OK | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
5 | BJBHT.A29L3 | BLECF_ | 646 | 0x | 0286 | 6 | Arc | BYPLM.A29L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
6 | BJBHT.A28L3 | BLECF_ | 702 | 0x | 02BE | 6 | Arc | BYPLM.A28L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 6X (ch5 75) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
7 | BJBHT.A27L3 | BLECF_ | 607 | 0x | 025F | 6 | Arc | BYPLM.A27L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
8 | BJBHT.A26L3 | BLECF_ | 623 | 0x | 026F | 6 | Arc | BYPLM.A26L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
9 | BJBHT.A25L3 | BLECF_ | 717 | 0x | 02CD | 6 | Arc | BYPLM.A25L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
10 | BJBHT.A24L3 | BLECF_ | 319 | 0x | 013F | 6 | Arc | BYPLM.A24L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | OK | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK
11 | BJBHT.A23L3 | BLECF_ | 60 | 0x | 003C | 6 | Arc | BYPLM.A23L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
12 | BJBHT.A22L3 | BLECF_ | 645 | 0x | 0285 | 6 | Arc | BYPLM.A22L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ↑ 8 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | in RSTGOH | [9 empty] | OK | OK | OK
13 | BJBHT.A21L3 | BLECF_ | 741 | 0x | 02E5 | 6 | Arc | BYPLM.A21L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
14 | BJBHT.A20L3 | BLECF_ | 614 | 0x | 0266 | 6 | Arc | BYPLM.A20L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
15 | BJBHT.A19L3 | BLECF_ | 620 | 0x | 026C | 6 | Arc | BYPLM.A19L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
16 | BJBHT.A18L3 | BLECF_ | 630 | 0x | 0276 | 6 | Arc | BYPLM.A18L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | OK | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
17 | BJBHT.A17L3 | BLECF_ | 668 | 0x | 029C | 6 | Arc | BYPLM.A17L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
18 | BJBHT.A16L3 | BLECF_ | 604 | 0x | 025C | 6 | Arc | BYPLM.A16L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | OK | OK |  | ↑ | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
19 | BJBHT.A15L3 | BLECF_ | 622 | 0x | 026E | 6 | Arc | BYPLM.A15L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | GOH_RST
20 | BJBHT.A14L3 | BLECF_ | 305 | 0x | 0131 | 6 | Arc | BYPLM.A14L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
21 | BJBHT.A13L3 | BLECF_ | 600 | 0x | 0258 | 10 | Arc | BYPLM.A13L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
22 | BJBHT.A12L3 | BLECF_ | 703 | 0x | 02BF | 10 | Arc | BYPLM.A12L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ↑ 7 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
23 | BJBHT.B11L3 | BLECF_ | 657 | 0x | 0291 | 7 | UJ33 | BY02 | 3 | L | 2 | 8 | C | 11-12 | SR3 | BY05 | C | 3 | 11-12 | BY02 | L | 1 | 7 | 4 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK |  |  |  | DAC_RST
24 | BJBHT.A11L3 | BLECF_ | 654 | 0x | 028E | 12 | UJ33 | BY02 | 3 | L | 2 | 7 | C | 9-10 | SR3 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 8 | 5 | 9-16 | no TC | no TC |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
25 | BJBHT.A10L3 | BLECF_ | 674 | 0x | 02A2 | 11 | UJ33 | BY02 | 3 | L | 2 | 6 | C | 7-8 | SR3 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
26 | BJBHT.B9L3 | BLECF_ | 754 | 0x | 02F2 | 11 | UJ33 | BY02 | 3 | L | 2 | 5 | C | 5-6 | SR3 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | in RSTDAC | [9 empty] | OK | OK | OK |  |  |  | DAC_RST
27 | BJBHT.A9L3 | BLECF_ | 609 | 0x | 0261 | 10 | UJ33 | BY02 | 3 | L | 2 | 4 | C | 3-4 | SR3 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 1-8 | no TC | no TC |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
28 | BJBHT.B8L3 | BLECF_ | 672 | 0x | 02A0 | 6 | UJ33 | BY02 | 3 | L | 2 | 3 | C | 1-2 | SR3 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 6 | 3 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK
29 | BJBHT.A8L3 | BLECF_ | 574 | 0x | 023E | 6 | UJ33 | BY02 | 3 | L | 2 | 2 | B | 23-24 | SR3 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 7 | 4 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | in RSTDAC | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
30 | BJBHT.B7L3 | BLECF_ | 576 | 0x | 0240 | 7 | UJ33 | BY02 | 3 | L | 2 | 1 | B | 21-22 | SR3 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
31 | BJBHT.A7L3 | BLECF_ | 637 | 0x | 027D | 1 | UJ33 | BY02 | 3 | L | 1 | 10 | B | 19-20 | SR3 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
32 | BJBHT.C6L3 | BLECF_ | 211 | 0x | 00D3 | 6 | UJ33 | BY02 | 3 | L | 1 | 9 | B | 17-18 | SR3 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 5 | 2 | 1-8 | OK | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
33 | BJBHT.B6L3 | BLECF_ | 629 | 0x | 0275 | 4 | UJ33 | BY02 | 3 | L | 1 | 8 | B | 15-16 | SR3 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 6 | 3 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
34 | BJBHT.A6L3 | BLECF_ | 536 | 0x | 0218 | 4 | UJ33 | BY02 | 3 | L | 1 | 7 | B | 13-14 | SR3 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 6 | 3 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
35 | BJBHT.D5L3 | BLECF_ | 175 | 0x | 00AF | 2 | UJ33 | BY02 | 3 | L | 1 | 6 | B | 11-12 | SR3 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
36 | BJBHT.C5L3 | BLECF_ | 577 | 0x | 0241 | 8 | UJ33 | BY02 | 3 | L | 1 | 5 | B | 9-10 | SR3 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
37 | BJBHT.B5L3 | BLECF_ | 575 | 0x | 023F | 6 | UJ33 | BY02 | 3 | L | 1 | 4 | B | 7-8 | SR3 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 1-8 | OK | OK |  | OK | OK | OK | ↑ | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
38 | BJBHT.A5L3 | BLECF_ | 602 | 0x | 025A | 6 | UJ33 | BY02 | 3 | L | 1 | 3 | B | 5-6 | SR3 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
39 | BJBHT.B4L3 | BLECF_ | 711 | 0x | 02C7 | 2 | UJ33 | BY02 | 3 | L | 1 | 2 | B | 3-4 | SR3 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 1-8 | OK | OK |  | OK | OK | OK | ↑ | ↑ | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
40 | BJBHT.A4L3 | BLECF_ | 729 | 0x | 02D9 | 8 | UJ33 | BY02 | 3 | L | 1 | 1 | B | 1-2 | SR3 | BY05 | C | 1 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X (ch6 3A) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
41 | BJBHT.A4R3 | BLECF_ | 543 | 0x | 021F | 8 | UJ33 | BY03 | 3 | R | 1 | 1 | C | 13-14 | SR3 | BY05 | C | 4 | 1-2 | BY02 | R | 3 | 4 | 1 | 1-8 | ↑ 32 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | little noise
42 | BJBHT.B4R3 | BLECF_ | 677 | 0x | 02A5 | 2 | UJ33 | BY03 | 3 | R | 1 | 2 | C | 15-16 | SR3 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 13 | 9 | 1-8 | ↑ AABF8 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0x (ch1 2A) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | little noise
43 | BJBHT.A5R3 | BLECF_ | 182 | 0x | 00B6 | 6 | UJ33 | BY03 | 3 | R | 1 | 3 | C | 17-18 | SR3 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-2X | bl | bl | bl | bl | bl | bl | bl | bl | ok | [9 empty] | OK | OK | OK |  |  |  | big noise -> solved ground problem
44 | BJBHT.B5R3 | BLECF_ | 326 | 0x | 0146 | 6 | UJ33 | BY03 | 3 | R | 1 | 4 | C | 19-20 | SR3 | BY05 | C | 4 | 7-8 | BY02 | C | 2 | 14 | 10 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X (ch1 27) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | little noise
45 | BJBHT.C5R3 | BLECF_ | 487 | 0x | 01E7 | 8 | UJ33 | BY03 | 3 | R | 1 | 5 | C | 21-22 | SR3 | BY05 | C | 4 | 9-10 | BY02 | R | 3 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | big noise -> solved ground problem
46 | BJBHT.D5R3 | BLECF_ | 393 | 0x | 0189 | 2 | UJ33 | BY03 | 3 | R | 1 | 6 | C | 23-24 | SR3 | BY05 | C | 4 | 11-12 | BY02 | C | 2 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK |  |  |  | big noise -> solved ground problem
47 | BJBHT.A6R3 | BLECF_ | 541 | 0x | 021D | 4 | UJ33 | BY03 | 3 | R | 1 | 7 | D | 1-2 | SR3 | BY05 | C | 5 | 1-2 | BY02 | C | 2 | 15 | 11 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | little noise
48 | BJBHT.B6R3 | BLECF_ | 341 | 0x | 0155 | 4 | UJ33 | BY03 | 3 | R | 1 | 8 | D | 3-4 | SR3 | BY05 | C | 5 | 3-4 | BY02 | C | 2 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
49 | BJBHT.C6R3 | BLECF_ | 339 | 0x | 0153 | 6 | UJ33 | BY03 | 3 | R | 1 | 9 | D | 5-6 | SR3 | BY05 | C | 5 | 5-6 | BY02 | R | 3 | 5 | 2 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | HV-Test on, TEST_CFC
50 | BJBHT.A7R3 | BLECF_ | 367 | 0x | 016F | 1 | UJ33 | BY03 | 3 | R | 1 | 10 | D | 7-8 | SR3 | BY05 | C | 5 | 7-8 | BY02 | C | 2 | 16 | 12 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | HV-Test on
51 | BJBHT.B7R3 | BLECF_ | 709 | 0x | 02C5 | 7 | UJ33 | BY03 | 3 | R | 2 | 1 | D | 9-10 | SR3 | BY05 | C | 5 | 9-10 | BY02 | R | 3 | 5 | 2 | 9-16 | LF | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | HV-Test on
52 | BJBHT.A8R3 | BLECF_ | 502 | 0x | 01F6 | 6 | UJ33 | BY03 | 3 | R | 2 | 2 | D | 11-12 | SR3 | BY05 | C | 5 | 11-12 | BY02 | C | 2 | 16 | 12 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | little noise
53 | BJBHT.B8R3 | BLECF_ | 183 | 0x | 00B7 | 6 | UJ33 | BY03 | 3 | R | 2 | 3 | D | 13-14 | SR3 | BY05 | C | 6 | 1-2 | BY02 | R | 3 | 6 | 3 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
54 | BJBHT.A9R3 | BLECF_ | 375 | 0x | 0177 | 10 | UJ33 | BY03 | 3 | R | 2 | 4 | D | 15-16 | SR3 | BY05 | C | 6 | 3-4 | BY02 | C | 2 | 17 | 13 | 1-8 | no TC | no TC |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | little noise
55 | BJBHT.B9R3 | BLECF_ | 410 | 0x | 019A | 11 | UJ33 | BY03 | 3 | R | 2 | 5 | D | 17-18 | SR3 | BY05 | C | 6 | 5-6 | BY02 | R | 3 | 6 | 3 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
56 | BJBHT.A10R3 | BLECF_ | 480 | 0x | 01E0 | 11 | UJ33 | BY03 | 3 | R | 2 | 6 | D | 19-20 | SR3 | BY05 | C | 6 | 7-8 | BY02 | R | 3 | 7 | 4 | 1-8 | ↑ 305A | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-2X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK
57 | BJBHT.A11R3 | BLECF_ | 562 | 0x | 0232 | 12 | UJ33 | BY03 | 3 | R | 2 | 7 | D | 21-22 | SR3 | BY05 | C | 6 | 9-10 | BY02 | C | 2 | 17 | 13 | 9-16 | no TC | no TC |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | medium noise, cabling checked
58 | BJBHT.B11R3 | BLECF_ | 459 | 0x | 01CB | 7 | UJ33 | BY03 | 3 | R | 2 | 8 | D | 23-24 | SR3 | BY05 | C | 6 | 11-12 | BY02 | R | 3 | 7 | 4 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | bl | bl | bl | bl | bl | bl | bl | bl | ok | [9 empty] | OK | OK | OK |  |  |  | medium noise, cabling checked, LEVEL
59 | BJBHT.A12R3 | BLECF_ | 472 | 0x | 01D8 | 10 | Arc | BYPLM.A12R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | OK | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
60 | BJBHT.A13R3 | BLECF_ | 611 | 0x | 0263 | 10 | Arc | BYPLM.A13R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
61 | BJBHT.A14R3 | BLECF_ | 504 | 0x | 01F8 | 6 | Arc | BYPLM.A14R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
62 | BJBHT.A15R3 | BLECF_ | 499 | 0x | 01F3 | 6 | Arc | BYPLM.A15R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-15 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
63 | BJBHT.A16R3 | BLECF_ | 514 | 0x | 0202 | 6 | Arc | BYPLM.A16R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
64 | BJBHT.A17R3 | BLECF_ | 493 | 0x | 01ED | 6 | Arc | BYPLM.A17R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
65 | BJBHT.A18R3 | BLECF_ | 496 | 0x | 01F0 | 6 | Arc | BYPLM.A18R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | OK | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X-55 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
66 | BJBHT.A19R3 | BLECF_ | 314 | 0x | 013A | 6 | Arc | BYPLM.A19R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X-55 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
67 | BJBHT.A20R3 | BLECF_ | 523 | 0x | 020B | 6 | Arc | BYPLM.A20R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | ↑↑↑ | ↑ 2 |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X-64 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
68 | BJBHT.A21R3 | BLECF_ | 185 | 0x | 00B9 | 6 | Arc | BYPLM.A21R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0x-27 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
69 | BJBHT.A22R3 | BLECF_ | 387 | 0x | 0183 | 6 | Arc | BYPLM.A22R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
70 | BJBHT.A23R3 | BLECF_ | 485 | 0x | 01E5 | 6 | Arc | BYPLM.A23R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
71 | BJBHT.A24R3 | BLECF_ | 309 | 0x | 0135 | 6 | Arc | BYPLM.A24R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | ↑↑↑ | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X-70 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
72 | BJBHT.A25R3 | BLECF_ | 184 | 0x | 00B8 | 6 | Arc | BYPLM.A25R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
73 | BJBHT.A26R3 | BLECF_ | 337 | 0x | 0151 | 6 | Arc | BYPLM.A26R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
74 | BJBHT.A27R3 | BLECF_ | 519 | 0x | 0207 | 6 | Arc | BYPLM.A27R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X-7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
75 | BJBHT.A28R3 | BLECF_ | 448 | 0x | 01C0 | 6 | Arc | BYPLM.A28R3 | 3 | R |  |  |  | 6-7 | SR3 | BY05 | H | 12 | 6-7 | BY02 | R | 3 | 17 | 13 | 1-8 | OK | LF |  | OK | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
76 | BJBHT.A29R3 | BLECF_ | 489 | 0x | 01E9 | 6 | Arc | BYPLM.A29R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
77 | BJBHT.A30R3 | BLECF_ | 401 | 0x | 0191 | 6 | Arc | BYPLM.A30R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | OK | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 00 | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
78 | BJBHT.A31R3 | BLECF_ | 488 | 0x | 01E8 | 6 | Arc | BYPLM.A31R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
79 | BJBHT.A32R3 | BLECF_ | 407 | 0x | 0197 | 6 | Arc | BYPLM.A32R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
80 | BJBHT.A33R3 | BLECF_ | 453 | 0x | 01C5 | 6 | Arc | BYPLM.A33R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
81 | BJBHT.A34R3 | BLECF_ | 109 | 0x | 006D | 6 | Arc | BYPLM.A34R3 | 3 | R |  |  |  | 6-7 | SR3 | BY05 | I | 12 | 6-7 | BY02 | R | 3 | 20 | 16 | 1-8 | OK | LF |  | OK | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 4X-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
last entry 19/6/08
| Door | [7 empty] | Transformer soldered | [16 empty] | ↑↑↑ | CRC Error consantly fast increasing
[26 empty] | ↑↑ | CRC Error sometimes fast increasing
[26 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number correct | [19 empty] | LF | Lost Frame
|  |  |  | 0x | 0000 | Card number needs to be checked | [19 empty] | no TC | No Threshold Comparator installed
|  |  |  | 0x | 0000 | Card number unknown need to be checked in the tunnel | [20 empty] | Opitcal link is working with Test system
|  | BLECF_ | 656 | Card was exchanged, Data base need to be checked for the correct number
Note on D98: effinger: 565 original number chnaged to 656
Note on AA6: effinger: 217uW
Note on AB6: effinger: 260uW
Note on AA16: effinger: 256uW
Note on AB16: effinger: 295uW
Note on AA26: effinger: 285uW
Note on AB26: effinger: 259uW
Note on AA29: effinger: 199uW
Note on AB29: effinger: 201uW
Note on AA32: effinger: 195uW
Note on AB32: effinger: 203uW
Note on AA47: effinger: 155uW
Note on AB47: effinger: 231uW
Note on AA48: effinger: 192uW
Note on AB48: effinger: 217uW
Note on AA57: effinger: 4uW
Note on AB57: effinger: 230uW
Note on AA60: effinger: 174uW
Note on AB60: effinger: 210uW
Note on AA62: effinger: 197uW
Note on AB62: effinger: 196uW
Note on AA63: effinger: 201uW
Note on AB63: effinger: 207uW
Note on AA68: effinger: 210uW
Note on AB68: effinger: 238uW
Note on AA74: effinger: 199uW
Note on AB74: effinger: 190uW
Note on AA78: effinger: 176uW
Note on AB78: effinger: 185uW
Note on AA82: effinger: 192uW
Note on AA88: effinger: 162uW
Note on AB82: effinger: 0W TB not o.k.
Note on AB88: effinger: 0W TB not o.k.
Note on AD29: effinger: Problem TC
Note on AD57: effinger: some LF, ERROR FID
Note on AC88: effinger: in the tunnel with TB o.k. --> Fibre broken tunnel to surface
Note on AC82: effinger: in the tunnel with TB o.k. --> Fibre broken tunnel to surface
Note on AF29: effinger: Problem TC --> card pluged and unpluged --> o.k now
Note on AC57: effinger: CFC-Card changed because of very low laser power and patch cord changed because of high attenuation
Note on AF32: effinger: Problem TC --> card pluged and unpluged --> o.k now
Note on AG32: effinger: Problem TC --> card pluged and unpluged --> o.k now
Note on AG29: effinger: Problem TC --> card pluged and unpluged --> o.k now
Note on AC30: effinger: CFC card was blocked -> reset done --> o.k. no LF

## Point 4
| Tunnel - position | [13 empty] | Surface position | [25 empty] | LEVEL | [8 empty] | Signal
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | Comment | ERRA | ERRB | ERRA | ERRB | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | HV | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Comment | Reset | Capa | 220V | Noise | HV | Source | other
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec. | 23.6.08 |  |  | 15.7.08 |  | 25.7.08 |  | 26-27.8.08 | [25 empty] | tested | tested |  | levels | tested | test | Comments
1 | BJBHT.A33L4 | BLECF_ | 452 | 0x | 01C4 | 6 | Arc | BYPLM.A33L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | OK | OK |  | OK | OK | OK | OK | ↑↑↑ | OK | N | N | ↑ | OK | OK | DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
2 | BJBHT.A32L4 | BLECF_ | 748 | 0x | 02EC | 6 | Arc | BYPLM.A32L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
3 | BJBHT.A31L4 | BLECF_ | 616 | 0x | 0268 | 6 | Arc | BYPLM.A31L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | ↑↑ | LF |  | ↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | DX-EX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
4 | BJBHT.A30L4 | BLECF_ | 662 | 0x | 0296 | 6 | Arc | BYPLM.A30L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | BX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
5 | BJBHT.A29L4 | BLECF_ | 38 | 0x | 0026 | 6 | Arc | BYPLM.A29L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | OK | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | CX-DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
6 | BJBHT.A28L4 | BLECF_ | 76 | 0x | 004C | 6 | Arc | BYPLM.A28L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | LF | LF |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK
7 | BJBHT.A27L4 | BLECF_ | 572 | 0x | 023C | 6 | Arc | BYPLM.A27L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ↑↑ | OK |  | ↑↑↑ | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 9X-AX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
8 | BJBHT.A26L4 | BLECF_ | 318 | 0x | 013E | 6 | Arc | BYPLM.A26L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | BX-C1 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
9 | BJBHT.A25L4 | BLECF_ | 613 | 0x | 0265 | 6 | Arc | BYPLM.A25L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ↑↑↑ | OK |  | OK | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | DX-EX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
10 | BJBHT.A24L4 | BLECF_ | 655 | 0x | 028F | 6 | Arc | BYPLM.A24L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X-9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
11 | BJBHT.A23L4 | BLECF_ | 631 | 0x | 0277 | 6 | Arc | BYPLM.A23L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ↑↑↑ | OK |  | ↑↑↑ | OK | OK | OK | ↑ | OK | N | N | ↑ | OK | OK | 2X-32 | ok | ok | ok | ok | ok | ok | ok | ok | in TCFC | [9 empty] | OK | OK | OK
12 | BJBHT.A22L4 | BLECF_ | 743 | 0x | 02E7 | 6 | Arc | BYPLM.A22L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX-DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
13 | BJBHT.A21L4 | BLECF_ | 664 | 0x | 0298 | 6 | Arc | BYPLM.A21L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | ↑↑↑ | OK |  | ↑↑↑ | OK | OK | OK | ↑ | OK | N | N | ↑ | OK | OK | DX-EX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
14 | BJBHT.A20L4 | BLECF_ | 471 | 0x | 01D7 | 6 | Arc | BYPLM.A20L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
15 | BJBHT.A19L4 | BLECF_ | 552 | 0x | 0228 | 6 | Arc | BYPLM.A19L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ↑ 2EF | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX-DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
16 | BJBHT.A18L4 | BLECF_ | 395 | 0x | 018B | 6 | Arc | BYPLM.A18L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX-DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
17 | BJBHT.A17L4 | BLECF_ | 615 | 0x | 0267 | 6 | Arc | BYPLM.A17L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | OK | OK |  | ↑↑↑ | OK | ↑ | OK | ↑ | OK | N | N | OK | OK | OK | DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
18 | BJBHT.A16L4 | BLECF_ | 591 | 0x | 024F | 6 | Arc | BYPLM.A16L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
19 | BJBHT.A15L4 | BLECF_ | 642 | 0x | 0282 | 6 | Arc | BYPLM.A15L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | ↑↑↑ | OK |  | ↑↑↑ | OK | OK | OK | ↑ | OK | N | N | ↑ | OK | OK | DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
20 | BJBHT.A14L4 | BLECF_ | 670 | 0x | 029E | 6 | Arc | BYPLM.A14L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
21 | BJBHT.A13L4 | BLECF_ | 599 | 0x | 0257 | 6 | Arc | BYPLM.A13L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ↑ 151 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | CX-DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
22 | BJBHT.A12L4 | BLECF_ | 659 | 0x | 0293 | 6 | Arc | BYPLM.A12L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | DX-E4 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
23 | BJBHT.B5L4 |  |  | 0x | 0000 | 0 | UA43 | BY12 | 4 | CL | 1 | 9 | A | 17-18 | SX4 | BY04 | C | 2 | 5-6
24 | BJBHT.A11L4 | BLECF_ | 643 | 0x | 0283 | 6 | UA43 | BY12 | 4 | CL | 1 | 8 | A | 15-16 | SX4 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 7 | 4 | 1-8 | LF | LF | Fibre missing, check surface TB, PM | LF | LF | OK | OK | ↑ | OK | N | N | OK | OK | OK | 3X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
25 | BJBHT.A10L4 | BLECF_ | 578 | 0x | 0242 | 6 | UA43 | BY12 | 4 | CL | 1 | 7 | A | 13-14 | SX4 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 6 | 3 | 9-16 | LF | LF | Fibre missing, check surface TB, PM | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
26 | BJBHT.A9L4 | BLECF_ | 408 | 0x | 0198 | 6 | UA43 | BY12 | 4 | CL | 1 | 6 | A | 11-12 | SX4 | BY05 | C | 1 | 11-12 | BY02 | L | 1 | 6 | 3 | 1-8 | ↑↑↑ | OK |  | ↑↑ | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 3X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
27 | BJBHT.A8L4 | BLECF_ | 651 | 0x | 028B | 6 | UA43 | BY12 | 4 | CL | 1 | 5 | A | 9-10 | SX4 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
28 | BJBHT.A7L4 | BLECF_ | 386 | 0x | 0182 | 6 | UA43 | BY12 | 4 | CL | 1 | 4 | A | 7-8 | SX4 | BY05 | C | 1 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | ↑↑↑ | OK |  | ↑↑ | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
29 | BJBHT.A6L4 | BLECF_ | 157 | 0x | 009D | 6 | UA43 | BY12 | 4 | CL | 1 | 3 | A | 5-6 | SX4 | BY05 | C | 1 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
30 | BJBHT.C5L4 | BLECF_ | 744 | 0x | 02E8 | 8 | UA43 | BY12 | 4 | CL | 1 | 2 | A | 3-4 | SX4 | BY05 | C | 1 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8 | ↑ 17FB | OK |  | ↑ | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
31 | BJBHT.B5L4 | BLECF_ | 186 | 0x | 00BA | 2 | UA43 | BY12 | 4 | CL | 1 | 1 | A | 1-2 | SX4 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | ↑ 7D6 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-2X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
32 | BJBHT.B5R4 | BLECF_ | 8 | 0x | 0008 | 2 | UA47 | BY09 | 4 | CR | 1 | 1 | A | 1-2 | SX4 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | ↑ 1131 | OK | wrong C no, check tunnel | ↑↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
33 | BJBHT.C5R4 | BLECF_ | 5 | 0x | 0005 | 8 | UA47 | BY09 | 4 | CR | 1 | 2 | A | 3-4 | SX4 | BY05 | C | 12 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8 | OK | OK | wrong C no, check tunnel | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
34 | BJBHT.A6R4 | BLECF_ | 316 | 0x | 013C | 6 | UA47 | BY09 | 4 | CR | 1 | 3 | A | 5-6 | SX4 | BY05 | C | 12 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | LF | LF | no link, check tunnel | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
35 | BJBHT.A7R4 | BLECF_ | 556 | 0x | 022C | 6 | UA47 | BY09 | 4 | CR | 1 | 4 | A | 7-8 | SX4 | BY05 | C | 12 | 7-8 | BY02 | R | 3 | 5 | 2 | 1-8 | LF | LF | no link, check tunnel | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
36 | BJBHT.A8R4 | BLECF_ | 584 | 0x | 0248 | 6 | UA47 | BY09 | 4 | CR | 1 | 5 | A | 9-10 | SX4 | BY05 | C | 12 | 9-10 | BY02 | R | 3 | 5 | 2 | 9-16 | LF | LF | no link, check tunnel | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 20 - 40 bit noise
37 | BJBHT.A9R4 | BLECF_ | 509 | 0x | 01FD | 6 | UA47 | BY09 | 4 | CR | 1 | 6 | A | 11-12 | SX4 | BY05 | C | 12 | 11-12 | BY02 | R | 3 | 6 | 3 | 1-8 | LF | LF | no link, check tunnel | ↑ | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 60 - 80 bit noise
38 | BJBHT.A10R4 | BLECF_ | 7 | 0x | 0007 | 6 | UA47 | BY09 | 4 | CR | 1 | 7 | A | 13-14 | SX4 | BY05 | C | 11 | 1-2 | BY02 | R | 3 | 6 | 3 | 9-16 | OK | OK | wrong C no, check tunnel | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch1 no noise , ch2-8 60 bit noise
39 | BJBHT.A11R4 | BLECF_ | 10 | 0x | 000A | 6 | UA47 | BY09 | 4 | CR | 1 | 8 | A | 15-16 | SX4 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 7 | 4 | 1-8 | ↑↑↑ | OK | wrong C no, check tunnel | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 20 -80 bit noise
40 | BJBHT.A5R4 |  |  | 0x | 0000 | 0 | UA47 | BY09 | 4 | CR | 1 | 9 | A | 17-18 | SX4 | BY05 | C | 11 | 5-6
41 | BJBHT.A12R4 | BLECF_ | 217 | 0x | 00D9 | 6 | Arc | BYPLM.A12R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 7 | 4 | 9-16 | LF | LF | Fibre missing, check surface TB, PM | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-3X-4C | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | reset problen
42 | BJBHT.A13R4 | BLECF_ | 531 | 0x | 0213 | 6 | Arc | BYPLM.A13R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | LF | LF | Fibre missing, check surface TB, PM | LF | LF | ↑ | OK | ↑ | OK | N | N | OK | OK | OK | DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
43 | BJBHT.A14R4 | BLECF_ | 568 | 0x | 0238 | 6 | Arc | BYPLM.A14R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | LF | LF | Fibre missing, check surface TB, PM | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | AX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
44 | BJBHT.A15R4 | BLECF_ | 57 | 0x | 0039 | 6 | Arc | BYPLM.A15R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | LF | LF | Fibre missing, check surface TB, PM | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
45 | BJBHT.A16R4 | BLECF_ | 644 | 0x | 0284 | 6 | Arc | BYPLM.A16R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | LF | LF | Fibre missing, check surface TB, PM | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
46 | BJBHT.A17R4 | BLECF_ | 403 | 0x | 0193 | 6 | Arc | BYPLM.A17R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | OK | ↑↑↑ |  | OK | OK | LF | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
47 | BJBHT.A18R4 | BLECF_ | 558 | 0x | 022E | 6 | Arc | BYPLM.A18R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | LF | LF | Card o.k. | ↑↑↑ | OK | OK | OK | OK | OK | N | Y | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
48 | BJBHT.A19R4 | BLECF_ | 199 | 0x | 00C7 | 6 | Arc | BYPLM.A19R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | OK | OK | Card didn't work, reset than o.k. | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X | bl | bl | bl | bl | bl | bl | bl | bl | in TCFC | [9 empty] | OK | OK | OK |  |  |  | reset problem
49 | BJBHT.A20R4 | BLECF_ | 585 | 0x | 0249 | 6 | Arc | BYPLM.A20R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | ↑↑↑ | OK |  | ↑↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | reset problem
50 | BJBHT.A21R4 | BLECF_ | 315 | 0x | 013B | 6 | Arc | BYPLM.A21R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | ↑↑↑ | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | test-cfc
51 | BJBHT.A22R4 | BLECF_ | 340 | 0x | 0154 | 6 | Arc | BYPLM.A22R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
52 | BJBHT.A23R4 | BLECF_ | 537 | 0x | 0219 | 6 | Arc | BYPLM.A23R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | OK | OK |  | ↑↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
53 | BJBHT.A24R4 | BLECF_ | 405 | 0x | 0195 | 6 | Arc | BYPLM.A24R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
54 | BJBHT.A25R4 | BLECF_ | 483 | 0x | 01E3 | 6 | Arc | BYPLM.A25R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
55 | BJBHT.A26R4 | BLECF_ | 438 | 0x | 01B6 | 6 | Arc | BYPLM.A26R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | LF | OK |  | ↑↑↑ ↑: LF; OK OK: OK; OK OK: LF; OK OK: OK; OK OK: OK; OK OK: OK; N N: Y; N N: N; OK OK: OK; OK OK: OK; OK OK: OK | [9 empty] | OK | OK | OK
56 | BJBHT.A27R4 | BLECF_ | 547 | 0x | 0223 | 6 | Arc | BYPLM.A27R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | LF | LF | Card didn't work, reset than o.k. | ↑↑↑ | OK | OK | OK | ↑ | OK | N | N | ↑ | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | reset problem
57 | BJBHT.A28R4 | BLECF_ | 551 | 0x | 0227 | 6 | Arc | BYPLM.A28R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
58 | BJBHT.A29R4 | BLECF_ | 83 | 0x | 0053 | 6 | Arc | BYPLM.A29R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | OK | OK |  | OK | OK | ↑ | OK | ↑ | OK | N | N | ↑ | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
59 | BJBHT.A30R4 | BLECF_ | 435 | 0x | 01B3 | 6 | Arc | BYPLM.A30R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | DX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | reset problem
60 | BJBHT.A31R4 | BLECF_ | 40 | 0x | 0028 | 6 | Arc | BYPLM.A31R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | ↑ 5 | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
61 | BJBHT.A32R4 | BLECF_ | 192 | 0x | 00C0 | 6 | Arc | BYPLM.A32R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | reset problem
62 | BJBHT.A33R4 | BLECF_ | 582 | 0x | 0246 | 6 | Arc | BYPLM.A33R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | OK | OK |  | ↑↑↑ | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
63 | BJBHT.A34R4 | BLECF_ | 389 | 0x | 0185 | 6 | Arc | BYPLM.A34R4 | 4 | R |  |  |  | 7-9 | SX4 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
last entry 16/7/08
[9 empty] | Transformer soldered | [16 empty] | ↑↑↑ | CRC Error consantly fast increasing
[26 empty] | ↑↑ | CRC Error sometimes fast increasing
[26 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number correct | [19 empty] | LF | Lost Frame
|  |  |  | 0x | 0000 | Card number needs to be checked | [19 empty] | no TC | No Threshold Comparator installed
|  |  |  | 0x | 0000 | Card number unknown need to be checked in the tunnel | [20 empty] | Opitcal link is working with Test system
|  | BLECF_ | 656 | Card was exchanged, Data base need to be checked for the correct number
Note on D80: effinger: 565 original number chnaged to 656
Note on AB51: effinger: Patch fibre missing
Note on AA7: effinger: 325uW
Note on AA11: effinger: 232uW , TB o.k.
Note on AA13: effinger: 181uW , TB o.k.
Note on AA15: effinger: 204uW , TB o.k.
Note on AA17: effinger: 305uW , TB o.k.
Note on AA19: effinger: 269uW , TB o.k.
Note on AA23: effinger: 204uW , TB o.k.
Note on AA25: effinger: 194uW , TB o.k.
Note on AA29: effinger: Patch fibre missing
Note on AB29: effinger: Patch fibre missing
Note on AA30: effinger: Patch fibre missing
Note on AB30: effinger: Patch fibre missing
Note on AA31: effinger: 84uW , TB o.k.
Note on AB31: effinger: 84uW , TB o.k.
Note on AA33: effinger: 83uW , TB o.k.
Note on AA35: effinger: 83uW , TB o.k.
Note on AA36: effinger: 84uW , TB o.k.
Note on AA38: effinger: 82uW , TB o.k.
Note on AA40: effinger: 0uW , no signal, TB not o.k.
Note on AB40: effinger: 0uW , no signal, TB not o.k.
Note on AA41: effinger: 0uW , no signal, TB not o.k.
Note on AA42: effinger: 0uW , no signal, TB not o.k.
Note on AA43: effinger: 0uW , no signal, TB not o.k.
Note on AB41: effinger: 0uW , no signal, TB not o.k.
Note on AB42: effinger: 0uW , no signal, TB not o.k.
Note on AB43: effinger: 0uW , no signal, TB not o.k.
Note on AA45: effinger: 83uW , TB o.k.
Note on AA48: effinger: Patch fibre missing
Note on AA49: effinger: Patch fibre missing
Note on AA50: effinger: Patch fibre missing
Note on AA51: effinger: Patch fibre missing
Note on AB48: effinger: Patch fibre missing
Note on AB49: effinger: Patch fibre missing
Note on AB50: effinger: Patch fibre missing
Note on AB53: effinger: 84uW , TB o.k.
Note on AA54: effinger: 83uW , TB o.k.
Note on AB54: effinger: 83uW , TB o.k.
Note on AA56: effinger: 83uW , TB o.k.
Note on AA57: effinger: 83uW , TB o.k.
Note on AA61: effinger: 84uW , TB o.k.
Note on AA62: effinger: 83uW , TB o.k.
Note on AA67: effinger: 83uW , TB o.k.
Note on AB7: effinger: 0uW , no signal, TB not o.k.
Note on AA10: effinger: 0uW , no signal, TB not o.k.
Note on AB10: effinger: 0uW , no signal, TB not o.k.
Note on AD29: effinger: Patch fibre missing
Note on AE29: effinger: Patch fibre missing
Note on AD30: effinger: Patch fibre missing
Note on AE30: effinger: Patch fibre missing
Note on AD48: effinger: Patch fibre missing
Note on AE48: effinger: Patch fibre missing
Note on AD49: effinger: Patch fibre missing
Note on AE49: effinger: Patch fibre missing
Note on AD50: effinger: Patch fibre missing
Note on AE50: effinger: Patch fibre missing
Note on AD51: effinger: Patch fibre missing
Note on AE51: effinger: Patch fibre missing
Note on AD56: effinger: Error CID, Error FID
Note on AD52: effinger: Patch fibre missing
Note on AE52: effinger: Patch fibre missing
Note on AA52: effinger: Patch fibre missing
Note on AB52: effinger: Patch fibre missing
Note on AC52: effinger: Card No. wrong -> reset -> o.k.
Note on AC48: effinger: Card didn't work, reset than o.k.
Note on AC38: effinger: Patch pannel C12 and C11 inversed -> we exchanged patch cords -> no it is o.k.
Note on AF62: effinger: Mezzanine card needs to be exchanged!

## Point 5
| Tunnel - position | [13 empty] | Surface position | [25 empty] | LEVEL | [8 empty] | Signal
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | Comment | ERRA | ERRB | ERRA | ERRB | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | HV | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Comment | Reset | Capa | 220V | Noise | HV | Source | other
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec. | 23.6.08 |  |  | 28.07.08 |  |  |  | 26-27.8.08 | [25 empty] | tested | tested |  | levels | tested | test | Comments
1 | BJBHT.A33L5 | BLECF_ | 335 | 0x | 014F | 6 | Arc | BYPLM.A33L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | reset problem
2 | BJBHT.A32L5 | BLECF_ | 330 | 0x | 014A | 6 | Arc | BYPLM.A32L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | LF | LF | Test in tunnel needed! | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
3 | BJBHT.A31L5 | BLECF_ | 650 | 0x | 028A | 6 | Arc | BYPLM.A31L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | OK | OK | Test in tunnel needed! | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
4 | BJBHT.A30L5 | BLECF_ | 414 | 0x | 019E | 6 | Arc | BYPLM.A30L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | ↑↑↑ | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
5 | BJBHT.A29L5 | BLECF_ | 307 | 0x | 0133 | 6 | Arc | BYPLM.A29L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
6 | BJBHT.A28L5 | BLECF_ | 447 | 0x | 01BF | 6 | Arc | BYPLM.A28L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | ↑↑ | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
7 | BJBHT.A27L5 | BLECF_ | 226 | 0x | 00E2 | 6 | Arc | BYPLM.A27L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
8 | BJBHT.A26L5 | BLECF_ | 680 | 0x | 02A8 | 6 | Arc | BYPLM.A26L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ↑↑ | OK | Test in tunnel needed! | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
9 | BJBHT.A25L5 | BLECF_ | 424 | 0x | 01A8 | 6 | Arc | BYPLM.A25L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
10 | BJBHT.A24L5 | BLECF_ | 419 | 0x | 01A3 | 6 | Arc | BYPLM.A24L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
11 | BJBHT.A23L5 | BLECF_ | 42 | 0x | 002A | 6 | Arc | BYPLM.A23L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | OK | OK | Test in tunnel needed! | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
12 | BJBHT.A22L5 | BLECF_ | 625 | 0x | 0271 | 6 | Arc | BYPLM.A22L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ↑5 | OK | Test in tunnel needed! | ↑ | OK |  |  | ↑ | OK | N | N | ↑ | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
13 | BJBHT.A21L5 | BLECF_ | 506 | 0x | 01FA | 6 | Arc | BYPLM.A21L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | OK | OK | Test in tunnel needed! | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
14 | BJBHT.A20L5 | BLECF_ | 589 | 0x | 024D | 6 | Arc | BYPLM.A20L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | OK | OK |  | ↑ | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
15 | BJBHT.A19L5 | BLECF_ | 546 | 0x | 0222 | 6 | Arc | BYPLM.A19L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
16 | BJBHT.A18L5 | BLECF_ | 473 | 0x | 01D9 | 6 | Arc | BYPLM.A18L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ↑↑ | OK | Test in tunnel needed! | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
17 | BJBHT.A17L5 | BLECF_ | 475 | 0x | 01DB | 6 | Arc | BYPLM.A17L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
18 | BJBHT.A16L5 | BLECF_ | 196 | 0x | 00C4 | 6 | Arc | BYPLM.A16L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | OK | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | test-cfc
19 | BJBHT.A15L5 | BLECF_ | 346 | 0x | 015A | 6 | Arc | BYPLM.A15L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
20 | BJBHT.A14L5 | BLECF_ | 48 | 0x | 0030 | 6 | Arc | BYPLM.A14L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X-2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
21 | BJBHT.A13L5 | BLECF_ | 640 | 0x | 0280 | 6 | Arc | BYPLM.A13L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | OK | OK | Test in tunnel needed! | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
22 | BJBHT.A12L5 | BLECF_ | 477 | 0x | 01DD | 6 | Arc | BYPLM.A12L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ↑↑↑ | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch 6 20 pA
23 | BJBHT.B11L5 | BLECF_ | 753 | 0x | 02F1 | 6 | RR53 | BY02 | 5 | CL | 2 | 9 | B | 13-14 | SR5 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | ↑↑↑ | OK |  | OK | ↑↑↑ |  |  | ↑↑↑ | OK | N | N | ↑↑ | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
24 | BJBHT.A11L5 | BLECF_ | 581 | 0x | 0245 | 7 | RR53 | BY02 | 5 | CL | 2 | 8 | B | 11-12 | SR5 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | test-cfc
25 | BJBHT.B10L5 | BLECF_ | 14 | 0x | 000E | 6 | RR53 | BY02 | 5 | CL | 2 | 7 | B | 9-10 | SR5 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 7 | 4 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch1 20 bit noise, ch 2 - 8 100 bit noise
26 | BJBHT.A10L5 | BLECF_ | 124 | 0x | 007C | 2 | RR53 | BY02 | 5 | CL | 2 | 6 | B | 7-8 | SR5 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch1 20 bit noise, ch 2 - 8 100 bit noise
27 | BJBHT.A9L5 | BLECF_ | 757 | 0x | 02F5 | 6 | RR53 | BY02 | 5 | CL | 2 | 5 | B | 5-6 | SR5 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | test-cfc
28 | BJBHT.B8L5 | BLECF_ | 190 | 0x | 00BE | 6 | RR53 | BY02 | 5 | CL | 2 | 4 | B | 3-4 | SR5 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch1 20 bit noise, ch 2 - 8 40 bit noise
29 | BJBHT.A8L5 | BLECF_ | 9 | 0x | 0009 | 6 | RR53 | BY02 | 5 | CL | 2 | 3 | B | 1-2 | SR5 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | ↑↑↑ | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 40 bit noise
30 | BJBHT.A7L5 | BLECF_ | 525 | 0x | 020D | 6 | RR53 | BY02 | 5 | CL | 2 | 2 | A | 23-24 | SR5 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 20 bit noise
31 | BJBHT.B6L5 | BLECF_ | 632 | 0x | 0278 | 7 | RR53 | BY02 | 5 | CL | 2 | 1 | A | 21-22 | SR5 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch1 no noise, ch 2 - 8 20 bit noise
32 | BJBHT.A6L5 | BLECF_ | 103 | 0x | 0067 | 1 | RR53 | BY02 | 5 | CL | 1 | 10 | A | 19-20 | SR5 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16 | LF | OK | Test in tunnel needed! | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 20 bit noise
33 | BJBHT.B5L5 | BLECF_ | 313 | 0x | 0139 | 6 | RR53 | BY02 | 5 | CL | 1 | 9 | A | 17-18 | SR5 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | [10 empty] | OK | OK | OK |  |  |  | no noise
34 | BJBHT.A5L5 | BLECF_ | 521 | 0x | 0209 | 2 | RR53 | BY02 | 5 | CL | 1 | 8 | A | 15-16 | SR5 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8 | OK | OK |  | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 1X-2X | ok | ok | ok | ok | ok | ok | ok | ok | [10 empty] | OK | OK | OK |  |  |  | test-cfc
35 | BJBHT.C4L5 | BLECF_ | 356 | 0x | 0164 | 6 | RR53 | BY02 | 5 | CL | 1 | 7 | A | 13-14 | SR5 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | ↑↑↑ | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X-2D | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | no noise
36 | BJBHT.B4L5 | BLECF_ | 370 | 0x | 0172 | 7 | RR53 | BY02 | 5 | CL | 1 | 6 | A | 11-12 | SR5 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X-35 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | no noise
37 | BJBHT.A4L5 | BLECF_ | 3 | 0x | 0003 | 2 | RR53 | BY02 | 5 | CL | 1 | 5 | A | 9-10 | SR5 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | ↑↑↑ | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK |  |  |  | test-cfc
38 | BJBHT.A3L5 | BLECF_ | 121 | 0x | 0079 | 5 | RR53 | BY02 | 5 | CL | 1 | 4 | A | 7-8 | SR5 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X-24 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch1 20 bit noise, ch 2 - 8 40 bit noise
39 | BJBHT.A2L5 | BLECF_ | 312 | 0x | 0138 | 8 | RR53 | BY02 | 5 | CL | 1 | 3 | A | 5-6 | SR5 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | OK | OK |  | OK | OK |  |  | ↑ | OK | N | N | ↑ | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch 2 peaks 150 - 200, rest 20 bit noise
40 | BJBHT.B1L5 | BLECF_ | 586 | 0x | 024A | 4 | RR53 | BY02 | 5 | CL | 1 | 2 | A | 3-4 | SR5 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch 1no noise , ch 2 -8 20 bit noise
41 | BJBHT.A1L5 | BLECF_ | 579 | 0x | 0243 | 2 | RR53 | BY02 | 5 | CL | 1 | 1 | A | 1-2 | SR5 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | ↑↑↑ | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0C-1X-3D | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | no noise
42 | BJBHT.A1R5 | BLECF_ | 193 | 0x | 00C1 | 2 | RR57 | BY02 | 5 | CR | 1 | 1 | A | 1-2 | SR5 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
43 | BJBHT.B1R5 | BLECF_ | 749 | 0x | 02ED | 6 | RR57 | BY02 | 5 | CR | 1 | 2 | A | 3-4 | SR5 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-2D | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
44 | BJBHT.A2R5 | BLECF_ | 671 | 0x | 029F | 6 | RR57 | BY02 | 5 | CR | 1 | 3 | A | 5-6 | SR5 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | ↑27E | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X-25 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
45 | BJBHT.A3R5 | BLECF_ | 750 | 0x | 02EE | 6 | RR57 | BY02 | 5 | CR | 1 | 4 | A | 7-8 | SR5 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | [10 empty] | OK | OK | OK
46 | BJBHT.A4R5 | BLECF_ | 726 | 0x | 02D6 | 2 | RR57 | BY02 | 5 | CR | 1 | 5 | A | 9-10 | SR5 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | OK | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
47 | BJBHT.B4R5 | BLECF_ | 663 | 0x | 0297 | 8 | RR57 | BY02 | 5 | CR | 1 | 6 | A | 11-12 | SR5 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
48 | BJBHT.C4R5 | BLECF_ | 710 | 0x | 02C6 | 6 | RR57 | BY02 | 5 | CR | 1 | 7 | A | 13-14 | SR5 | BY05 | C | 11 | 1-2 | BY02 | R | 3 | 4 | 1 | 1-8 | OK | OK |  | ↑↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
49 | BJBHT.A5R5 | BLECF_ | 598 | 0x | 0256 | 2 | RR57 | BY02 | 5 | CR | 1 | 8 | A | 15-16 | SR5 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
50 | BJBHT.B5R5 | BLECF_ | 697 | 0x | 02B9 | 6 | RR57 | BY02 | 5 | CR | 1 | 9 | A | 17-18 | SR5 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
51 | BJBHT.A6R5 | BLECF_ | 648 | 0x | 0288 | 2 | RR57 | BY02 | 5 | CR | 1 | 10 | A | 19-20 | SR5 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-19 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
52 | BJBHT.B6R5 | BLECF_ | 656 | 0x | 0290 | 6 | RR57 | BY02 | 5 | CR | 2 | 1 | A | 21-22 | SR5 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 5 | 2 | 1-8 | ↑↑↑ | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
53 | BJBHT.A7R5 | BLECF_ | 618 | 0x | 026A | 6 | RR57 | BY02 | 5 | CR | 2 | 2 | A | 23-24 | SR5 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
54 | BJBHT.A8R5 | BLECF_ | 596 | 0x | 0254 | 6 | RR57 | BY02 | 5 | CR | 2 | 3 | B | 1-2 | SR5 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | OK | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
55 | BJBHT.B8R5 | BLECF_ | 660 | 0x | 0294 | 6 | RR57 | BY02 | 5 | CR | 2 | 4 | B | 3-4 | SR5 | BY05 | C | 10 | 3-4 | BY02 | R | 3 | 6 | 3 | 1-8 | ↑↑↑ | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
56 | BJBHT.A9R5 | BLECF_ | 149 | 0x | 0095 | 6 | RR57 | BY02 | 5 | CR | 2 | 5 | B | 5-6 | SR5 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
57 | BJBHT.A10R5 | BLECF_ | 605 | 0x | 025D | 2 | RR57 | BY02 | 5 | CR | 2 | 6 | B | 7-8 | SR5 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
58 | BJBHT.B10R5 | BLECF_ | 673 | 0x | 02A1 | 6 | RR57 | BY02 | 5 | CR | 2 | 7 | B | 9-10 | SR5 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 7 | 4 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
59 | BJBHT.A11R5 | BLECF_ | 621 | 0x | 026D | 8 | RR57 | BY02 | 5 | CR | 2 | 8 | B | 11-12 | SR5 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
60 | BJBHT.B11R5 | BLECF_ | 603 | 0x | 025B | 6 | RR57 | BY02 | 5 | CR | 2 | 9 | B | 13-14 | SR5 | BY05 | C | 9 | 1-2 | BY02 | R | 3 | 7 | 4 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
61 | BJBHT.A12R5 | BLECF_ | 699 | 0x | 02BB | 6 | Arc | BYPLM.A12R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | LF | LF | Test in tunnel needed! | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-12 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
62 | BJBHT.A13R5 | BLECF_ | 728 | 0x | 02D8 | 6 | Arc | BYPLM.A13R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
63 | BJBHT.A14R5 | BLECF_ | 698 | 0x | 02BA | 6 | Arc | BYPLM.A14R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | ↑↑ | OK |  | OK | OK |  |  | ↑ | OK | N | N | ↑ | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
64 | BJBHT.A15R5 | BLECF_ | 737 | 0x | 02E1 | 6 | Arc | BYPLM.A15R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
65 | BJBHT.A16R5 | BLECF_ | 733 | 0x | 02DD | 6 | Arc | BYPLM.A16R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
66 | BJBHT.A17R5 | BLECF_ | 634 | 0x | 027A | 6 | Arc | BYPLM.A17R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
67 | BJBHT.A18R5 | BLECF_ | 732 | 0x | 02DC | 6 | Arc | BYPLM.A18R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | ↑↑↑ | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-11 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
68 | BJBHT.A19R5 | BLECF_ | 742 | 0x | 02E6 | 6 | Arc | BYPLM.A19R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-14 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
69 | BJBHT.A20R5 | BLECF_ | 695 | 0x | 02B7 | 6 | Arc | BYPLM.A20R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | LF | LF | Test in tunnel needed! | OK | LF |  |  | OK | OK | Y | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
70 | BJBHT.A21R5 | BLECF_ | 688 | 0x | 02B0 | 6 | Arc | BYPLM.A21R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
71 | BJBHT.A22R5 | BLECF_ | 739 | 0x | 02E3 | 6 | Arc | BYPLM.A22R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | OK | OK |  | OK | OK |  |  | ↑ | OK | N | N | ↑ | OK | OK | 0X-12 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
72 | BJBHT.A23R5 | BLECF_ | 691 | 0x | 02B3 | 6 | Arc | BYPLM.A23R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | LF | OK | Test in tunnel needed! | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-12 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
73 | BJBHT.A24R5 | BLECF_ | 617 | 0x | 0269 | 6 | Arc | BYPLM.A24R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | LF | OK |  | OK | OK |  |  | ↑ | OK | N | N | ↑ | OK | OK | 0X-11 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
74 | BJBHT.A25R5 | BLECF_ | 736 | 0x | 02E0 | 6 | Arc | BYPLM.A25R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-15 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
75 | BJBHT.A26R5 | BLECF_ | 683 | 0x | 02AB | 6 | Arc | BYPLM.A26R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
76 | BJBHT.A27R5 | BLECF_ | 490 | 0x | 01EA | 6 | Arc | BYPLM.A27R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-10 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
77 | BJBHT.A28R5 | BLECF_ | 734 | 0x | 02DE | 6 | Arc | BYPLM.A28R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | OK | LF |  | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
78 | BJBHT.A29R5 | BLECF_ | 701 | 0x | 02BD | 6 | Arc | BYPLM.A29R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | OK | ↑↑↑ |  | OK | ↑↑↑ |  |  | OK | ↑↑↑ | N | N | ↑↑↑ | ↑↑↑ | ↑↑↑ | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
79 | BJBHT.A30R5 | BLECF_ | 704 | 0x | 02C0 | 6 | Arc | BYPLM.A30R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | OK | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-13 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
80 | BJBHT.A31R5 | BLECF_ | 693 | 0x | 02B5 | 6 | Arc | BYPLM.A31R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
81 | BJBHT.A32R5 | BLECF_ | 735 | 0x | 02DF | 6 | Arc | BYPLM.A32R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | OK | OK |  | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 0X-14 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
82 | BJBHT.A33R5 | BLECF_ | 731 | 0x | 02DB | 6 | Arc | BYPLM.A33R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-12 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
83 | BJBHT.A34R5 | BLECF_ | 661 | 0x | 0295 | 6 | Arc | BYPLM.A34R5 | 5 | R |  |  |  | 7-9 | SR5 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | OK | OK |  | ↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-13 | ok | ok | ok | ok | ok | ok | ok | ok | [10 empty] | OK | OK | OK
last entry 22/5/08
| Door | [7 empty] | Transformer soldered | [16 empty] | ↑↑↑ | CRC Error consantly fast increasing
[26 empty] | ↑↑ | CRC Error sometimes fast increasing
[26 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number correct | [19 empty] | LF | Lost Frame
|  |  |  | 0x | 0000 | Card number needs to be checked | [19 empty] | no TC | No Threshold Comparator installed
|  |  |  | 0x | 0000 | Card number unknown need to be checked in the tunnel | [20 empty] | Opitcal link is working with Test system
|  | BLECF_ | 656 | Card was exchanged, Data base need to be checked for the correct number
Note on D58: effinger: 565 original number chnaged to 656
Note on D100: effinger: 565 original number chnaged to 656
Note on AA6: effinger: not working with the TB, 0W
Note on AB6: effinger: not working with the TB, 0W
Note on AA8: effinger: 219uW
Note on AB8: effinger: 28uW
Note on AA10: effinger: 154uW
Note on AB10: effinger: 228uW
Note on AA12: effinger: 89uW, TB not o.k
Note on AB12: effinger: 0uW, TB not o.k
Note on AA16: effinger: 0W, TB not o.k.
Note on AB16: effinger: 0W, TB not o.k.
Note on AA20: effinger: 0W, TB not o.k.
Note on AB20: effinger: 0W, TB not o.k.
Note on AA26: effinger: 249uW
Note on AB26: effinger: 259uW
Note on AA28: effinger: 306uW
Note on AB28: effinger: 310uW
Note on AA34: effinger: 278uW
Note on AB34: effinger: 258uW
Note on AA37: effinger: 0W, TB not o.k.
Note on AB37: effinger: 243uW
Note on AA40: effinger: 212uW
Note on AB40: effinger: 256uW
Note on AA42: effinger: 260uW
Note on AB42: effinger: 214uW
Note on AA46: effinger: 338uW
Note on AB46: effinger: 237uW
Note on AA50: effinger: 226uW
Note on AB50: effinger: 231uW
Note on AA58: effinger: 295uW
Note on AB58: effinger: 280uW
Note on AA61: effinger: 319uW
Note on AB61: effinger: 245uW
Note on AA68: effinger: TB not o.k 0W
Note on AB68: effinger: TB not o.k 0W
Note on AA70: effinger: 305uW
Note on AB70: effinger: 267uW
Note on AA74: effinger: 286uW
Note on AB74: effinger: 321uW
Note on AB76: effinger: 273uW
Note on AA76: effinger: TB not o.k. 0W
Note on AA79: effinger: TB not o.k 229nW
Note on AB79: effinger: 300uW
Note on AA80: effinger: 234uW
Note on AB80: effinger: 246uW
Note on AA84: effinger: 94uW
Note on AB84: effinger: 304uW
Note on AA85: effinger: 240uW
Note on AB85: effinger: 258uW
Note on AC68: effinger: both jumper not correctly pluged, Card o.k.,
Note on AC76: effinger: jumper not correctly pluged, Card o.k.,
Note on AC79: effinger: jumper not correctly pluged, Card o.k.,
Note on AD37: effinger: Fibre broken from tunnel to surface --> Spare fibre used; B15 - C4 3

## Point 6
| Tunnel - position | [13 empty] | Surface position | [25 empty] | LEVEL | [8 empty] | Signal
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | Comment | ERRA | ERRB | ERRA | ERRB | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | HV | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Comment | Reset | Capa | 220V | Noise | HV | Source | other
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec. | [7 empty] | 26-27.8.08 | [25 empty] | tested | tested |  | levels | tested | test | Comments
1 | BJBHT.A33L6 | BLECF_ | 533 | 0x | 0215 | 6 | Arc | BYPLM.A33L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
2 | BJBHT.A32L6 | BLECF_ | 117 | 0x | 0075 | 6 | Arc | BYPLM.A32L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | OK | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
3 | BJBHT.A31L6 | BLECF_ | 232 | 0x | 00E8 | 6 | Arc | BYPLM.A31L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
4 | BJBHT.A30L6 | BLECF_ | 93 | 0x | 005D | 6 | Arc | BYPLM.A30L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
5 | BJBHT.A29L6 | BLECF_ | 163 | 0x | 00A3 | 6 | Arc | BYPLM.A29L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
6 | BJBHT.A28L6 | BLECF_ | 549 | 0x | 0225 | 6 | Arc | BYPLM.A28L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ↑↑↑ | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
7 | BJBHT.A27L6 | BLECF_ | 166 | 0x | 00A6 | 6 | Arc | BYPLM.A27L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
8 | BJBHT.A26L6 | BLECF_ | 266 | 0x | 010A | 6 | Arc | BYPLM.A26L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ↑↑↑ | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
9 | BJBHT.A25L6 | BLECF_ | 162 | 0x | 00A2 | 6 | Arc | BYPLM.A25L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
10 | BJBHT.A24L6 | BLECF_ | 168 | 0x | 00A8 | 6 | Arc | BYPLM.A24L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
11 | BJBHT.A23L6 | BLECF_ | 177 | 0x | 00B1 | 6 | Arc | BYPLM.A23L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
12 | BJBHT.A22L6 | BLECF_ | 197 | 0x | 00C5 | 6 | Arc | BYPLM.A22L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
13 | BJBHT.A21L6 | BLECF_ | 170 | 0x | 00AA | 6 | Arc | BYPLM.A21L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
14 | BJBHT.A20L6 | BLECF_ | 179 | 0x | 00B3 | 6 | Arc | BYPLM.A20L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | LF | OK | Test in tunnel needed! |  |  |  |  | ↑ | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
15 | BJBHT.A19L6 | BLECF_ | 178 | 0x | 00B2 | 6 | Arc | BYPLM.A19L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
16 | BJBHT.A18L6 | BLECF_ | 207 | 0x | 00CF | 6 | Arc | BYPLM.A18L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
17 | BJBHT.A17L6 | BLECF_ | 125 | 0x | 007D | 6 | Arc | BYPLM.A17L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
18 | BJBHT.A16L6 | BLECF_ | 708 | 0x | 02C4 | 6 | Arc | BYPLM.A16L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | OK | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
19 | BJBHT.A15L6 | BLECF_ | 299 | 0x | 012B | 6 | Arc | BYPLM.A15L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
20 | BJBHT.A14L6 | BLECF_ | 281 | 0x | 0119 | 6 | Arc | BYPLM.A14L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ↑↑↑ | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
21 | BJBHT.A13L6 | BLECF_ | 268 | 0x | 010C | 6 | Arc | BYPLM.A13L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 2A-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
22 | BJBHT.A12L6 | BLECF_ | 172 | 0x | 00AC | 6 | Arc | BYPLM.A12L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | ↑↑↑ | OK | Power loss could be checked |  |  |  |  | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
23 | UD62 | BLECF_ | 187 | 0x | 00BB | 8 | UJ63 | BY03 | 6 | CL | 3 | 1 | B | 17-18 | SR6 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 11 | 8 | 1-8 | ↑↑ | OK | Power loss could be checked |  |  |  |  | OK | OK | N | N | OK | OK | OK | 5X-9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
24 | TD62.2 | BLECF_ | 255 | 0x | 00FF | 8 | UJ63 | BY03 | 6 | CL | 2 | 10 | B | 15-16 | SR6 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X-8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
25 | TD62.1 | BLECF_ | 511 | 0x | 01FF | 8 | UJ63 | BY03 | 6 | CL | 2 | 9 | B | 13-14 | SR6 | BY05 | C | 4 | 1-2 | BY02 | C | 2 | 10 | 7 | 1-8 | ↑↑↑ | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X-FF | ok | ok | ok | ok | ok | bl |  |  | ok | [9 empty] | OK | OK | OK
26 | BJBHT.A11L6 | BLECF_ | 108 | 0x | 006C | 6 | UJ63 | BY03 | 6 | CL | 2 | 8 | B | 11-12 | SR6 | BY05 | C | 3 | 11-12 | BY02 | L | 1 | 6 | 3 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
27 | BJBHT.A10L6 | BLECF_ | 91 | 0x | 005B | 6 | UJ63 | BY03 | 6 | CL | 2 | 7 | B | 9-10 | SR6 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 1-8 | ↑4 | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
28 | BJBHT.A9L6 | BLECF_ | 161 | 0x | 00A1 | 6 | UJ63 | BY03 | 6 | CL | 2 | 6 | B | 7-8 | SR6 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 5 | 2 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X-83 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
29 | BJBHT.A8L6 | BLECF_ | 115 | 0x | 0073 | 6 | UJ63 | BY03 | 6 | CL | 2 | 5 | B | 5-6 | SR6 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 5 | 2 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
30 | BJBHT.C5L6 | BLECF_ | 49 | 0x | 0031 | 1 | UJ63 | BY03 | 6 | CL | 2 | 4 | B | 3-4 | SR6 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 9 | 6 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X ch1 74 ch3 26 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
31 | BJBHT.B5L6 | BLECF_ | 429 | 0x | 01AD | 8 | UJ63 | BY03 | 6 | CL | 2 | 3 | B | 1-2 | SR6 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 4 | 1 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X-8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKD
32 | BJBHT.A5L6 | BLECF_ | 107 | 0x | 006B | 3 | UJ63 | BY03 | 6 | CL | 2 | 2 | A | 23-24 | SR6 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 1X-4X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK |  |  |  | MKD, TEST_CFC
33 | BJBHT.K4L6 | BLECF_ | 223 | 0x | 00DF | 8 | UJ63 | BY03 | 6 | CL | 2 | 1 | A | 21-22 | SR6 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 4 | 1 | 1-8 | OK | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 7X-8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKD
34 | BJBHT.J4L6 | BLECF_ | 548 | 0x | 0224 | 5 | UJ63 | BY03 | 6 | CL | 1 | 10 | A | 19-20 | SR6 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
35 | BJBHT.I4L6 | BLECF_ | 508 | 0x | 01FC | 5 | UJ63 | BY03 | 6 | CL | 1 | 9 | A | 17-18 | SR6 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 8 | 5 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X-7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
36 | BJBHT.H4L6 | BLECF_ | 336 | 0x | 0150 | 3 | UJ63 | BY03 | 6 | CL | 1 | 8 | A | 15-16 | SR6 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 1X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKB
37 | BJBHT.G4L6 | BLECF_ | 66 | 0x | 0042 | 3 | UJ63 | BY03 | 6 | CL | 1 | 7 | A | 13-14 | SR6 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X-BC | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKB
38 | BJBHT.F4L6 | BLECF_ | 610 | 0x | 0262 | 2 | UJ63 | BY03 | 6 | CL | 1 | 6 | A | 11-12 | SR6 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X-37-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
39 | BJBHT.E4L6 | BLECF_ | 392 | 0x | 0188 | 2 | UJ63 | BY03 | 6 | CL | 1 | 5 | A | 9-10 | SR6 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | OK | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 0X-2X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
40 | BJBHT.D4L6 | BLECF_ | 612 | 0x | 0264 | 8 | UJ63 | BY03 | 6 | CL | 1 | 4 | A | 7-8 | SR6 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MSD
41 | BJBHT.C4L6 | BLECF_ | 12 | 0x | 000C | 8 | UJ63 | BY03 | 6 | CL | 1 | 3 | A | 5-6 | SR6 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | OK | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 7X-8X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK |  |  |  | MSD
42 | BJBHT.B4L6 | BLECF_ | 436 | 0x | 01B4 | 7 | UJ63 | BY03 | 6 | CL | 1 | 2 | A | 3-4 | SR6 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 4X-71 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MSD
43 | BJBHT.A4L6 | BLECF_ | 144 | 0x | 0090 | 7 | UJ63 | BY03 | 6 | CL | 1 | 1 | A | 1-2 | SR6 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X-9D | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MSD
44 | BJBHT.A4R6 | BLECF_ | 102 | 0x | 0066 | 7 | UJ67 | BY03 | 6 | CR | 1 | 1 | A | 1-2 | SR6 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | OK | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 7X-95 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MSD
45 | BJBHT.B4R6 | BLECF_ | 116 | 0x | 0074 | 7 | UJ67 | BY03 | 6 | CR | 1 | 2 | A | 3-4 | SR6 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | LF | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 4X-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  | high noise |  | MSD
46 | BJBHT.C4R6 | BLECF_ | 156 | 0x | 009C | 8 | UJ67 | BY03 | 6 | CR | 1 | 3 | A | 5-6 | SR6 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | OK | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 7X-8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  | high noise |  | MSD
47 | BJBHT.D4R6 | BLECF_ | 202 | 0x | 00CA | 8 | UJ67 | BY03 | 6 | CR | 1 | 4 | A | 7-8 | SR6 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 4X-51 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MSD
48 | BJBHT.E4R6 | BLECF_ | 167 | 0x | 00A7 | 2 | UJ67 | BY03 | 6 | CR | 1 | 5 | A | 9-10 | SR6 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | ↑↑↑ | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 1X-7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
49 | BJBHT.F4R6 | BLECF_ | 545 | 0x | 0221 | 2 | UJ67 | BY03 | 6 | CR | 1 | 6 | A | 11-12 | SR6 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok |  | ok | ok | RSTDAC | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
50 | BJBHT.G4R6 | BLECF_ | 219 | 0x | 00DB | 3 | UJ67 | BY03 | 6 | CR | 1 | 7 | A | 13-14 | SR6 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 36-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKB
51 | BJBHT.H4R6 | BLECF_ | 45 | 0x | 002D | 3 | UJ67 | BY03 | 6 | CR | 1 | 8 | A | 15-16 | SR6 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKB
52 | BJBHT.I4R6 | BLECF_ | 173 | 0x | 00AD | 5 | UJ67 | BY03 | 6 | CR | 1 | 9 | A | 17-18 | SR6 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 17 | 13 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
53 | BJBHT.J4R6 | BLECF_ | 262 | 0x | 0106 | 4 | UJ67 | BY03 | 6 | CR | 1 | 10 | A | 19-20 | SR6 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 4X-57 | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK |  |  |  | MKD, TEST_CFC
54 | BJBHT.K4R6 | BLECF_ | 198 | 0x | 00C6 | 8 | UJ67 | BY03 | 6 | CR | 2 | 1 | A | 21-22 | SR6 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 4 | 1 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok |  | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKD, TEST_CFC
55 | BJBHT.A5R6 | BLECF_ | 289 | 0x | 0121 | 3 | UJ67 | BY03 | 6 | CR | 2 | 2 | A | 23-24 | SR6 | BY05 | C | 11 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 1X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKD
56 | BJBHT.B5R6 | BLECF_ | 165 | 0x | 00A5 | 8 | UJ67 | BY03 | 6 | CR | 2 | 3 | B | 1-2 | SR6 | BY05 | C | 10 | 1-2 | BY02 | R | 3 | 4 | 1 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
57 | BJBHT.C5R6 | BLECF_ | 88 | 0x | 0058 | 1 | UJ67 | BY03 | 6 | CR | 2 | 4 | B | 3-4 | SR6 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 18 | 14 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X-1X (ch1 75) | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
58 | BJBHT.A8R6 | BLECF_ | 159 | 0x | 009F | 6 | UJ67 | BY03 | 6 | CR | 2 | 5 | B | 5-6 | SR6 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 5 | 2 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X-9A | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
59 | BJBHT.A9R6 | BLECF_ | 542 | 0x | 021E | 6 | UJ67 | BY03 | 6 | CR | 2 | 6 | B | 7-8 | SR6 | BY05 | C | 10 | 7-8 | BY02 | R | 3 | 5 | 2 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
60 | BJBHT.A10R6 | BLECF_ | 164 | 0x | 00A4 | 6 | UJ67 | BY03 | 6 | CR | 2 | 7 | B | 9-10 | SR6 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 6 | 3 | 1-8 | OK | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 6X-7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
61 | BJBHT.A11R6 | BLECF_ | 72 | 0x | 0048 | 6 | UJ67 | BY03 | 6 | CR | 2 | 8 | B | 11-12 | SR6 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 6 | 3 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
62 | TD68.1 | BLECF_ | 391 | 0x | 0187 | 8 | UJ67 | BY03 | 6 | CR | 2 | 9 | B | 13-14 | SR6 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 19 | 15 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 1A-7X-85 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
63 | TD68.2 | BLECF_ | 56 | 0x | 0038 | 8 | UJ67 | BY03 | 6 | CR | 2 | 10 | B | 15-16 | SR6 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 7X-FF | ok | ok | ok | ok | ok | ok | ok | [11 empty] | OK | OK | OK
64 | UD68 | BLECF_ | 82 | 0x | 0052 | 8 | UJ67 | BY03 | 6 | CR | 3 | 1 | B | 17-18 | SR6 | BY05 | C | 9 | 5-6 | BY02 | C | 2 | 20 | 16 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X-7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
65 | BJBHT.A12R6 | BLECF_ | 527 | 0x | 020F | 6 | Arc | BYPLM.A12R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 7 | 4 | 1-8 | OK | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  | bit noise |  | Fiber exchanged
66 | BJBHT.A13R6 | BLECF_ | 230 | 0x | 00E6 | 6 | Arc | BYPLM.A13R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 7 | 4 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 00-15 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [8 empty] | Tuersday 2 sept. replug of the card/ modulation ok | OK | OK | OK |  | bit noise |  | Fiber exchanged
67 | BJBHT.A14R6 | BLECF_ | 532 | 0x | 0214 | 6 | Arc | BYPLM.A14R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | ↑186 | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  | bit noise
68 | BJBHT.A15R6 | BLECF_ | 569 | 0x | 0239 | 6 | Arc | BYPLM.A15R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
69 | BJBHT.A16R6 | BLECF_ | 274 | 0x | 0112 | 6 | Arc | BYPLM.A16R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | ↑↑↑ | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
70 | BJBHT.A17R6 | BLECF_ | 301 | 0x | 012D | 6 | Arc | BYPLM.A17R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
71 | BJBHT.A18R6 | BLECF_ | 19 | 0x | 0013 | 6 | Arc | BYPLM.A18R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | OK | OK | [5 empty] | ↑ | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
72 | BJBHT.A19R6 | BLECF_ | 563 | 0x | 0233 | 6 | Arc | BYPLM.A19R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
73 | BJBHT.A20R6 | BLECF_ | 566 | 0x | 0236 | 6 | Arc | BYPLM.A20R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | ↑↑ | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
74 | BJBHT.A21R6 | BLECF_ | 295 | 0x | 0127 | 6 | Arc | BYPLM.A21R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
75 | BJBHT.A22R6 | BLECF_ | 28 | 0x | 001C | 6 | Arc | BYPLM.A22R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | ↑↑↑ | OK | Test in tunnel needed! |  |  |  |  | ↑ | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
76 | BJBHT.A23R6 | BLECF_ | 540 | 0x | 021C | 6 | Arc | BYPLM.A23R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
77 | BJBHT.A24R6 | BLECF_ | 364 | 0x | 016C | 6 | Arc | BYPLM.A24R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | ↑↑↑ | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
78 | BJBHT.A25R6 | BLECF_ | 555 | 0x | 022B | 6 | Arc | BYPLM.A25R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
79 | BJBHT.A26R6 | BLECF_ | 564 | 0x | 0234 | 6 | Arc | BYPLM.A26R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | OK | OK | Test in tunnel needed! |  |  |  |  | OK | OK | N | N | OK | OK | OK | 5X | bl | bl | bl | bl | bl | bl | bl | bl | ok | [9 empty] | OK | OK | OK |  |  |  | TEST_CFC
80 | BJBHT.A27R6 | BLECF_ | 557 | 0x | 022D | 6 | Arc | BYPLM.A27R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK
81 | BJBHT.A28R6 | BLECF_ | 63 | 0x | 003F | 6 | Arc | BYPLM.A28R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
82 | BJBHT.A29R6 | BLECF_ | 561 | 0x | 0231 | 6 | Arc | BYPLM.A29R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
83 | BJBHT.A30R6 | BLECF_ | 565 | 0x | 0235 | 6 | Arc | BYPLM.A30R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | ↑↑↑ | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
84 | BJBHT.A31R6 | BLECF_ | 210 | 0x | 00D2 | 6 | Arc | BYPLM.A31R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | Earth problem
85 | BJBHT.A32R6 | BLECF_ | 180 | 0x | 00B4 | 6 | Arc | BYPLM.A32R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | all channels overshoot
86 | BJBHT.A33R6 | BLECF_ | 538 | 0x | 021A | 6 | Arc | BYPLM.A33R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | OK | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK |  |  |  | ch 5 200pA from the cabling, TEST_CFC
87 | BJBHT.A34R6 | BLECF_ | 176 | 0x | 00B0 | 6 | Arc | BYPLM.A34R6 | 6 | R |  |  |  | 7-9 | SR6 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | ↑4 | OK | [5 empty] | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
last entry 23/6/08
| Door | [7 empty] | Transformer soldered | [16 empty] | ↑↑↑ | CRC Error consantly fast increasing
[26 empty] | ↑↑ | CRC Error sometimes fast increasing
[26 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number correct | [19 empty] | LF | Lost Frame
|  |  |  | 0x | 0000 | Card number needs to be checked | [19 empty] | no TC | No Threshold Comparator installed
|  |  |  | 0x | 0000 | Card number unknown need to be checked in the tunnel | [20 empty] | Opitcal link is working with Test system
|  | BLECF_ | 656 | Card was exchanged, Data base need to be checked for the correct number
Note on D22: effinger: BLECF_0279 repalced by BLECF_0708 at the 5.6.08 by Ion
Note on D104: effinger: 565 original number chnaged to 656
Note on AA10: effinger: 249.6uW
Note on AB10: effinger: 281.4uW
Note on AA12: effinger: 213.6uW
Note on AB12: effinger: 255.8uW
Note on AA18: effinger: 3.5uW on the patch, not o.k with TB
Note on AB18: effinger: 310uW
Note on AA24: effinger: 225uW
Note on AB24: effinger: 278uW
Note on AA26: effinger: 246uW
Note on AB26: effinger: 16.5uW
Note on AA32: effinger: 274uW
Note on AB32: effinger: 206uW
Note on AA28: effinger: 65uW
Note on AB28: effinger: 288uW
Note on AA30: effinger: 260uW
Note on AB30: effinger: 232uW
Note on AA51: effinger: 251uW
Note on AB51: effinger: 249uW
Note on AA54: effinger: 194uW
Note on AB54: effinger: 258uW
Note on AA74: effinger: 304uW
Note on AB74: effinger: 288uW
Note on AA76: effinger: 275uW
Note on AB76: effinger: 302uW
Note on AA80: effinger: 251uW
Note on AB80: effinger: 313uW
Note on AA82: effinger: TB not o.k. 0W
Note on AB82: effinger: TB not o.k 0W
Note on AA84: effinger: 242uW
Note on AB84: effinger: 269uW
Note on AA90: effinger: 260uW
Note on AB90: effinger: 139uW
Note on AA94: effinger: 189uW
Note on AB94: effinger: 234uW
Note on AC82: effinger: CFC Card not working correctly-> reset -> o.k
Note on AC86: effinger: CFC Card not working correctly-> reset -> o.k
Note on AC18: effinger: A20L6; Jumper was not conneted properly; Card working corectly: reception o.k.

## Point 7
| Tunnel - position | [13 empty] | Surface position | [25 empty] | LEVEL | [8 empty] | Signal
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | Comment | ERRA | ERRB | ERRA | ERRB | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Comment | Reset | Capa | 220V | Noise | HV | Source | other
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec. | 23.6.08 |  |  | 15.7.08 |  | 24.7.08 |  | 26-27.8.08 | [25 empty] | tested | tested |  | levels | tested | test | Comments
1 | BJBHT.A33L7 | BLECF_ | 69 | 0x | 0045 | 6 | Arc | BYPLM.A33L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
2 | BJBHT.A32L7 | BLECF_ | 122 | 0x | 007A | 6 | Arc | BYPLM.A32L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 220 V supply missing,?!
3 | BJBHT.A31L7 | BLECF_ | 426 | 0x | 01AA | 6 | Arc | BYPLM.A31L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
4 | BJBHT.A30L7 | BLECF_ | 580 | 0x | 0244 | 6 | Arc | BYPLM.A30L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | LF | LF | to be check in tunnel with TB | LF | LF | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | DACRST | [9 empty] | OK | OK | OK
5 | BJBHT.A29L7 | BLECF_ | 123 | 0x | 007B | 6 | Arc | BYPLM.A29L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | ↑↑↑ | OK |  | ↑↑↑ | OK | ↑↑↑ | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
6 | BJBHT.A28L7 | BLECF_ | 20 | 0x | 0014 | 6 | Arc | BYPLM.A28L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | ↑ | ↑↑↑ |  | ↑ | ↑↑↑ | OK | ↑↑↑ | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
7 | BJBHT.A27L7 | BLECF_ | 140 | 0x | 008C | 6 | Arc | BYPLM.A27L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | BLM missing
8 | BJBHT.A26L7 | BLECF_ | 126 | 0x | 007E | 6 | Arc | BYPLM.A26L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | OK | OK |  | ↑ | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
9 | BJBHT.A25L7 | BLECF_ | 43 | 0x | 002B | 6 | Arc | BYPLM.A25L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
10 | BJBHT.A24L7 | BLECF_ | 454 | 0x | 01C6 | 6 | Arc | BYPLM.A24L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
11 | BJBHT.A23L7 | BLECF_ | 17 | 0x | 0011 | 6 | Arc | BYPLM.A23L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
12 | BJBHT.A22L7 | BLECF_ | 47 | 0x | 002F | 6 | Arc | BYPLM.A22L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ↑ | OK |  | ↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
13 | BJBHT.A21L7 | BLECF_ | 87 | 0x | 0057 | 6 | Arc | BYPLM.A21L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
14 | BJBHT.A20L7 | BLECF_ | 74 | 0x | 004A | 6 | Arc | BYPLM.A20L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | OK | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 6X-7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | BLM missing
15 | BJBHT.A19L7 | BLECF_ | 466 | 0x | 01D2 | 6 | Arc | BYPLM.A19L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X-8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
16 | BJBHT.A18L7 | BLECF_ | 571 | 0x | 023B | 6 | Arc | BYPLM.A18L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | LF | LF | to be check in tunnel with TB | LF | LF | ↑ | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
17 | BJBHT.A17L7 | BLECF_ | 430 | 0x | 01AE | 6 | Arc | BYPLM.A17L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
18 | BJBHT.A16L7 | BLECF_ | 39 | 0x | 0027 | 6 | Arc | BYPLM.A16L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
19 | BJBHT.A15L7 | BLECF_ | 534 | 0x | 0216 | 6 | Arc | BYPLM.A15L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X-8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
20 | BJBHT.A14L7 | BLECF_ | 461 | 0x | 01CD | 6 | Arc | BYPLM.A14L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
21 | BJBHT.A13L7 | BLECF_ | 449 | 0x | 01C1 | 6 | Arc | BYPLM.A13L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
22 | BJBHT.A12L7 | BLECF_ | 68 | 0x | 0044 | 6 | Arc | BYPLM.A12L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | BLM missing
23 | BJBHT.B11L7 | BLECF_ | 469 | 0x | 01D5 | 8 | RR73 | BY01 | 7 | CL | 3 | 8 | F | 1-2 | SR7 | BY05 | C | 5 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | LF | LF | was wrong conneted on patch | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK |  |  |  | ch1 no noise ch 2 - 8 60 bit noise, fiber missing
24 | BJBHT.A11L7 | BLECF_ | 679 | 0x | 02A7 | 8 | RR73 | BY01 | 7 | CL | 3 | 7 | E | 5-6 | SR7 | BY05 | C | 5 | 5-6 | BY02 | C | 4 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 9X-FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | ch1 no noise ch2 - 8 100 bit noise
25 | BJBHT.A10L7 | BLECF_ | 544 | 0x | 0220 | 8 | RR73 | BY01 | 7 | CL | 3 | 6 | E | 3-4 | SR7 | BY05 | C | 5 | 3-4 | BY02 | L | 1 | 7 | 4 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK |  |  |  | ch 1 -2 no noise ch3 - 8 60 bit noise
26 | BJBHT.B9L7 | BLECF_ | 678 | 0x | 02A6 | 8 | RR73 | BY01 | 7 | CL | 3 | 5 | E | 1-2 | SR7 | BY05 | C | 5 | 1-2 | BY02 | L | 1 | 6 | 3 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X-9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 20 bit noise
27 | BJBHT.A9L7 | BLECF_ | 559 | 0x | 022F | 8 | RR73 | BY01 | 7 | CL | 3 | 4 | D | 23-24 | SR7 | BY05 | C | 4 | 11-12 | BY02 | E | 4 | 5 | 2 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 6X-7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | test-cfc
28 | BJBHT.B8L7 | BLECF_ | 649 | 0x | 0289 | 6 | RR73 | BY01 | 7 | CL | 3 | 3 | D | 21-22 | SR7 | BY05 | C | 4 | 9-10 | BY02 | L | 1 | 6 | 3 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 6X-7X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK |  |  |  | test-cfc
29 | BJBHT.A8L7 | BLECF_ | 462 | 0x | 01CE | 3 | RR73 | BY01 | 7 | CL | 3 | 2 | D | 19-20 | SR7 | BY05 | C | 4 | 7-8 | BY02 | E | 4 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-4X | ok | ok | ok | ok | ok | ok | ok | ok | DACRST | [9 empty] | OK | OK | OK |  |  |  | test-cfc
30 | BJBHT.B7L7 | BLECF_ | 513 | 0x | 0201 | 7 | RR73 | BY01 | 7 | CL | 3 | 1 | D | 17-18 | SR7 | BY05 | C | 4 | 5-6 | BY02 | L | 1 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | RSTGOH | [9 empty] | OK | OK | OK |  |  |  | ch 6 -7 200 bit peaks; test-cfc
31 | BJBHT.A7L7 | BLECF_ | 411 | 0x | 019B | 2 | RR73 | BY01 | 7 | CL | 2 | 10 | D | 15-16 | SR7 | BY05 | C | 4 | 3-4 | BY02 | E | 4 | 4 | 1 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 00 |  | ok | ok | ok | ok | ok | ok | ok | DACRST | [9 empty] | OK | OK | OK |  |  |  | test-cfc
32 | BJBHT.I6L7 | BLECF_ | 587 | 0x | 024B | 7 | RR73 | BY01 | 7 | CL | 2 | 9 | D | 13-14 | SR7 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 5 | 2 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 6X-7X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | test-cfc
33 | BJBHT.H6L7 | BLECF_ | 118 | 0x | 0076 | 4 | RR73 | BY01 | 7 | CL | 2 | 8 | D | 11-12 | SR7 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 11 | 8 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-3X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
34 | BJBHT.G6L7 | BLECF_ | 191 | 0x | 00BF | 4 | RR73 | BY01 | 7 | CL | 2 | 7 | D | 9-10 | SR7 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 11 | 8 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK |  |  |  | test-cfc
35 | BJBHT.F6L7 | BLECF_ | 434 | 0x | 01B2 | 7 | RR73 | BY01 | 7 | CL | 2 | 6 | D | 7-8 | SR7 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 10 | 7 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
36 | BJBHT.E6L7 | BLECF_ | 560 | 0x | 0230 | 7 | RR73 | BY01 | 7 | CL | 2 | 5 | D | 5-6 | SR7 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 10 | 7 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK |  |  |  | test-cfc
37 | BJBHT.D6L7 | BLECF_ | 355 | 0x | 0163 | 3 | RR73 | BY01 | 7 | CL | 2 | 4 | D | 3-4 | SR7 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 9 | 6 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | no noise
38 | BJBHT.C6L7 | BLECF_ | 553 | 0x | 0229 | 3 | RR73 | BY01 | 7 | CL | 2 | 3 | D | 1-2 | SR7 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 9 | 6 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
39 | BJBHT.B6L7 | BLECF_ | 682 | 0x | 02AA | 7 | RR73 | BY01 | 7 | CL | 2 | 2 | C | 23-24 | SR7 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 8 | 5 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 10 bit noise
40 | BJBHT.A6L7 | BLECF_ | 465 | 0x | 01D1 | 7 | RR73 | BY01 | 7 | CL | 2 | 1 | C | 21-22 | SR7 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X-9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 10 bit noise
41 | BJBHT.E5L7 | BLECF_ | 96 | 0x | 0060 | 6 | RR73 | BY01 | 7 | CL | 1 | 10 | C | 19-20 | SR7 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X-9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 20 bit noise
42 | BJBHT.D5L7 | BLECF_ | 94 | 0x | 005E | 6 | RR73 | BY01 | 7 | CL | 1 | 9 | C | 17-18 | SR7 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 7 | 4 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK |  |  |  | test-cfc
43 | BJBHT.C5L7 | BLECF_ | 273 | 0x | 0111 | 6 | RR73 | BY01 | 7 | CL | 1 | 8 | C | 15-16 | SR7 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8 | OK | OK |  | ↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X-9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 50 bit noise
44 | BJBHT.B5L7 | BLECF_ | 590 | 0x | 024E | 8 | RR73 | BY01 | 7 | CL | 1 | 7 | C | 13-14 | SR7 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 6 | 3 | 9-16 | OK | OK |  | OK | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | test-cfc
45 | BJBHT.A5L7 | BLECF_ | 1 | 0x | 0001 | 8 | RR73 | BY01 | 7 | CL | 1 | 6 | C | 11-12 | SR7 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK |  |  |  | 20 bit noise
46 | BJBHT.E4L6 | BLECF_ | 372 | 0x | 0174 | 6 | RR73 | BY01 | 7 | CL | 1 | 5 | C | 9-10 | SR7 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 4 | 1 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X-9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 4096 bit noise ch 5 - 6
47 | BJBHT.D4L6 | BLECF_ | 445 | 0x | 01BD | 6 | RR73 | BY01 | 7 | CL | 1 | 4 | C | 7-8 | SR7 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 20 bit noise
48 | BJBHT.C4L6 | BLECF_ | 369 | 0x | 0171 | 6 | RR73 | BY01 | 7 | CL | 1 | 3 | C | 5-6 | SR7 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-5X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK |  |  |  | test-cfc
49 | BJBHT.B4L6 | BLECF_ | 633 | 0x | 0279 | 6 | RR73 | BY01 | 7 | CL | 1 | 2 | C | 3-4 | SR7 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 3X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 10 bit noise
50 | BJBHT.A4L6 | BLECF_ | 308 | 0x | 0134 | 6 | RR73 | BY01 | 7 | CL | 1 | 1 | C | 1-2 | SR7 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | ↑↑ | OK |  | ↑↑ | OK | ↑ | OK | OK | OK | N | N | OK | OK | OK | 8X-9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | 100 bit noise
51 | BJBHT.A4R7 | BLECF_ | 443 | 0x | 01BB | 6 | RR77 | BY02 | 7 | CR | 1 | 1 | C | 1-2 | SR7 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
52 | BJBHT.B4R7 | BLECF_ | 455 | 0x | 01C7 | 6 | RR77 | BY02 | 7 | CR | 1 | 2 | C | 3-4 | SR7 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
53 | BJBHT.C4R7 | BLECF_ | 351 | 0x | 015F | 6 | RR77 | BY02 | 7 | CR | 1 | 3 | C | 5-6 | SR7 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-13 | ok | ok | ok | ok | ok | ok | ok | ok | IN DACRST | [9 empty] | OK | OK | OK
54 | BJBHT.D4R7 | BLECF_ | 431 | 0x | 01AF | 6 | RR77 | BY02 | 7 | CR | 1 | 4 | C | 7-8 | SR7 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
55 | BJBHT.E4R7 | BLECF_ | 420 | 0x | 01A4 | 6 | RR77 | BY02 | 7 | CR | 1 | 5 | C | 9-10 | SR7 | BY05 | C | 12 | 9-10 | BY02 | R | 3 | 4 | 1 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | AX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
56 | BJBHT.A5R7 | BLECF_ | 412 | 0x | 019C | 8 | RR77 | BY02 | 7 | CR | 1 | 6 | C | 11-12 | SR7 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | AX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
57 | BJBHT.B5R7 | BLECF_ | 446 | 0x | 01BE | 8 | RR77 | BY02 | 7 | CR | 1 | 7 | C | 13-14 | SR7 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
58 | BJBHT.C5R7 | BLECF_ | 421 | 0x | 01A5 | 6 | RR77 | BY02 | 7 | CR | 1 | 8 | C | 15-16 | SR7 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 9X-AX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
59 | BJBHT.D5R7 | BLECF_ | 415 | 0x | 019F | 6 | RR77 | BY02 | 7 | CR | 1 | 9 | C | 17-18 | SR7 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 16 | 12 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 4X-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
60 | BJBHT.E5R7 | BLECF_ | 368 | 0x | 0170 | 6 | RR77 | BY02 | 7 | CR | 1 | 10 | C | 19-20 | SR7 | BY05 | C | 11 | 7-8 | BY02 | R | 3 | 4 | 1 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 7X-8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
61 | BJBHT.A6R7 | BLECF_ | 334 | 0x | 014E | 4 | RR77 | BY02 | 7 | CR | 2 | 1 | C | 21-22 | SR7 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 17 | 13 | 1-8 | ↑↑↑ | OK |  | ↑ | OK | ↑ | OK | ↑ | OK | N | N | OK | OK | OK | 1X-FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
62 | BJBHT.B6R7 | BLECF_ | 31 | 0x | 001F | 4 | RR77 | BY02 | 7 | CR | 2 | 2 | C | 23-24 | SR7 | BY05 | C | 11 | 11-12 | BY02 | C | 2 | 17 | 13 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
63 | BJBHT.C6R7 | BLECF_ | 350 | 0x | 015E | 5 | RR77 | BY02 | 7 | CR | 2 | 3 | D | 1-2 | SR7 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 18 | 14 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-FF | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
64 | BJBHT.D6R7 | BLECF_ | 81 | 0x | 0051 | 5 | RR77 | BY02 | 7 | CR | 2 | 4 | D | 3-4 | SR7 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 18 | 14 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-5X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
65 | BJBHT.E6R7 | BLECF_ | 394 | 0x | 018A | 8 | RR77 | BY02 | 7 | CR | 2 | 5 | D | 5-6 | SR7 | BY05 | C | 10 | 5-6 | BY02 | C | 2 | 19 | 15 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | EX-FX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
66 | BJBHT.F6R7 | BLECF_ | 13 | 0x | 000D | 8 | RR77 | BY02 | 7 | CR | 2 | 6 | D | 7-8 | SR7 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 19 | 15 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 5X-6X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
67 | BJBHT.G6R7 | BLECF_ | 433 | 0x | 01B1 | 4 | RR77 | BY02 | 7 | CR | 2 | 7 | D | 9-10 | SR7 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 20 | 16 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-EX | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
68 | BJBHT.H6R7 | BLECF_ | 482 | 0x | 01E2 | 4 | RR77 | BY02 | 7 | CR | 2 | 8 | D | 11-12 | SR7 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 20 | 16 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 1X-5X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
69 | BJBHT.I6R7 | BLECF_ | 464 | 0x | 01D0 | 7 | RR77 | BY02 | 7 | CR | 2 | 9 | D | 13-14 | SR7 | BY05 | C | 9 | 1-2 | BY02 | R | 3 | 5 | 2 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 9X-A2 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
70 | BJBHT.A7R7 | BLECF_ | 359 | 0x | 0167 | 2 | RR77 | BY02 | 7 | CR | 2 | 10 | D | 15-16 | SR7 | BY05 | C | 9 | 3-4 | BY02 | E | 4 | 13 | 9 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-3X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
71 | BJBHT.B7R7 | BLECF_ | 470 | 0x | 01D6 | 7 | RR77 | BY02 | 7 | CR | 3 | 1 | D | 17-18 | SR7 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 5 | 2 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
72 | BJBHT.A8R7 | BLECF_ | 700 | 0x | 02BC | 3 | RR77 | BY02 | 7 | CR | 3 | 2 | D | 19-20 | SR7 | BY05 | C | 9 | 7-8 | BY02 | E | 4 | 13 | 9 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
73 | BJBHT.B8R7 | BLECF_ | 358 | 0x | 0166 | 6 | RR77 | BY02 | 7 | CR | 3 | 3 | D | 21-22 | SR7 | BY05 | C | 9 | 9-10 | BY02 | R | 3 | 6 | 3 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK
74 | BJBHT.A9R7 | BLECF_ | 505 | 0x | 01F9 | 8 | RR77 | BY02 | 7 | CR | 3 | 4 | D | 23-24 | SR7 | BY05 | C | 9 | 11-12 | BY02 | E | 4 | 14 | 10 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | LF | OK | OK | N | N | OK | OK | OK | BX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
75 | BJBHT.B9R7 | BLECF_ | 345 | 0x | 0159 | 8 | RR77 | BY02 | 7 | CR | 3 | 5 | E | 1-2 | SR7 | BY05 | C | 8 | 1-2 | BY02 | R | 3 | 6 | 3 | 9-16 | OK | LF | Patch cord in the tunnel broken | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
76 | BJBHT.A10R7 | BLECF_ | 381 | 0x | 017D | 8 | RR77 | BY02 | 7 | CR | 3 | 6 | E | 3-4 | SR7 | BY05 | C | 8 | 3-4 | BY02 | R | 3 | 7 | 4 | 1-8 | ↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
77 | BJBHT.A11R7 | BLECF_ | 428 | 0x | 01AC | 8 | RR77 | BY02 | 7 | CR | 3 | 7 | E | 5-6 | SR7 | BY05 | C | 8 | 5-6 | BY02 | E | 4 | 14 | 10 | 9-16 | ↑↑↑ | LF | Patch cord surface broken | ↑↑↑ | LF | ↑↑↑ | LF | OK | OK | N | N | OK | OK | OK | BX-CX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
78 | BJBHT.B11R7 | BLECF_ | 441 | 0x | 01B9 | 7 | RR77 | BY02 | 7 | CR | 3 | 8 | F | 1-2 | SR7 | BY05 | C | 7 | 1-2 | BY02 | R | 3 | 7 | 4 | 9-16 | LF | LF | was wrong conneted on patch | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | spare fibers from BPM
79 | BJBHT.A12R7 | BLECF_ | 160 | 0x | 00A0 | 6 | Arc | BYPLM.A12R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-39 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
80 | BJBHT.A13R7 | BLECF_ | 119 | 0x | 0077 | 6 | Arc | BYPLM.A13R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | BX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
81 | BJBHT.A14R7 | BLECF_ | 44 | 0x | 002C | 6 | Arc | BYPLM.A14R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | BX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
82 | BJBHT.A15R7 | BLECF_ | 264 | 0x | 0108 | 6 | Arc | BYPLM.A15R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | BX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
83 | BJBHT.A16R7 | BLECF_ | 413 | 0x | 019D | 6 | Arc | BYPLM.A16R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | BX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
84 | BJBHT.A17R7 | BLECF_ | 467 | 0x | 01D3 | 6 | Arc | BYPLM.A17R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X-96 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
85 | BJBHT.A18R7 | BLECF_ | 263 | 0x | 0107 | 6 | Arc | BYPLM.A18R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | ↑↑↑ | OK |  | ↑ | OK | ↑ | OK | ↑ | OK | N | N | OK | OK | OK | 00 | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK
86 | BJBHT.A19R7 | BLECF_ | 291 | 0x | 0123 | 6 | Arc | BYPLM.A19R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X-91 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
87 | BJBHT.A20R7 | BLECF_ | 236 | 0x | 00EC | 6 | Arc | BYPLM.A20R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | ↑↑↑ | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | AX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
88 | BJBHT.A21R7 | BLECF_ | 158 | 0x | 009E | 6 | Arc | BYPLM.A21R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | OK | ↑↑↑ |  | OK | ↑↑↑ | OK | ↑↑↑ | OK | OK | N | N | OK | OK | OK | 9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
89 | BJBHT.A22R7 | BLECF_ | 256 | 0x | 0100 | 6 | Arc | BYPLM.A22R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
90 | BJBHT.A23R7 | BLECF_ | 65 | 0x | 0041 | 6 | Arc | BYPLM.A23R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
91 | BJBHT.A24R7 | BLECF_ | 152 | 0x | 0098 | 6 | Arc | BYPLM.A24R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | ↑↑↑ | OK |  | ↑↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
92 | BJBHT.A25R7 | BLECF_ | 220 | 0x | 00DC | 6 | Arc | BYPLM.A25R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X-BX | ok | ok | ok | ok | ok | ok | ok | ok | [10 empty] | OK | OK | OK
93 | BJBHT.A26R7 | BLECF_ | 133 | 0x | 0085 | 6 | Arc | BYPLM.A26R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | ↑↑ | OK |  | ↑↑↑ | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
94 | BJBHT.A27R7 | BLECF_ | 138 | 0x | 008A | 6 | Arc | BYPLM.A27R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-38 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
95 | BJBHT.A28R7 | BLECF_ | 238 | 0x | 00EE | 6 | Arc | BYPLM.A28R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | BX | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
96 | BJBHT.A29R7 | BLECF_ | 216 | 0x | 00D8 | 6 | Arc | BYPLM.A29R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
97 | BJBHT.A30R7 | BLECF_ | 174 | 0x | 00AE | 6 | Arc | BYPLM.A30R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
98 | BJBHT.A31R7 | BLECF_ | 132 | 0x | 0084 | 6 | Arc | BYPLM.A31R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X-9X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
99 | BJBHT.A32R7 | BLECF_ | 397 | 0x | 018D | 6 | Arc | BYPLM.A32R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | ↑↑↑ | OK |  | ↑ | OK | OK | OK | ↑ | OK | N | N | OK | OK | OK | 3X-42 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
100 | BJBHT.A33R7 | BLECF_ | 2 | 0x | 0002 | 6 | Arc | BYPLM.A33R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | OK | OK |  | OK | OK | OK | OK | OK | OK | N | N | OK | OK | OK | 8X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
101 | BJBHT.A34R7 | BLECF_ | 486 | 0x | 01E6 | 6 | Arc | BYPLM.A34R7 | 7 | R |  |  |  | 7-9 | SR7 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | ↑↑↑ | OK |  | ↑↑↑ | OK | ↑↑ | OK | OK | OK | N | N | OK | OK | OK | 3X-42 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
last entry 22/5/08
| Door | [7 empty] | Transformer soldered | [16 empty] | ↑↑↑ | CRC Error consantly fast increasing
[26 empty] | ↑↑ | CRC Error sometimes fast increasing
[26 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number correct | [19 empty] | LF | Lost Frame
|  |  |  | 0x | 0000 | Card number needs to be checked | [19 empty] | no TC | No Threshold Comparator installed
|  |  |  | 0x | 0000 | Card number unknown need to be checked in the tunnel | [20 empty] | Opitcal link is working with Test system
|  | BLECF_ | 656 | Card was exchanged, Data base need to be checked for the correct number
Note on D118: effinger: 565 original number chnaged to 656
Note on AD20: effinger: CFC card was not working -> reset -> o.k
Note on AE20: effinger: CFC card was not working -> reset -> o.k
Note on AC20: effinger: CFC card was not working -> reset -> o.k
Note on AE8: effinger: CFC card was not working -> reset -> o.k
Note on AD8: effinger: CFC card was not working -> reset -> o.k
Note on AC8: effinger: CFC card was not working -> reset -> o.k
Note on AH94: effinger: 29.8.08 TC exchnaged link working correctly

## Point 8
| Tunnel - position | [13 empty] | Surface position | [25 empty] | LEVEL | [8 empty] | Signal
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | Comment | ERRA | ERRB | ERRA | ERRB | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Comment | Reset | Capa | 220V | Noise | HV | Source | other
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec. | 30.07.2008 |  |  | 23.08.2008 |  |  |  | 26-27.8.08 | [25 empty] | tested | tested |  | levels | tested | test | Comments
1 | BJBHT.A33L8 | BLECF_ | 229 | 0x | 00E5 | 6 | Arc | BYPLM.A33L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
2 | BJBHT.A32L8 | BLECF_ | 237 | 0x | 00ED | 6 | Arc | BYPLM.A32L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
3 | BJBHT.A31L8 | BLECF_ | 292 | 0x | 0124 | 6 | Arc | BYPLM.A31L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
4 | BJBHT.A30L8 | BLECF_ | 257 | 0x | 0101 | 6 | Arc | BYPLM.A30L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
5 | BJBHT.A29L8 | BLECF_ | 215 | 0x | 00D7 | 6 | Arc | BYPLM.A29L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
6 | BJBHT.A28L8 | BLECF_ | 213 | 0x | 00D5 | 6 | Arc | BYPLM.A28L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
7 | BJBHT.A27L8 | BLECF_ | 275 | 0x | 0113 | 6 | Arc | BYPLM.A27L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
8 | BJBHT.A26L8 | BLECF_ | 277 | 0x | 0115 | 6 | Arc | BYPLM.A26L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
9 | BJBHT.A25L8 | BLECF_ | 247 | 0x | 00F7 | 6 | Arc | BYPLM.A25L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | LF | LF | check in tunnel needed | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
10 | BJBHT.A24L8 | BLECF_ | 137 | 0x | 0089 | 6 | Arc | BYPLM.A24L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | LF | OK | check in tunnel needed | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
11 | BJBHT.A23L8 | BLECF_ | 145 | 0x | 0091 | 6 | Arc | BYPLM.A23L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
12 | BJBHT.A22L8 | BLECF_ | 129 | 0x | 0081 | 6 | Arc | BYPLM.A22L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | OK | OK |  | OK | OK |  |  | OK | ↑ | N | N | ↑ | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
13 | BJBHT.A21L8 | BLECF_ | 112 | 0x | 0070 | 6 | Arc | BYPLM.A21L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
14 | BJBHT.A20L8 | BLECF_ | 100 | 0x | 0064 | 6 | Arc | BYPLM.A20L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | OK | OK |  | ↑↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
15 | BJBHT.A19L8 | BLECF_ | 114 | 0x | 0072 | 6 | Arc | BYPLM.A19L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
16 | BJBHT.A18L8 | BLECF_ | 293 | 0x | 0125 | 6 | Arc | BYPLM.A18L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | OK | OK |  | ↑↑↑ | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
17 | BJBHT.A17L8 | BLECF_ | 150 | 0x | 0096 | 6 | Arc | BYPLM.A17L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
18 | BJBHT.A16L8 | BLECF_ | 130 | 0x | 0082 | 6 | Arc | BYPLM.A16L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | OK | OK |  | ↑↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-4D | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
19 | BJBHT.A15L8 | BLECF_ | 101 | 0x | 0065 | 6 | Arc | BYPLM.A15L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
20 | BJBHT.A14L8 | BLECF_ | 105 | 0x | 0069 | 6 | Arc | BYPLM.A14L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | OK | OK |  | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
21 | BJBHT.A13L8 | BLECF_ | 99 | 0x | 0063 | 6 | Arc | BYPLM.A13L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-35 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
22 | BJBHT.A12L8 | BLECF_ | 98 | 0x | 0062 | 6 | Arc | BYPLM.A12L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ↑ | OK | check on surface with TB | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 2X-33 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
23 | BJBHT.C11L8 | BLECF_ | 95 | 0x | 005F | 6 | UA83 | BY05 | 8 | CL | 2 | 9 | C | 13-14 | SR8 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X | ok | ok | ok | ok | ok | ok | ok | ok | TCFC | [9 empty] | OK | OK | OK
24 | BJBHT.B11L8 | BLECF_ | 92 | 0x | 005C | 6 | UA83 | BY05 | 8 | CL | 2 | 8 | C | 11-12 | SR8 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | OK | OK |  | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
25 | BJBHT.A11L8 | BLECF_ | 86 | 0x | 0056 | 6 | UA83 | BY05 | 8 | CL | 2 | 7 | C | 9-10 | SR8 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 8 | 5 | 9-16 | ↑↑↑ | OK | check on surface with TB | ↑↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
26 | BJBHT.A10L8 | BLECF_ | 67 | 0x | 0043 | 6 | UA83 | BY05 | 8 | CL | 2 | 6 | C | 7-8 | SR8 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
27 | BJBHT.A9L8 | BLECF_ | 33 | 0x | 0021 | 6 | UA83 | BY05 | 8 | CL | 2 | 5 | C | 5-6 | SR8 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
28 | BJBHT.C8L8 | BLECF_ | 224 | 0x | 00E0 | 6 | UA83 | BY05 | 8 | CL | 2 | 4 | C | 3-4 | SR8 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8 | OK | OK |  | ↑↑↑ | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
29 | BJBHT.B8L8 | BLECF_ | 205 | 0x | 00CD | 6 | UA83 | BY05 | 8 | CL | 2 | 3 | C | 1-2 | SR8 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
30 | BJBHT.A8L8 | BLECF_ | 151 | 0x | 0097 | 6 | UA83 | BY05 | 8 | CL | 2 | 2 | B | 23-24 | SR8 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 7 | 4 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
31 | BJBHT.A7L8 | BLECF_ | 221 | 0x | 00DD | 6 | UA83 | BY05 | 8 | CL | 2 | 1 | B | 21-22 | SR8 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
32 | BJBHT.B6L8 | BLECF_ | 239 | 0x | 00EF | 7 | UA83 | BY05 | 8 | CL | 1 | 10 | B | 19-20 | SR8 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | OK | OK |  | ↑↑↑ | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
33 | BJBHT.A6L8 | BLECF_ | 231 | 0x | 00E7 | 1 | UA83 | BY05 | 8 | CL | 1 | 9 | B | 17-18 | SR8 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 7 | 4 | 1-8 | ↑ | OK | check on surface with TB | ↑↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-28 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
34 | BJBHT.A5L8 | BLECF_ | 240 | 0x | 00F0 | 6 | UA83 | BY05 | 8 | CL | 1 | 8 | B | 15-16 | SR8 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
35 | BJBHT.C4L8 | BLECF_ | 246 | 0x | 00F6 | 8 | UA83 | BY05 | 8 | CL | 1 | 7 | B | 13-14 | SR8 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
36 | BJBHT.B4L8 | BLECF_ | 248 | 0x | 00F8 | 1 | UA83 | BY05 | 8 | CL | 1 | 6 | B | 11-12 | SR8 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | OK | ↑↑↑ | check on surface with TB | OK | ↑↑↑ |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X-2C | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
37 | BJBHT.A4L8 | BLECF_ | 243 | 0x | 00F3 | 2 | UA83 | BY05 | 8 | CL | 1 | 5 | B | 9-10 | SR8 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | ↑↑↑ | OK | check on surface with TB | ↑↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-2X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
38 | BJBHT.A3L8 | BLECF_ | 250 | 0x | 00FA | 8 | UA83 | BY05 | 8 | CL | 1 | 4 | B | 7-8 | SR8 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
39 | BJBHT.A2L8 | BLECF_ | 261 | 0x | 0105 | 8 | UA83 | BY05 | 8 | CL | 1 | 3 | B | 5-6 | SR8 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | OK | OK |  | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X-8A | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
40 | BJBHT.B1L8 | BLECF_ | 269 | 0x | 010D | 5 | UA83 | BY05 | 8 | CL | 1 | 2 | B | 3-4 | SR8 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | OK | ↑↑↑ | check on surface with TB | OK | ↑↑↑ |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
41 | BJBHT.A1L9 | BLECF_ | 265 | 0x | 0109 | 1 | UA83 | BY05 | 8 | CL | 1 | 1 | B | 1-2 | SR8 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-2D | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
42 | BJBHT.A1R8 | BLECF_ | 78 | 0x | 004E | 1 | UA87 | BY06 | 8 | CR | 1 | 1 | B | 1-2 | SR8 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X-2X-32 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
43 | BJBHT.B1R8 | BLECF_ | 75 | 0x | 004B | 5 | UA87 | BY06 | 8 | CR | 1 | 2 | B | 3-4 | SR8 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-4X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
44 | BJBHT.A2R8 | BLECF_ | 70 | 0x | 0046 | 8 | UA87 | BY06 | 8 | CR | 1 | 3 | B | 5-6 | SR8 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X-4C | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
45 | BJBHT.A3R8 | BLECF_ | 228 | 0x | 00E4 | 4 | UA87 | BY06 | 8 | CR | 1 | 4 | B | 7-8 | SR8 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X-2X-47 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
46 | BJBHT.A4R8 | BLECF_ | 71 | 0x | 0047 | 7 | UA87 | BY06 | 8 | CR | 1 | 5 | B | 9-10 | SR8 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | LF | OK | check in tunnel needed | OK | OK |  |  | OK | OK | N | Y | OK | OK | OK | 2X-9F | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
47 | BJBHT.B4R8 | BLECF_ | 287 | 0x | 011F | 4 | UA87 | BY06 | 8 | CR | 1 | 6 | B | 11-12 | SR8 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | OK | ↑↑↑ | check on surface with TB | OK | ↑↑↑ |  |  | OK | OK | N | N | OK | OK | OK | 1X-29 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
48 | BJBHT.C4R8 | BLECF_ | 278 | 0x | 0116 | 1 | UA87 | BY06 | 8 | CR | 1 | 7 | B | 13-14 | SR8 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X-26 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
49 | BJBHT.D4R8 | BLECF_ | 209 | 0x | 00D1 | 8 | UA87 | BY06 | 8 | CR | 1 | 8 | B | 15-16 | SR8 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8 | ↑ | OK | check on surface with TB | ↑↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKI
50 | BJBHT.A5R8 | BLECF_ | 233 | 0x | 00E9 | 6 | UA87 | BY06 | 8 | CR | 1 | 9 | B | 17-18 | SR8 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 6X-79 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKI
51 | BJBHT.A6R8 | BLECF_ | 267 | 0x | 010B | 6 | UA87 | BY06 | 8 | CR | 1 | 10 | B | 19-20 | SR8 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKI
52 | BJBHT.B6R8 | BLECF_ | 29 | 0x | 001D | 6 | UA87 | BY06 | 8 | CR | 2 | 1 | B | 21-22 | SR8 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 17 | 13 | 1-8 | ↑ | OK | check on surface with TB | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK |  |  |  | MKI
53 | BJBHT.C6R8 | BLECF_ | 27 | 0x | 001B | 6 | UA87 | BY06 | 8 | CR | 2 | 2 | B | 23-24 | SR8 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 1-8 | OK | OK |  | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 2X-33 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
54 | BJBHT.A7R8 | BLECF_ | 26 | 0x | 001A | 6 | UA87 | BY06 | 8 | CR | 2 | 3 | C | 1-2 | SR8 | BY05 | C | 10 | 1-2 | BY02 | R | 3 | 5 | 2 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
55 | BJBHT.A8R8 | BLECF_ | 24 | 0x | 0018 | 6 | UA87 | BY06 | 8 | CR | 2 | 4 | C | 3-4 | SR8 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
56 | BJBHT.B8R8 | BLECF_ | 25 | 0x | 0019 | 6 | UA87 | BY06 | 8 | CR | 2 | 5 | C | 5-6 | SR8 | BY05 | C | 10 | 5-6 | BY02 | C | 2 | 18 | 14 | 1-8 | ↑ | OK | check on surface with TB | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
57 | BJBHT.C8R8 | BLECF_ | 127 | 0x | 007F | 6 | UA87 | BY06 | 8 | CR | 2 | 6 | C | 7-8 | SR8 | BY05 | C | 10 | 7-8 | BY02 | R | 3 | 6 | 3 | 1-8 | OK | OK |  | ↑↑↑ | OK |  |  | OK | ↑ | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
58 | BJBHT.A9R8 | BLECF_ | 80 | 0x | 0050 | 6 | UA87 | BY06 | 8 | CR | 2 | 7 | C | 9-10 | SR8 | BY05 | C | 10 | 9-10 | BY02 | R | 3 | 6 | 3 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
59 | BJBHT.A10R8 | BLECF_ | 18 | 0x | 0012 | 6 | UA87 | BY06 | 8 | CR | 2 | 8 | C | 11-12 | SR8 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 7 | 4 | 1-8 | ↑ | OK | check on surface with TB | ↑↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
60 | BJBHT.A11R8 | BLECF_ | 120 | 0x | 0078 | 6 | UA87 | BY06 | 8 | CR | 2 | 9 | C | 13-14 | SR8 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
61 | BJBHT.B11R8 | BLECF_ | 113 | 0x | 0071 | 6 | UA87 | BY06 | 8 | CR | 2 | 10 | C | 15-16 | SR8 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | IN TCFC | [9 empty] | OK | OK | OK
62 | BJBHT.C11R8 | BLECF_ | 300 | 0x | 012C | 6 | UA87 | BY06 | 8 | CR | 3 | 1 | C | 17-18 | SR8 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 7 | 4 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
63 | BJBHT.A12R8 | BLECF_ | 272 | 0x | 0110 | 6 | Arc | BYPLM.A12R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
64 | BJBHT.A13R8 | BLECF_ | 51 | 0x | 0033 | 6 | Arc | BYPLM.A13R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
65 | BJBHT.A14R8 | BLECF_ | 54 | 0x | 0036 | 6 | Arc | BYPLM.A14R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | OK | ↑↑↑ | check on surface with TB | OK | ↑↑↑ |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
66 | BJBHT.A15R8 | BLECF_ | 30 | 0x | 001E | 6 | Arc | BYPLM.A15R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | ↑↑ | OK | check on surface with TB | ↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
67 | BJBHT.A16R8 | BLECF_ | 142 | 0x | 008E | 6 | Arc | BYPLM.A16R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
68 | BJBHT.A17R8 | BLECF_ | 61 | 0x | 003D | 6 | Arc | BYPLM.A17R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
69 | BJBHT.A18R8 | BLECF_ | 148 | 0x | 0094 | 6 | Arc | BYPLM.A18R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | OK | OK |  | OK | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 2X-31 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
70 | BJBHT.A19R8 | BLECF_ | 15 | 0x | 000F | 6 | Arc | BYPLM.A19R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
71 | BJBHT.A20R8 | BLECF_ | 294 | 0x | 0126 | 6 | Arc | BYPLM.A20R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | ↑ | OK | check on surface with TB | ↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
72 | BJBHT.A21R8 | BLECF_ | 286 | 0x | 011E | 6 | Arc | BYPLM.A21R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
73 | BJBHT.A22R8 | BLECF_ | 206 | 0x | 00CE | 6 | Arc | BYPLM.A22R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | ↑↑↑ | OK | check on surface with TB | ↑↑↑ | OK |  |  | ↑ | OK | N | N | OK | OK | OK | 1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
74 | BJBHT.A23R8 | BLECF_ | 270 | 0x | 010E | 6 | Arc | BYPLM.A23R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
75 | BJBHT.A24R8 | BLECF_ | 131 | 0x | 0083 | 6 | Arc | BYPLM.A24R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | ↑↑ | OK | check on surface with TB | ↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 00 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
76 | BJBHT.A25R8 | BLECF_ | 227 | 0x | 00E3 | 6 | Arc | BYPLM.A25R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-11 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
77 | BJBHT.A26R8 | BLECF_ | 253 | 0x | 00FD | 6 | Arc | BYPLM.A26R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
78 | BJBHT.A27R8 | BLECF_ | 208 | 0x | 00D0 | 6 | Arc | BYPLM.A27R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
79 | BJBHT.A28R8 | BLECF_ | 153 | 0x | 0099 | 6 | Arc | BYPLM.A28R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
80 | BJBHT.A29R8 | BLECF_ | 106 | 0x | 006A | 6 | Arc | BYPLM.A29R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
81 | BJBHT.A30R8 | BLECF_ | 252 | 0x | 00FC | 6 | Arc | BYPLM.A30R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 0X-1X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
82 | BJBHT.A31R8 | BLECF_ | 104 | 0x | 0068 | 6 | Arc | BYPLM.A31R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
83 | BJBHT.A32R8 | BLECF_ | 271 | 0x | 010F | 6 | Arc | BYPLM.A32R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | ↑↑ | OK | check on surface with TB | ↑↑ | OK |  |  | OK | ↑ | N | N | ↑ | OK | OK | 2X-35 | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
84 | BJBHT.A33R8 | BLECF_ | 225 | 0x | 00E1 | 6 | Arc | BYPLM.A33R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | OK | OK |  | OK | OK |  |  | OK | OK | N | N | OK | OK | OK | 2X-3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
85 | BJBHT.A34R8 | BLECF_ | 297 | 0x | 0129 | 6 | Arc | BYPLM.A34R8 | 8 | R |  |  |  | 7-9 | SR8 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8 | OK | OK |  | ↑↑ | OK |  |  | OK | OK | N | N | OK | OK | OK | 3X | ok | ok | ok | ok | ok | ok | ok | ok | ok | [9 empty] | OK | OK | OK
last entry 14/4/08
[9 empty] | Transformer soldered | [16 empty] | ↑↑↑ | CRC Error consantly fast increasing
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [9 empty] | Splitter used
[26 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number correct | [19 empty] | LF | Lost Frame
|  |  |  | 0x | 0000 | Card number needs to be checked | [19 empty] | no TC | No Threshold Comparator installed
|  |  |  | 0x | 0000 | Card number unknown need to be checked in the tunnel | [20 empty] | Opitcal link is working with Test system
|  | BLECF_ | 656 | Card was exchanged, Data base need to be checked for the correct number
Note on D102: effinger: 565 original number chnaged to 656
Note on AA33: effinger: ERROR FID
Note on AB72: effinger: ERROR C ERROR FID
Note on AA80: effinger: ERROR C ERROR FID ERROR CID
Note on AB45: effinger: ERROR C ERROR FID
Note on AA42: effinger: ERROR C ERROR FID
Note on AB41: effinger: ERROR C ERROR FID
Note on AA30: effinger: ERROR C ERROR FID
Note on AB53: effinger: ERROR C ERRO CID ERROR FID
Note on AE45: effinger: ERROR C ERROR FID
Note on AE41: effinger: ERROR C ERROR FID
Note on AD42: effinger: ERROR C ERROR FID
Note on AD30: effinger: ERROR C ERROR FID
Note on AE53: effinger: ERROR C ERRO CID ERROR FID
Note on AE72: effinger: ERROR C ERROR FID
Note on AD80: effinger: ERROR C ERROR FID ERROR CID
Note on AD82: effinger: ERROR FID ERROR C
Note on AH14: effinger: 29.8.08 TC exchanged -> links are working correctly
Note on AK52: effinger: 28.8.08 Need to be checke in the tunnel

## Overview
267 | 0 | 457
551 | BLECF_ | #REF!
578 | BLECF_ | 757
435 | BLECF_ | 754
574 | BLECF_ | 753
596 | BLECF_ | 750
594 | BLECF_ | 749
492 | BLECF_ | 748
519 | BLECF_ | 744
502 | BLECF_ | 743
619 | BLECF_ | 742
422 | BLECF_ | 741
622 | BLECF_ | 739
615 | BLECF_ | 737
625 | BLECF_ | 736
632 | BLECF_ | 735
628 | BLECF_ | 734
616 | BLECF_ | 733
618 | BLECF_ | 732
633 | BLECF_ | 731
449 | BLECF_ | 729
613 | BLECF_ | 728
412 | BLECF_ | 727
597 | BLECF_ | 726
404 | BLECF_ | 725
409 | BLECF_ | 724
418 | BLECF_ | 717
448 | BLECF_ | 711
599 | BLECF_ | 710
64 | BLECF_ | 708
413 | BLECF_ | 705
630 | BLECF_ | 704
431 | BLECF_ | 703
415 | BLECF_ | 702
629 | BLECF_ | 701
612 | BLECF_ | 699
614 | BLECF_ | 698
601 | BLECF_ | 697
620 | BLECF_ | 695
402 | BLECF_ | 694
631 | BLECF_ | 693
623 | BLECF_ | 691
406 | BLECF_ | 690
621 | BLECF_ | 688
388 | BLECF_ | 687
391 | BLECF_ | 686
403 | BLECF_ | 685
389 | BLECF_ | 684
626 | BLECF_ | 683
651 | BLECF_ | 682
387 | BLECF_ | 681
559 | BLECF_ | 680
636 | BLECF_ | 679
638 | BLECF_ | 678
451 | BLECF_ | 677
285 | BLECF_ | 676
434 | BLECF_ | 674
609 | BLECF_ | 673
437 | BLECF_ | 672
595 | BLECF_ | 671
510 | BLECF_ | 670
426 | BLECF_ | 668
262 | BLECF_ | 667
444 | BLECF_ | 175
503 | BLECF_ | 664
598 | BLECF_ | 663
494 | BLECF_ | 662
634 | BLECF_ | 661
606 | BLECF_ | 660
512 | BLECF_ | 659
432 | BLECF_ | 657
603 | BLECF_ | 656
500 | BLECF_ | 655
433 | BLECF_ | 654
286 | BLECF_ | 653
516 | BLECF_ | 651
554 | BLECF_ | 650
640 | BLECF_ | 649
602 | BLECF_ | 648
294 | BLECF_ | 647
414 | BLECF_ | 646
421 | BLECF_ | 645
533 | BLECF_ | 644
513 | BLECF_ | 643
509 | BLECF_ | 642
572 | BLECF_ | 640
280 | BLECF_ | 639
440 | BLECF_ | 637
284 | BLECF_ | 636
279 | BLECF_ | 635
617 | BLECF_ | 634
661 | BLECF_ | 633
582 | BLECF_ | 632
501 | BLECF_ | 631
425 | BLECF_ | 630
442 | BLECF_ | 629
563 | BLECF_ | 625
417 | BLECF_ | 623
428 | BLECF_ | 622
610 | BLECF_ | 621
424 | BLECF_ | 620
410 | BLECF_ | 619
604 | BLECF_ | 618
624 | BLECF_ | 617
493 | BLECF_ | 616
507 | BLECF_ | 615
423 | BLECF_ | 614
499 | BLECF_ | 613
1 | BLECF_ | 612
469 | BLECF_ | 611
2 | BLECF_ | 610
436 | BLECF_ | 609
416 | BLECF_ | 607
608 | BLECF_ | 605
427 | BLECF_ | 604
611 | BLECF_ | 603
447 | BLECF_ | 602
411 | BLECF_ | 601
430 | BLECF_ | 600
511 | BLECF_ | 599
600 | BLECF_ | 598
273 | BLECF_ | 597
605 | BLECF_ | 596
276 | BLECF_ | 595
263 | BLECF_ | 594
277 | BLECF_ | 593
278 | BLECF_ | 592
508 | BLECF_ | 591
656 | BLECF_ | 590
565 | BLECF_ | 589
268 | BLECF_ | 588
644 | BLECF_ | 587
591 | BLECF_ | 586
537 | BLECF_ | 585
524 | BLECF_ | 584
549 | BLECF_ | 582
575 | BLECF_ | 581
120 | BLECF_ | 580
592 | BLECF_ | 579
514 | BLECF_ | 578
445 | BLECF_ | 577
439 | BLECF_ | 576
446 | BLECF_ | 575
438 | BLECF_ | 574
497 | BLECF_ | 572
3 | BLECF_ | 571
301 | BLECF_ | 570
4 | BLECF_ | 569
531 | BLECF_ | 568
297 | BLECF_ | 567
5 | BLECF_ | 566
6 | BLECF_ | 565
7 | BLECF_ | 564
8 | BLECF_ | 563
466 | BLECF_ | 562
9 | BLECF_ | 561
648 | BLECF_ | 560
639 | BLECF_ | 559
535 | BLECF_ | 558
10 | BLECF_ | 557
523 | BLECF_ | 556
11 | BLECF_ | 555
650 | BLECF_ | 553
505 | BLECF_ | 552
544 | BLECF_ | 551
265 | BLECF_ | 550
12 | BLECF_ | 549
13 | BLECF_ | 548
543 | BLECF_ | 547
566 | BLECF_ | 546
14 | BLECF_ | 545
637 | BLECF_ | 544
450 | BLECF_ | 543
15 | BLECF_ | 542
456 | BLECF_ | 541
16 | BLECF_ | 540
299 | BLECF_ | 539
17 | BLECF_ | 538
540 | BLECF_ | 537
443 | BLECF_ | 536
18 | BLECF_ | 534
19 | BLECF_ | 533
20 | BLECF_ | 532
530 | BLECF_ | 531
272 | BLECF_ | 530
21 | BLECF_ | 527
581 | BLECF_ | 525
476 | BLECF_ | 523
308 | BLECF_ | 522
585 | BLECF_ | 521
326 | BLECF_ | 520
483 | BLECF_ | 519
266 | BLECF_ | 518
349 | BLECF_ | 517
361 | BLECF_ | 516
305 | BLECF_ | 515
472 | BLECF_ | 514
642 | BLECF_ | 513
292 | BLECF_ | 512
23 | BLECF_ | 511
364 | BLECF_ | 510
525 | BLECF_ | 509
24 | BLECF_ | 508
564 | BLECF_ | 506
686 | BLECF_ | 505
470 | BLECF_ | 504
461 | BLECF_ | 502
283 | BLECF_ | 501
312 | BLECF_ | 500
25 | BLECF_ | 499
471 | BLECF_ | 499
324 | BLECF_ | 497
474 | BLECF_ | 496
318 | BLECF_ | 495
344 | BLECF_ | 494
473 | BLECF_ | 493
460 | BLECF_ | 709
352 | BLECF_ | 491
627 | BLECF_ | 490
485 | BLECF_ | 489
487 | BLECF_ | 488
454 | BLECF_ | 487
26 | BLECF_ | 486
479 | BLECF_ | 485
390 | BLECF_ | 484
680 | BLECF_ | 482
306 | BLECF_ | 481
465 | BLECF_ | 480
313 | BLECF_ | 479
314 | BLECF_ | 478
573 | BLECF_ | 477
365 | BLECF_ | 476
568 | BLECF_ | 475
356 | BLECF_ | 474
567 | BLECF_ | 473
468 | BLECF_ | 472
504 | BLECF_ | 471
683 | BLECF_ | 470
635 | BLECF_ | 469
332 | BLECF_ | 468
27 | BLECF_ | 467
29 | BLECF_ | 466
652 | BLECF_ | 465
681 | BLECF_ | 464
393 | BLECF_ | 463
641 | BLECF_ | 462
28 | BLECF_ | 461
329 | BLECF_ | 460
467 | BLECF_ | 459
333 | BLECF_ | 456
664 | BLECF_ | 455
30 | BLECF_ | 454
489 | BLECF_ | 453
491 | BLECF_ | 452
396 | BLECF_ | 451
31 | BLECF_ | 449
484 | BLECF_ | 448
557 | BLECF_ | 447
669 | BLECF_ | 446
659 | BLECF_ | 445
340 | BLECF_ | 444
663 | BLECF_ | 443
289 | BLECF_ | 442
690 | BLECF_ | 441
331 | BLECF_ | 440
327 | BLECF_ | 439
542 | BLECF_ | 438
288 | BLECF_ | 437
32 | BLECF_ | 436
546 | BLECF_ | 435
647 | BLECF_ | 434
679 | BLECF_ | 433
325 | BLECF_ | 432
666 | BLECF_ | 431
33 | BLECF_ | 430
34 | BLECF_ | 429
689 | BLECF_ | 428
338 | BLECF_ | 427
35 | BLECF_ | 426
36 | BLECF_ | 425
560 | BLECF_ | 424
310 | BLECF_ | 423
684 | BLECF_ | 700
670 | BLECF_ | 421
667 | BLECF_ | 420
561 | BLECF_ | 419
399 | BLECF_ | 418
320 | BLECF_ | 417
671 | BLECF_ | 415
555 | BLECF_ | 414
37 | BLECF_ | 413
668 | BLECF_ | 412
643 | BLECF_ | 411
464 | BLECF_ | 410
264 | BLECF_ | 409
515 | BLECF_ | 408
488 | BLECF_ | 407
541 | BLECF_ | 405
397 | BLECF_ | 404
534 | BLECF_ | 403
38 | BLECF_ | 402
486 | BLECF_ | 401
39 | BLECF_ | 400
322 | BLECF_ | 399
274 | BLECF_ | 398
40 | BLECF_ | 397
506 | BLECF_ | 395
677 | BLECF_ | 394
455 | BLECF_ | 393
41 | BLECF_ | 392
395 | BLECF_ | 382
42 | BLECF_ | 391
400 | BLECF_ | 390
550 | BLECF_ | 389
330 | BLECF_ | 388
478 | BLECF_ | 387
517 | BLECF_ | 386
339 | BLECF_ | 385
319 | BLECF_ | 384
290 | BLECF_ | 383
688 | BLECF_ | 381
341 | BLECF_ | 380
394 | BLECF_ | 379
346 | BLECF_ | 378
281 | BLECF_ | 377
287 | BLECF_ | 376
463 | BLECF_ | 375
275 | BLECF_ | 374
309 | BLECF_ | 373
658 | BLECF_ | 372
358 | BLECF_ | 371
587 | BLECF_ | 370
660 | BLECF_ | 369
672 | BLECF_ | 368
459 | BLECF_ | 367
337 | BLECF_ | 366
22 | BLECF_ | 364
348 | BLECF_ | 363
353 | BLECF_ | 362
300 | BLECF_ | 361
407 | BLECF_ | 360
682 | BLECF_ | 359
685 | BLECF_ | 358
350 | BLECF_ | 357
586 | BLECF_ | 356
649 | BLECF_ | 355
311 | BLECF_ | 352
665 | BLECF_ | 351
675 | BLECF_ | 350
298 | BLECF_ | 349
323 | BLECF_ | 348
342 | BLECF_ | 347
570 | BLECF_ | 346
687 | BLECF_ | 345
355 | BLECF_ | 344
345 | BLECF_ | 343
328 | BLECF_ | 342
457 | BLECF_ | 341
539 | BLECF_ | 340
458 | BLECF_ | 339
296 | BLECF_ | 338
482 | BLECF_ | 337
43 | BLECF_ | 336
552 | BLECF_ | 335
673 | BLECF_ | 334
351 | BLECF_ | 333
282 | BLECF_ | 331
553 | BLECF_ | 330
317 | BLECF_ | 329
357 | BLECF_ | 328
453 | BLECF_ | 326
293 | BLECF_ | 325
362 | BLECF_ | 324
291 | BLECF_ | 323
269 | BLECF_ | 322
321 | BLECF_ | 321
360 | BLECF_ | 320
419 | BLECF_ | 319
498 | BLECF_ | 318
307 | BLECF_ | 317
522 | BLECF_ | 316
538 | BLECF_ | 315
475 | BLECF_ | 314
584 | BLECF_ | 313
590 | BLECF_ | 312
303 | BLECF_ | 311
316 | BLECF_ | 310
480 | BLECF_ | 309
662 | BLECF_ | 308
556 | BLECF_ | 307
302 | BLECF_ | 306
429 | BLECF_ | 305
315 | BLECF_ | 304
304 | BLECF_ | 303
133 | BLECF_ | 301
44 | BLECF_ | 300
45 | BLECF_ | 299
46 | BLECF_ | 298
47 | BLECF_ | 297
48 | BLECF_ | 295
49 | BLECF_ | 294
50 | BLECF_ | 293
51 | BLECF_ | 292
52 | BLECF_ | 291
53 | BLECF_ | 290
54 | BLECF_ | 289
55 | BLECF_ | 288
56 | BLECF_ | 287
57 | BLECF_ | 286
58 | BLECF_ | 285
59 | BLECF_ | 284
60 | BLECF_ | 212
61 | BLECF_ | 282
62 | BLECF_ | 281
63 | BLECF_ | 280
65 | BLECF_ | 278
66 | BLECF_ | 277
67 | BLECF_ | 276
68 | BLECF_ | 275
69 | BLECF_ | 274
655 | BLECF_ | 273
70 | BLECF_ | 272
71 | BLECF_ | 271
72 | BLECF_ | 270
73 | BLECF_ | 269
74 | BLECF_ | 268
75 | BLECF_ | 267
76 | BLECF_ | 266
77 | BLECF_ | 265
78 | BLECF_ | 264
79 | BLECF_ | 263
80 | BLECF_ | 262
81 | BLECF_ | 261
82 | BLECF_ | 260
83 | BLECF_ | 258
84 | BLECF_ | 257
85 | BLECF_ | 256
86 | BLECF_ | 255
87 | BLECF_ | 253
88 | BLECF_ | 252
89 | BLECF_ | 251
90 | BLECF_ | 250
91 | BLECF_ | 249
92 | BLECF_ | 248
93 | BLECF_ | 247
94 | BLECF_ | 246
95 | BLECF_ | 245
96 | BLECF_ | 244
97 | BLECF_ | 243
98 | BLECF_ | 242
99 | BLECF_ | 241
100 | BLECF_ | 240
101 | BLECF_ | 239
102 | BLECF_ | 238
103 | BLECF_ | 237
104 | BLECF_ | 236
105 | BLECF_ | 235
106 | BLECF_ | 233
107 | BLECF_ | 232
108 | BLECF_ | 231
109 | BLECF_ | 230
110 | BLECF_ | 229
111 | BLECF_ | 228
112 | BLECF_ | 227
558 | BLECF_ | 226
113 | BLECF_ | 225
114 | BLECF_ | 224
115 | BLECF_ | 223
116 | BLECF_ | 221
117 | BLECF_ | 220
118 | BLECF_ | 219
119 | BLECF_ | 218
529 | BLECF_ | 217
121 | BLECF_ | 216
122 | BLECF_ | 215
401 | BLECF_ | 214
123 | BLECF_ | 213
441 | BLECF_ | 211
124 | BLECF_ | 210
125 | BLECF_ | 209
126 | BLECF_ | 208
127 | BLECF_ | 207
128 | BLECF_ | 206
129 | BLECF_ | 205
130 | BLECF_ | 204
131 | BLECF_ | 202
132 | BLECF_ | 201
343 | BLECF_ | 200
536 | BLECF_ | 199
134 | BLECF_ | 198
135 | BLECF_ | 197
569 | BLECF_ | 196
405 | BLECF_ | 195
270 | BLECF_ | 194
593 | BLECF_ | 193
548 | BLECF_ | 192
646 | BLECF_ | 191
579 | BLECF_ | 190
354 | BLECF_ | 189
363 | BLECF_ | 188
136 | BLECF_ | 187
520 | BLECF_ | 186
477 | BLECF_ | 185
481 | BLECF_ | 184
462 | BLECF_ | 183
452 | BLECF_ | 182
359 | BLECF_ | 181
137 | BLECF_ | 180
138 | BLECF_ | 179
139 | BLECF_ | 178
140 | BLECF_ | 177
141 | BLECF_ | 176
142 | BLECF_ | 174
143 | BLECF_ | 173
144 | BLECF_ | 172
145 | BLECF_ | 171
146 | BLECF_ | 170
147 | BLECF_ | 169
148 | BLECF_ | 168
149 | BLECF_ | 167
150 | BLECF_ | 166
151 | BLECF_ | 165
152 | BLECF_ | 164
153 | BLECF_ | 163
154 | BLECF_ | 162
155 | BLECF_ | 161
156 | BLECF_ | 160
157 | BLECF_ | 159
158 | BLECF_ | 158
518 | BLECF_ | 157
159 | BLECF_ | 156
160 | BLECF_ | 153
161 | BLECF_ | 152
162 | BLECF_ | 151
163 | BLECF_ | 150
607 | BLECF_ | 149
164 | BLECF_ | 148
165 | BLECF_ | 145
166 | BLECF_ | 144
392 | BLECF_ | 143
167 | BLECF_ | 142
168 | BLECF_ | 141
169 | BLECF_ | 140
170 | BLECF_ | 138
171 | BLECF_ | 138
172 | BLECF_ | 137
347 | BLECF_ | 134
173 | BLECF_ | 133
174 | BLECF_ | 132
175 | BLECF_ | 131
176 | BLECF_ | 130
177 | BLECF_ | 129
178 | BLECF_ | 127
179 | BLECF_ | 126
180 | BLECF_ | 125
577 | BLECF_ | 124
181 | BLECF_ | 123
182 | BLECF_ | 122
589 | BLECF_ | 121
183 | BLECF_ | 120
184 | BLECF_ | 119
645 | BLECF_ | 118
185 | BLECF_ | 117
186 | BLECF_ | 116
187 | BLECF_ | 115
188 | BLECF_ | 114
189 | BLECF_ | 113
190 | BLECF_ | 112
334 | BLECF_ | 111
191 | BLECF_ | 110
490 | BLECF_ | 109
192 | BLECF_ | 108
193 | BLECF_ | 107
194 | BLECF_ | 106
195 | BLECF_ | 105
196 | BLECF_ | 104
583 | BLECF_ | 103
197 | BLECF_ | 102
198 | BLECF_ | 101
199 | BLECF_ | 100
200 | BLECF_ | 99
201 | BLECF_ | 98
653 | BLECF_ | 96
202 | BLECF_ | 95
654 | BLECF_ | 94
203 | BLECF_ | 93
204 | BLECF_ | 92
205 | BLECF_ | 91
206 | BLECF_ | 88
207 | BLECF_ | 87
208 | BLECF_ | 86
295 | BLECF_ | 85
398 | BLECF_ | 84
545 | BLECF_ | 83
209 | BLECF_ | 82
676 | BLECF_ | 81
210 | BLECF_ | 80
211 | BLECF_ | 79
212 | BLECF_ | 78
213 | BLECF_ | 77
496 | BLECF_ | 76
214 | BLECF_ | 75
215 | BLECF_ | 74
216 | BLECF_ | 72
217 | BLECF_ | 71
218 | BLECF_ | 70
219 | BLECF_ | 69
220 | BLECF_ | 68
221 | BLECF_ | 67
222 | BLECF_ | 66
223 | BLECF_ | 65
224 | BLECF_ | 63
225 | BLECF_ | 61
420 | BLECF_ | 60
226 | BLECF_ | 59
227 | BLECF_ | 58
532 | BLECF_ | 57
228 | BLECF_ | 56
229 | BLECF_ | 54
230 | BLECF_ | 53
231 | BLECF_ | 52
232 | BLECF_ | 51
233 | BLECF_ | 49
571 | BLECF_ | 48
234 | BLECF_ | 47
235 | BLECF_ | 46
236 | BLECF_ | 45
237 | BLECF_ | 44
238 | BLECF_ | 43
562 | BLECF_ | 42
239 | BLECF_ | 41
547 | BLECF_ | 40
240 | BLECF_ | 39
495 | BLECF_ | 38
241 | BLECF_ | 37
242 | BLECF_ | 36
243 | BLECF_ | 35
244 | BLECF_ | 34
245 | BLECF_ | 33
674 | BLECF_ | 31
246 | BLECF_ | 30
247 | BLECF_ | 29
248 | BLECF_ | 28
249 | BLECF_ | 27
250 | BLECF_ | 26
251 | BLECF_ | 25
252 | BLECF_ | 24
408 | BLECF_ | 23
253 | BLECF_ | 22
254 | BLECF_ | 21
255 | BLECF_ | 20
256 | BLECF_ | 19
257 | BLECF_ | 18
258 | BLECF_ | 17
259 | BLECF_ | 15
576 | BLECF_ | 14
678 | BLECF_ | 13
260 | BLECF_ | 12
335 | BLECF_ | 11
527 | BLECF_ | 10
580 | BLECF_ | 9
526 | BLECF_ | 7
521 | BLECF_ | 5
588 | BLECF_ | 3
261 | BLECF_ | 2
657 | BLECF_ | 1
271 | BLECF_ | 638
336 | BLECF_ | 715
366 | BLECF_ | 706
367 | BLECF_ | 606
368 | BLECF_ | 675
369 | BLECF_ | 641
370 | BLECF_ | 719
371 | BLECF_ | 738
372 | BLECF_ | 652
373 | BLECF_ | 665
374 | BLECF_ | 740
375 | BLECF_ | 666
376 | BLECF_ | 722
377 | BLECF_ | 658
378 | BLECF_ | 626
379 | BLECF_ | 406
380 | BLECF_ | 4
381 | BLECF_ | 279
382 | BLECF_ | 97
383 | BLECF_ | 721
384 | BLECF_ | 128
385 | BLECF_ | 627
386 | BLECF_ | 689
528 | BLECF_ | 0

## Spare
| Tunnel - position | [11 empty] | Surface position
| HV box | CFC-Card |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  | VME-TC position
Pos. | position | No. | Hex. | Ch | Place | Rack | P. | Pos. | Crate | Slot |  | No. | Bldg. | Crate | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | [21 empty] | Installed LHC | 669
2 | [21 empty] | Installed SPS, PS, Experiments | 8
3
4 | [21 empty] | Installed CMS, Atlas, LHCb, Alice | 16
5
6
7 | [21 empty] | In the Lab:
8 | [21 empty] | O.K. | 11 | in the lab (sum) | 81
9 | [21 empty] | Endtest missing | 12
10 | [21 empty] | 10pA calibration missing | 0
11 | [21 empty] | little problem | 31
12 | [21 empty] | big problem | 9
13 | [21 empty] | Short circuit | 5
14 | [21 empty] | in the lab not mounted | 11
15 | [21 empty] | in testbench cris | 1
16
17 | [21 empty] | Ewald office | 1
18
19 | [21 empty] | total | 774
20
21 | [21 empty] | date | 2008-08-05
22
last entry 11/2/08

## SPS
| Tunnel - position | [13 empty] | Surface position
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  | VME-TC position
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Crate | Slot |  | No. | Bldg. | Crate | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BA5 | BLECF_ | 235 | 0x | 00EB |  |  | BA5 | 5 | [5 empty] | BA5
2 | BA5 | BLECF_ | 242 | 0x | 00F2 |  |  | BA5 | 5 | [5 empty] | BA5
3 | BA5 | BLECF_ | 244 | 0x | 00F4 |  |  | BA5 | 5 | [5 empty] | BA5
4 | BA5 | BLECF_ | 245 | 0x | 00F5 |  |  | BA5 | 5 | [5 empty] | BA5
5 | BA5 | BLECF_ | 251 | 0x | 00FB |  |  | BA5 | 5 | [5 empty] | BA5
6 | BA5 | BLECF_ | 288 | 0x | 0120 |  |  | BA6 | 6 | [5 empty] | BA5
7 | BA1 | BLECF_ | 499 | 0x | 01F3 |  |  | BA1 | 1 | [5 empty] | BA1
8 | BA1 | BLECF_ | 425 | 0x | 01A9 |  |  | BA1 | 1 | [5 empty] | BA1
last entry 12/2/08

## CPS
| Tunnel - position | [11 empty] | Surface position
| HV box | CFC-Card |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  | VME-TC position
Pos. | position | No. | Hex. | Ch | Place | Rack | P. | Pos. | Crate | Slot |  | No. | Bldg. | Crate | No. | Rack | VME | Crate | Slot | Card | Dec.
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
last entry 11/2/08

## CMS
| Tunnel - position | [13 empty] | Surface position
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  | VME-TC position
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Crate | Slot |  | No. | Bldg. | Crate | No. | Rack | VME | Crate | Slot | Card | Dec.
1 |  | BLECF_ | 298 | 0x | 012A
2 |  | BLECF_ | 290 | 0x | 0122
3 |  | BLECF_ | 258 | 0x | 0102
4 |  | BLECF_ | 260 | 0x | 0104
5 |  | BLECF_ | 169 | 0x | 00A9
6 |  | BLECF_ | 171 | 0x | 00AB
7 |  | BLECF_ | 41 | 0x | 0029
8 |  | BLECF_ | 46 | 0x | 002E
9 |  | BLECF_ | 138 | 0x | 008A | [18 empty] | old version
10 |  | BLECF_ |  | 0x | 0000
11 |  | BLECF_ |  | 0x | 0000
12 |  | BLECF_ |  | 0x | 0000
13 |  | BLECF_ |  | 0x | 0000
last entry 11/2/08

## LHCb + Alice
| Tunnel - position | [13 empty] | Surface position
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  | VME-TC position
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Crate | Slot |  | No. | Bldg. | Crate | No. | Rack | VME | Crate | Slot | Card | Dec.
1 |  | BLECF_ | 37 | 0x | 0025
2 |  | BLECF_ | 53 | 0x | 0035
3 |  | BLECF_ | 400 | 0x | 0190
4 |  | BLECF_ | 402 | 0x | 0192
5 |  | BLECF_ | 110 | 0x | 006E
6
7
8
last entry 11/2/08

## Atlas
| Tunnel - position | [12 empty] | Surface position
| HV box | CFC-Card |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  | VME-TC position
Pos. | position | No. |  | Hex. | Ch | Place | Rack | P. | Pos. | Crate | Slot |  | No. | Bldg. | Crate | No. | Rack | VME | Crate | Slot | Card | Dec.
1 |  |  | 332
2 |  |  | 302
3
4
5
6
7
8
last entry 11/2/08

## Labo
| Tunnel - position | [11 empty] | Surface position
| HV box | CFC-Card |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  | VME-TC position
Pos. | position | No. | Hex. | Ch | Place | Rack | P. | Pos. | Crate | Slot |  | No. | Bldg. | Crate | No. | Rack | VME | Crate | Slot | Card | Dec.
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
last entry 11/2/08