## Installation overview
BLECF installation (tunnel electronic)
| ARC |  |  |  |  | straight section
| BLECF |  |  |  |  | BLECF |  |  |  |  | Source test planning
| Dismounted | Calibrated | Installed | Checked and working | Comments | Dismounted | Calibrated | Installed | Checked and working | Comments
1 LEFT | 22 | 22 | 22 | 22 |  | 19 | 19 | 19 | 19 |  | 4
1 RIGHT | 23 | 23 | 23 | 23 |  | 19 | 19 | 19 | 19 |  | 3
2 LEFT | 22 | 22 | 22 | 22 |  | 22 | 22 | 22 | 22 |  | 3
2 RIGHT | 23 | 23 | 23 | 23 |  | 21 | 21 | 21 | 21 |  | 8
3 LEFT | 22 | 22 | 22 | 22 |  | 18 | 18 | 18 | 18 |  | 8
3 RIGHT | 23 | 23 | 23 | 23 |  | 18 | 18 | 18 | 18 |  | 2
4 LEFT | 22 | 22 | 22 | 22 |  | 8 | 8 | 8 | 8 |  | 2
4 RIGHT | 23 | 23 | 23 | 23 |  | 8 | 8 | 8 | 8 |  | 6
5 LEFT | 22 | 22 | 22 | 22 |  | 19 | 19 | 19 | 19 |  | 6
5 RIGHT | 23 | 23 | 23 | 23 |  | 19 | 19 | 19 | 19 |  | 7
6 LEFT | 22 | 22 | 22 | 22 |  | 21 | 21 | 21 | 21 |  | 7
6 RIGHT | 23 | 23 | 23 | 23 |  | 21 | 21 | 21 | 21 |  | 1
7 LEFT | 22 | 22 | 22 | 22 |  | 28 | 28 | 28 | 28 |  | 1
7 RIGHT | 23 | 23 | 23 | 23 |  | 28 | 28 | 28 | 28 |  | 5
8 LEFT | 22 | 22 | 22 | 22 |  | 19 | 19 | 19 | 19 |  | 5
8 RIGHT | 23 | 23 | 23 | 23 |  | 21 | 21 | 20 | 21 | C11.R8 (BLECF_0300) not installed | 4
Total | 360 | 360 | 360 | 360 | 0 | 309 | 309 | 308 | 309
in % |  | 100 | 100 | 100 | 0 |  | 100 | 99.6764 | 100
Missing |  | 0 | 0 | 0 | 360 |  | 0 | 1 | 0
Sum BLECF complete installation |  |  |  |  | 669
Installed BLECF |  |  |  |  | 669 | 100
Missing BLECF complete installation |  |  |  |  | 0 | 0

## Point 1
| Tunnel - position | [15 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | NES18 | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 1024 | HV | Comments
Pos. | position |  |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBAP.A33L1 |  |  | BLECF_ | 285 | 0x | 011D | 6 | Arc | BYPLM.A33L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
2 | BJBAP.A32L1 |  |  | BLECF_ | 280 | 0x | 0118 | 6 | Arc | BYPLM.A32L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
3 | BJBAP.A31L1 |  |  | BLECF_ | 276 | 0x | 0114 | 6 | Arc | BYPLM.A31L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
4 | BJBAP.A30L1 |  |  | BLECF_ | 201 | 0x | 00C9 | 6 | Arc | BYPLM.A30L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
5 | BJBAP.A29L1 |  |  | BLECF_ | 141 | 0x | 008D | 6 | Arc | BYPLM.A29L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
6 | BJBAP.A28L1 |  |  | BLECF_ | 249 | 0x | 00F9 | 6 | Arc | BYPLM.A28L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
7 | BJBAP.A27L1 |  |  | BLECF_ | 241 | 0x | 00F1 | 6 | Arc | BYPLM.A27L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
8 | BJBAP.A26L1 |  |  | BLECF_ | 282 | 0x | 011A | 6 | Arc | BYPLM.A26L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
9 | BJBAP.A25L1 |  |  | BLECF_ | 738 | 0x | 02E2 | 6 | Arc | BYPLM.A25L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
10 | BJBAP.A24L1 |  |  | BLECF_ | 204 | 0x | 00CC | 6 | Arc | BYPLM.A24L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
11 | BJBAP.A23L1 |  |  | BLECF_ | 212 | 0x | 00D4 | 6 | Arc | BYPLM.A23L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
12 | BJBAP.A22L1 |  |  | BLECF_ | 284 | 0x | 011C | 6 | Arc | BYPLM.A22L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
13 | BJBAP.A21L1 |  |  | BLECF_ | 22 | 0x | 0016 | 6 | Arc | BYPLM.A21L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
14 | BJBAP.A20L1 |  |  | BLECF_ | 754 | 0x | 02F2 | 6 | Arc | BYPLM.A20L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
15 | BJBAP.A19L1 |  |  | BLECF_ | 35 | 0x | 0023 | 6 | Arc | BYPLM.A19L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
16 | BJBAP.A18L1 |  |  | BLECF_ | 400 | 0x | 0190 | 6 | Arc | BYPLM.A18L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
17 | BJBAP.A17L1 |  |  | BLECF_ | 36 | 0x | 0024 | 6 | Arc | BYPLM.A17L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
18 | BJBAP.A16L1 |  |  | BLECF_ | 64 | 0x | 0040 | 6 | Arc | BYPLM.A16L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
19 | BJBAP.A15L1 |  |  | BLECF_ | 59 | 0x | 003B | 6 | Arc | BYPLM.A15L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
20 | BJBAP.A14L1 |  |  | BLECF_ | 58 | 0x | 003A | 6 | Arc | BYPLM.A14L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
21 | BJBAP.A13L1 |  |  | BLECF_ | 77 | 0x | 004D | 8 | Arc | BYPLM.A13L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
22 | BJBAP.A12L1 |  |  | BLECF_ | 79 | 0x | 004F | 6 | Arc | BYPLM.A12L1 | 1 | L |  |  |  | 7-8 | SR1 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | doesn’t reset remotely
23 | BJBAP.B11L1 |  |  | BLECF_ | 667 | 0x | 029B | 8 | RR13 | BY02 | 1 | CL | 2 | 9 | B | 13-14 | SR1 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
24 | BJBAP.A11L1 |  |  | BLECF_ | 594 | 0x | 0252 | 8 | RR13 | BY02 | 1 | CL | 2 | 8 | B | 11-12 | SR1 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
25 | BJBAP.B10L1 |  |  | BLECF_ | 409 | 0x | 0199 | 6 | RR13 | BY02 | 1 | CL | 2 | 7 | B | 9-10 | SR1 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
26 | BJBAP.A10L1 |  |  | BLECF_ | 550 | 0x | 0226 | 2 | RR13 | BY02 | 1 | CL | 2 | 6 | B | 7-8 | SR1 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
27 | BJBAP.A9L1 |  |  | BLECF_ | 518 | 0x | 0206 | 6 | RR13 | BY02 | 1 | CL | 2 | 5 | B | 5-6 | SR1 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
28 | BJBAP.B8L1 |  |  | BLECF_ | 457 | 0x | 01C9 | 6 | RR13 | BY02 | 1 | CL | 2 | 4 | B | 3-4 | SR1 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
29 | BJBAP.A8L1 |  |  | BLECF_ | 588 | 0x | 024C | 6 | RR13 | BY02 | 1 | CL | 2 | 3 | B | 1-2 | SR1 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
30 | BJBAP.B7L1 |  |  | BLECF_ | 322 | 0x | 0142 | 7 | RR13 | BY02 | 1 | CL | 2 | 2 | A | 23-24 | SR1 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
31 | BJBAP.A7L1 |  |  | BLECF_ | 194 | 0x | 00C2 | 1 | RR13 | BY02 | 1 | CL | 2 | 1 | A | 21-22 | SR1 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
32 | BJBAP.A6L1 |  | NES18 | BLECF_ | 89 | 0x | 0059 | 6 | RR13 | BY02 | 1 | CL | 1 | 10 | A | 19-20 | SR1 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
33 | BJBAP.B5L1 |  |  | BLECF_ | 530 | 0x | 0212 | 7 | RR13 | BY02 | 1 | CL | 1 | 9 | A | 17-18 | SR1 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
34 | BJSAP.A5L1 |  |  | BLECF_ | 597 | 0x | 0255 | 1 | RR13 | BY02 | 1 | CL | 1 | 8 | A | 15-16 | SR1 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
35 | BJBAP.C4L1 |  |  | BLECF_ | 398 | 0x | 018E | 6 | RR13 | BY02 | 1 | CL | 1 | 7 | A | 13-14 | SR1 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
36 | BJSAP.B4L1 |  |  | BLECF_ | 374 | 0x | 0176 | 4 | RR13 | BY02 | 1 | CL | 1 | 6 | A | 11-12 | SR1 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
37 | BJBAP.A4L1 |  |  | BLECF_ | 401 | 0x | 0191 | 4 | RR13 | BY02 | 1 | CL | 1 | 5 | A | 9-10 | SR1 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
38 | BJBAP.A3L1 |  | NES18 | BLECF_ | 593 | 0x | 0251 | 5 | RR13 | BY02 | 1 | CL | 1 | 4 | A | 7-8 | SR1 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
39 | BJBAP.A2L1 |  | NES18 | BLECF_ | 592 | 0x | 0250 | 8 | RR13 | BY02 | 1 | CL | 1 | 3 | A | 5-6 | SR1 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
40 | BJBAP.B1L1 |  | NES18 | BLECF_ | 635 | 0x | 027B | 5 | RR13 | BY02 | 1 | CL | 1 | 2 | A | 3-4 | SR1 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
41 | BJSAP.A1L1 | yes |  | BLECF_ | 251 | 0x | 00FB | 1 | RR13 | BY02 | 1 | CL | 1 | 1 | A | 1-2 | SR1 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
42 | BJSAP.A1R1 | yes |  | BLECF_ | 377 | 0x | 0179 | 1 | RR17 | BY02 | 1 | CR | 1 | 1 | A | 1-2 | SR1 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
43 | BJBAP.B1R1 |  |  | BLECF_ | 331 | 0x | 014B | 5 | RR17 | BY02 | 1 | CR | 1 | 2 | A | 3-4 | SR1 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
44 | BJBAP.A2R1 |  | NES18 | BLECF_ | 501 | 0x | 01F5 | 8 | RR17 | BY02 | 1 | CR | 1 | 3 | A | 5-6 | SR1 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
45 | BJBAP.A3R1 |  | NES18 | BLECF_ | 636 | 0x | 027C | 5 | RR17 | BY02 | 1 | CR | 1 | 4 | A | 7-8 | SR1 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
46 | BJBAP.A4R1 |  |  | BLECF_ | 676 | 0x | 02A4 | 4 | RR17 | BY02 | 1 | CR | 1 | 5 | A | 9-10 | SR1 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
47 | BJSAP.B4R1 |  |  | BLECF_ | 653 | 0x | 028D | 4 | RR17 | BY02 | 1 | CR | 1 | 6 | A | 11-12 | SR1 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
48 | BJBAP.C4R1 |  |  | BLECF_ | 376 | 0x | 0178 | 6 | RR17 | BY02 | 1 | CR | 1 | 7 | A | 13-14 | SR1 | BY05 | C | 11 | 1-2 | BY01 | R | 3 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
49 | BJSAP.A5R1 |  |  | BLECF_ | 437 | 0x | 01B5 | 1 | RR17 | BY02 | 1 | CR | 1 | 8 | A | 15-16 | SR1 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
50 | BJBAP.B5R1 |  | NES18 | BLECF_ | 442 | 0x | 01BA | 7 | RR17 | BY02 | 1 | CR | 1 | 9 | A | 17-18 | SR1 | BY05 | C | 11 | 5-6 | BY01 | R | 3 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
51 | BJBAP.A6R1 |  | NES18 | BLECF_ | 355 | 0x | 0163 | 6 | RR17 | BY02 | 1 | CR | 1 | 10 | A | 19-20 | SR1 | BY05 | C | 11 | 7-8 | BY01 | R | 3 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
52 | BJBAP.A7R1 |  |  | BLECF_ | 323 | 0x | 0143 | 1 | RR17 | BY02 | 1 | CR | 2 | 1 | A | 21-22 | SR1 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
53 | BJBAP.B7R1 |  |  | BLECF_ | 512 | 0x | 0200 | 7 | RR17 | BY02 | 1 | CR | 2 | 2 | A | 23-24 | SR1 | BY05 | C | 11 | 11-12 | BY01 | R | 3 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
54 | BJBAP.A8R1 |  |  | BLECF_ | 325 | 0x | 0145 | 6 | RR17 | BY02 | 1 | CR | 2 | 3 | B | 1-2 | SR1 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
55 | BJBAP.B8R1 |  |  | BLECF_ | 647 | 0x | 0287 | 6 | RR17 | BY02 | 1 | CR | 2 | 4 | B | 3-4 | SR1 | BY05 | C | 10 | 3-4 | BY01 | R | 3 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
56 | BJBAP.A9R1 |  | NES18 | BLECF_ | 85 | 0x | 0055 | 6 | RR17 | BY02 | 1 | CR | 2 | 5 | B | 5-6 | SR1 | BY05 | C | 10 | 5-6 | BY01 | R | 3 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
57 | BJBAP.A10R1 |  |  | BLECF_ | 338 | 0x | 0152 | 2 | RR17 | BY02 | 1 | CR | 2 | 6 | B | 7-8 | SR1 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
58 | BJBAP.B10R1 |  |  | BLECF_ | 567 | 0x | 0237 | 6 | RR17 | BY02 | 1 | CR | 2 | 7 | B | 9-10 | SR1 | BY05 | C | 10 | 9-10 | BY01 | R | 3 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
59 | BJBAP.A11R1 |  |  | BLECF_ | 349 | 0x | 015D | 8 | RR17 | BY02 | 1 | CR | 2 | 8 | B | 11-12 | SR1 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
60 | BJBAP.B11R1 |  |  | BLECF_ | 327 | 0x | 0147 | 8 | RR17 | BY03 | 1 | CR | 2 | 9 | B | 13-14 | SR2 | BY06 | C | 9 | 1-2 | BY01 | R | 3 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
61 | BJBAP.A12R1 |  |  | BLECF_ | 539 | 0x | 021B | 6 | Arc | BYPLM.A12R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 8 | 7-8 | BY01 | R | 3 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
62 | BJBAP.A13R1 |  |  | BLECF_ | 361 | 0x | 0169 | 8 | Arc | BYPLM.A13R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 9 | 7-8 | BY01 | R | 3 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
63 | BJBAP.A14R1 |  |  | BLECF_ | 570 | 0x | 023A | 6 | Arc | BYPLM.A14R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 10 | 7-8 | BY01 | R | 3 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
64 | BJBAP.A15R1 |  |  | BLECF_ | 306 | 0x | 0132 | 6 | Arc | BYPLM.A15R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 11 | 7-8 | BY01 | R | 3 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
65 | BJBAP.A16R1 |  |  | BLECF_ | 311 | 0x | 0137 | 6 | Arc | BYPLM.A16R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | F | 12 | 7-8 | BY01 | R | 3 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
66 | BJBAP.A17R1 |  |  | BLECF_ | 303 | 0x | 012F | 6 | Arc | BYPLM.A17R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 7 | 7-8 | BY01 | R | 3 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
67 | BJBAP.A18R1 |  |  | BLECF_ | 515 | 0x | 0203 | 6 | Arc | BYPLM.A18R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 8 | 7-8 | BY01 | R | 3 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
68 | BJBAP.A19R1 |  |  | BLECF_ | 481 | 0x | 01E1 | 6 | Arc | BYPLM.A19R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 9 | 7-8 | BY01 | R | 3 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
69 | BJBAP.A20R1 |  |  | BLECF_ | 81 | 0x | 0051 | 6 | Arc | BYPLM.A20R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 10 | 7-8 | BY01 | R | 3 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
70 | BJBAP.A21R1 |  |  | BLECF_ | 522 | 0x | 020A | 6 | Arc | BYPLM.A21R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 11 | 7-8 | BY01 | R | 3 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
71 | BJBAP.A22R1 |  |  | BLECF_ | 373 | 0x | 0175 | 6 | Arc | BYPLM.A22R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | G | 12 | 7-8 | BY01 | R | 3 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
72 | BJBAP.A23R1 |  |  | BLECF_ | 423 | 0x | 01A7 | 6 | Arc | BYPLM.A23R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 7 | 7-8 | BY01 | R | 3 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
73 | BJBAP.A24R1 |  |  | BLECF_ | 352 | 0x | 0160 | 6 | Arc | BYPLM.A24R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 8 | 7-8 | BY01 | R | 3 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
74 | BJBAP.A25R1 |  |  | BLECF_ | 500 | 0x | 01F4 | 6 | Arc | BYPLM.A25R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 9 | 7-8 | BY01 | R | 3 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
75 | BJBAP.A26R1 |  |  | BLECF_ | 479 | 0x | 01DF | 6 | Arc | BYPLM.A26R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 10 | 7-8 | BY01 | R | 3 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
76 | BJBAP.A27R1 |  |  | BLECF_ | 478 | 0x | 01DE | 6 | Arc | BYPLM.A27R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 11 | 7-8 | BY01 | R | 3 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
77 | BJBAP.A28R1 |  |  | BLECF_ | 317 | 0x | 013D | 6 | Arc | BYPLM.A28R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | H | 12 | 7-8 | BY01 | R | 3 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
78 | BJBAP.A29R1 |  |  | BLECF_ | 310 | 0x | 0136 | 6 | Arc | BYPLM.A29R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 7 | 7-8 | BY01 | R | 3 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
79 | BJBAP.A30R1 |  |  | BLECF_ | 329 | 0x | 0149 | 6 | Arc | BYPLM.A30R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 8 | 7-8 | BY01 | R | 3 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
80 | BJBAP.A31R1 |  |  | BLECF_ | 495 | 0x | 01EF | 6 | Arc | BYPLM.A31R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 9 | 7-8 | BY01 | R | 3 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
81 | BJBAP.A32R1 |  |  | BLECF_ | 384 | 0x | 0180 | 6 | Arc | BYPLM.A32R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 10 | 7-8 | BY01 | R | 3 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
82 | BJBAP.A33R1 |  |  | BLECF_ | 417 | 0x | 01A1 | 6 | Arc | BYPLM.A33R1 | 1 | R |  |  |  | 7-8 | SR1 | BY05 | I | 11 | 7-8 | BY01 | R | 3 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
83 | BJBAP.A34R1 |  |  | BLECF_ | 321 | 0x | 0141 | 6 | Arc | BYPLM.A34R1 | 1 | R |  |  |  | 7-9 | SR1 | BY05 | I | 12 | 7-8 | BY01 | R | 3 | 20 | 16 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
Door
[28 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[28 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[8 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[8 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[8 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[28 empty] | O.K. | No problems observed
Note on V14: effinger: 13.08.08 need to check if the fibre are changed to 5-6 --> 18.8.08 was connected to 5&6 reconnected to 7&8
Note on V79: effinger: 13.08.08 need to check if the fibre are changed to 5-6 --> checked 18.8.08 connected to 7-8

## Point 2
| Tunnel - position | [15 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | NES18 | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 1024 | HV | Comments
Pos. | position |  |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBAP.A33L2 |  |  | BLECF_ | 399 | 0x | 018F | 6 | Arc | BYPLM.A33L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
2 | BJBAP.A32L2 |  |  | BLECF_ | 348 | 0x | 015C | 6 | Arc | BYPLM.A32L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
3 | BJBAP.A31L2 |  |  | BLECF_ | 497 | 0x | 01F1 | 6 | Arc | BYPLM.A31L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
4 | BJBAP.A30L2 |  |  | BLECF_ | 432 | 0x | 01B0 | 6 | Arc | BYPLM.A30L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
5 | BJBAP.A29L2 |  |  | BLECF_ | 520 | 0x | 0208 | 6 | Arc | BYPLM.A29L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
6 | BJBAP.A28L2 |  |  | BLECF_ | 439 | 0x | 01B7 | 6 | Arc | BYPLM.A28L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
7 | BJBAP.A27L2 |  |  | BLECF_ | 342 | 0x | 0156 | 6 | Arc | BYPLM.A27L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
8 | BJBAP.A26L2 |  |  | BLECF_ | 460 | 0x | 01CC | 6 | Arc | BYPLM.A26L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
9 | BJBAP.A25L2 |  |  | BLECF_ | 388 | 0x | 0184 | 6 | Arc | BYPLM.A25L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
10 | BJBAP.A24L2 |  |  | BLECF_ | 440 | 0x | 01B8 | 6 | Arc | BYPLM.A24L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
11 | BJBAP.A23L2 |  |  | BLECF_ | 468 | 0x | 01D4 | 6 | Arc | BYPLM.A23L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
12 | BJBAP.A22L2 |  |  | BLECF_ | 456 | 0x | 01C8 | 6 | Arc | BYPLM.A22L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
13 | BJBAP.A21L2 |  |  | BLECF_ | 111 | 0x | 006F | 6 | Arc | BYPLM.A21L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
14 | BJBAP.A20L2 |  |  | BLECF_ | 11 | 0x | 000B | 6 | Arc | BYPLM.A20L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
15 | BJBAP.A19L2 |  |  | BLECF_ | 715 | 0x | 02CB | 6 | Arc | BYPLM.A19L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
16 | BJBAP.A18L2 |  |  | BLECF_ | 366 | 0x | 016E | 6 | Arc | BYPLM.A18L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
17 | BJBAP.A17L2 |  |  | BLECF_ | 427 | 0x | 01AB | 6 | Arc | BYPLM.A17L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
18 | BJBAP.A16L2 |  |  | BLECF_ | 385 | 0x | 0181 | 6 | Arc | BYPLM.A16L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
19 | BJBAP.A15L2 |  |  | BLECF_ | 444 | 0x | 01BC | 6 | Arc | BYPLM.A15L2 | 2 | L |  |  |  | 5-6 | SR2 | BY05 | F | 4 | 5-6 | BY02 | L | 1 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
20 | BJBAP.A14L2 |  |  | BLECF_ | 380 | 0x | 017C | 6 | Arc | BYPLM.A14L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
21 | BJBAP.A13L2 |  |  | BLECF_ | 347 | 0x | 015B | 6 | Arc | BYPLM.A13L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
22 | BJBAP.A12L2 |  |  | BLECF_ | 200 | 0x | 00C8 | 8 | Arc | BYPLM.A12L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
23 | BJBAP.C11L2 |  |  | BLECF_ | 583 | 0x | 0247 | 6 | UA23 | BY02 | 2 | CL | 3 | 2 | C | 19-20 | SR2 | BY05 | C | 4 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
24 | BJSAP.B11L2 | yes |  | BLECF_ | 343 | 0x | 0157 | 6 | UA23 | BY02 | 2 | CL | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
25 | BJBAP.A11L2 |  |  | BLECF_ | 378 | 0x | 017A | 6 | UA23 | BY02 | 2 | CL | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
26 | BJBAP.B10L2 |  |  | BLECF_ | 98 | 0x | 0062 | 6 | UA23 | BY02 | 2 | CL | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
27 | BJBAP.A10L2 |  |  | BLECF_ | 363 | 0x | 016B | 7 | UA23 | BY02 | 2 | CL | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 9-16 | ok | ok | bad | bad | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok |  | ok
28 | BJBAP.A9L2 |  |  | BLECF_ | 517 | 0x | 0205 | 6 | UA23 | BY02 | 2 | CL | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
29 | BJBAP.C8L2 | yes |  | BLECF_ | 357 | 0x | 0165 | 6 | UA23 | BY02 | 2 | CL | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 6 | 3 | 1-8 | ok | ok | bad | bad | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok |  | ok
| BJBAP.B8L2 |  |  | BLECF_ | 333 | 0x | 014D | 6 | UA23 | BY02 | 2 | CL | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 9 | 6 | 1-8
30 | BJSAP.B8L2 | yes |  | BLECF_ | 333 | 0x | 014D | 6 | UA23 | BY02 | 2 | CL | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 3 | 5-6 | BY01 | I | 4 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
31 | BJBAP.A8L2 | yes |  | BLECF_ | 491 | 0x | 01EB | 6 | UA23 | BY02 | 2 | CL | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 9-16 | ok | ok | bad | bad | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok |  | ok
32 | BJBAP.A7L2 | yes |  | BLECF_ | 362 | 0x | 016A | 6 | UA23 | BY02 | 2 | CL | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
33 | BJBAP.C6L2 | yes | NES18 | BLECF_ | 189 | 0x | 00BD | 6 | UA23 | BY02 | 2 | CL | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
34 | BJSAP.B6L2 | yes |  | BLECF_ | 344 | 0x | 0158 | 6 | UA23 | BY02 | 2 | CL | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
| BJBAP.A6L2 |  |  | BLECF_ | 474 | 0x | 01DA | 6 | UA23 | BY02 | 2 | CL | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16
35 | BJBAP.A6L2 | yes |  | BLECF_ | 474 | 0x | 01DA | 6 | UA23 | BY02 | 2 | CL | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 2 | 7-8 | BY01 | I | 4 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
36 | BJBAP.A5L2 |  |  | BLECF_ | 328 | 0x | 0148 | 6 | UA23 | BY02 | 2 | CL | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
37 | BJBAP.D4L2 |  |  | BLECF_ | 371 | 0x | 0173 | 8 | UA23 | BY02 | 2 | CL | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
38 | BJSAP.C4L2 | yes |  | BLECF_ | 181 | 0x | 00B5 | 1 | UA23 | BY02 | 2 | CL | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
| BJBAP.B4L2 |  |  | BLECF_ | 320 | 0x | 0140 | 4 | UA23 | BY02 | 2 | CL | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16
39 | BJSAP.B4L2 | yes |  | BLECF_ | 320 | 0x | 0140 | 4 | UA23 | BY02 | 2 | CL | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 1 | 11-12 | BY01 | I | 4 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
40 | BJBAP.A4L2 | yes | NES18 | BLECF_ | 516 | 0x | 0204 | 7 | UA23 | BY02 | 2 | CL | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
41 | BJBAP.A3L2 |  |  | BLECF_ | 324 | 0x | 0144 | 4 | UA23 | BY02 | 2 | CL | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
42 | BJBAP.A2L2 |  | NES18 | BLECF_ | 188 | 0x | 00BC | 8 | UA23 | BY02 | 2 | CL | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
43 | BJBAP.B1L2 |  |  | BLECF_ | 510 | 0x | 01FE | 5 | UA23 | BY02 | 2 | CL | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
44 | BJSAP.A1L2 | yes |  | BLECF_ | 476 | 0x | 01DC | 1 | UA23 | BY02 | 2 | CL | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
45 | BJSAP.A1R2 | yes |  | BLECF_ | 706 | 0x | 02C2 | 1 | UA27 | BY03 | 2 | CR | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
46 | BJBAP.B1R2 |  | NES18 | BLECF_ | 606 | 0x | 025E | 5 | UA27 | BY03 | 2 | CR | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
47 | BJBAP.A2R2 |  | NES18 | BLECF_ | 675 | 0x | 02A3 | 8 | UA27 | BY03 | 2 | CR | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
48 | BJBAP.A3R2 |  |  | BLECF_ | 641 | 0x | 0281 | 4 | UA27 | BY03 | 2 | CR | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
49 | BJBAP.A4R2 |  |  | BLECF_ | 719 | 0x | 02CF | 5 | UA27 | BY03 | 2 | CR | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
50 | BJSAP.B4R2 | yes |  | BLECF_ | 258 | 0x | 0102 | 2 | UA27 | BY03 | 2 | CR | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
51 | BJSAP.C4R2 | yes |  | BLECF_ | 652 | 0x | 028C | 1 | UA27 | BY03 | 2 | CR | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
52 | BJBAP.D4R2 |  |  | BLECF_ | 665 | 0x | 0299 | 8 | UA27 | BY03 | 2 | CR | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 11 | 3-4 | BY01 | R | 3 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
53 | BJBAP.A5R2 |  | NES18 | BLECF_ | 740 | 0x | 02E4 | 6 | UA27 | BY03 | 2 | CR | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 11 | 5-6 | BY01 | R | 3 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
54 | BJSAP.A6R2 | yes |  | BLECF_ | 666 | 0x | 029A | 1 | UA27 | BY03 | 2 | CR | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | erre 37
55 | BJBAP.B6R2 | yes | NES18 | BLECF_ | 722 | 0x | 02D2 | 7 | UA27 | BY03 | 2 | CR | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 11 | 9-10 | BY01 | R | 3 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
56 | BJBAP.A7R2 |  |  | BLECF_ | 658 | 0x | 0292 | 6 | UA27 | BY03 | 2 | CR | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 11 | 11-12 | BY01 | R | 3 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
57 | BJBAP.A8R2 |  |  | BLECF_ | 626 | 0x | 0272 | 6 | UA27 | BY03 | 2 | CR | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
58 | BJSAP.B8R2 | yes |  | BLECF_ | 406 | 0x | 0196 | 6 | UA27 | BY03 | 2 | CR | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
59 | BJBAP.C8R2 |  |  | BLECF_ | 4 | 0x | 0004 | 6 | UA27 | BY03 | 2 | CR | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 10 | 5-6 | BY01 | R | 3 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
60 | BJBAP.A9R2 |  |  | BLECF_ | 279 | 0x | 0117 | 6 | UA27 | BY03 | 2 | CR | 2 | 6 | C | 7 | SR2 | BY05 | C | 10 | 7 | BY01 | R | 3 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
61 | BJBAP.A10R2 |  |  | BLECF_ | 97 | 0x | 0061 | 6 | UA27 | BY03 | 2 | CR | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
62 | BJBAP.B10R2 |  |  | BLECF_ | 721 | 0x | 02D1 | 6 | UA27 | BY03 | 2 | CR | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 10 | 11-12 | BY01 | R | 3 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
63 | BJBAP.A11R2 |  | NES18 | BLECF_ | 128 | 0x | 0080 | 6 | UA27 | BY03 | 2 | CR | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
64 | BJSAP.B11R2 | yes |  | BLECF_ | 627 | 0x | 0273 | 6 | UA27 | BY03 | 2 | CR | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
65 | BJBAP.C11R2 |  |  | BLECF_ | 689 | 0x | 02B1 | 6 | UA27 | BY03 | 2 | CR | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 9 | 5-6 | BY01 | R | 3 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
66 | BJBAP.A12R2 |  |  | BLECF_ | 681 | 0x | 02A9 | 8 | Arc | BYPLM.A12R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 8 | 7-8 | BY01 | R | 3 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
67 | BJBAP.A13R2 |  |  | BLECF_ | 687 | 0x | 02AF | 6 | Arc | BYPLM.A13R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 9 | 7-8 | BY01 | R | 3 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
68 | BJBAP.A14R2 |  |  | BLECF_ | 684 | 0x | 02AC | 6 | Arc | BYPLM.A14R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 10 | 7-8 | BY01 | R | 3 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
69 | BJBAP.A15R2 |  |  | BLECF_ | 484 | 0x | 01E4 | 6 | Arc | BYPLM.A15R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 11 | 7-8 | BY01 | R | 3 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
70 | BJBAP.A16R2 |  |  | BLECF_ | 686 | 0x | 02AE | 6 | Arc | BYPLM.A16R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 12 | 7-8 | BY01 | R | 3 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
71 | BJBAP.A17R2 |  |  | BLECF_ | 143 | 0x | 008F | 6 | Arc | BYPLM.A17R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 7 | 7-8 | BY01 | R | 3 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
72 | BJBAP.A18R2 |  |  | BLECF_ | 463 | 0x | 01CF | 6 | Arc | BYPLM.A18R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 8 | 7-8 | BY01 | R | 3 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
73 | BJBAP.A19R2 |  |  | BLECF_ | 379 | 0x | 017B | 6 | Arc | BYPLM.A19R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 9 | 7-8 | BY01 | R | 3 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
74 | BJBAP.A20R2 |  |  | BLECF_ | 382 | 0x | 017E | 6 | Arc | BYPLM.A20R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 10 | 7-8 | BY01 | R | 3 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
75 | BJBAP.A21R2 |  |  | BLECF_ | 451 | 0x | 01C3 | 6 | Arc | BYPLM.A21R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 11 | 7-8 | BY01 | R | 3 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
76 | BJBAP.A22R2 |  |  | BLECF_ | 404 | 0x | 0194 | 6 | Arc | BYPLM.A22R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 12 | 7-8 | BY01 | R | 3 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
77 | BJBAP.A23R2 |  |  | BLECF_ | 55 | 0x | 0037 | 6 | Arc | BYPLM.A23R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 7 | 7-8 | BY01 | R | 3 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
78 | BJBAP.A24R2 |  |  | BLECF_ | 418 | 0x | 01A2 | 6 | Arc | BYPLM.A24R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 8 | 7-8 | BY01 | R | 3 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
79 | BJBAP.A25R2 |  |  | BLECF_ | 390 | 0x | 0186 | 6 | Arc | BYPLM.A25R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 9 | 7-8 | BY01 | R | 3 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
80 | BJBAP.A26R2 |  |  | BLECF_ | 214 | 0x | 00D6 | 6 | Arc | BYPLM.A26R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 10 | 7-8 | BY01 | R | 3 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
81 | BJBAP.A27R2 |  |  | BLECF_ | 694 | 0x | 02B6 | 6 | Arc | BYPLM.A27R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 11 | 7-8 | BY01 | R | 3 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
82 | BJBAP.A28R2 |  |  | BLECF_ | 685 | 0x | 02AD | 6 | Arc | BYPLM.A28R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 12 | 7-8 | BY01 | R | 3 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
83 | BJBAP.A29R2 |  |  | BLECF_ | 725 | 0x | 02D5 | 6 | Arc | BYPLM.A29R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 7 | 7-8 | BY01 | R | 3 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
84 | BJBAP.A30R2 |  |  | BLECF_ | 195 | 0x | 00C3 | 6 | Arc | BYPLM.A30R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 8 | 7-8 | BY01 | R | 3 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
85 | BJBAP.A31R2 |  |  | BLECF_ | 690 | 0x | 02B2 | 6 | Arc | BYPLM.A31R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 9 | 7-8 | BY01 | R | 3 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
86 | BJBAP.A32R2 |  |  | BLECF_ | 82 | 0x | 0052 | 6 | Arc | BYPLM.A32R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 10 | 7-8 | BY01 | R | 3 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
87 | BJBAP.A33R2 |  |  | BLECF_ | 751 | 0x | 02EF | 6 | Arc | BYPLM.A33R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 11 | 7-8 | BY01 | R | 3 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
88 | BJBAP.A34R2 |  |  | BLECF_ | 724 | 0x | 02D4 | 6 | Arc | BYPLM.A34R2 | 2 | R |  |  |  | 7-9 | SR2 | BY05 | I | 12 | 7-8 | BY01 | R | 3 | 20 | 16 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
Door
[28 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[28 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[8 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[8 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[8 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[28 empty] | O.K. | No problems observed
Note on V23: effinger: Fibre 7 & 8 broken, spare fibre used! Still the case after LS1!!!!
Note on Q69: Ewald Effinger: Second fibre: C 19
Note on V69: Ewald Effinger: Second fiber : C 9 - 7

## Point 3
| Tunnel - position | [15 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | NES18 | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 1024 | HV | Comments
Pos. | position |  |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBAP.A33L3 |  |  | BLECF_ | 619 | 0x | 026B | 6 | Arc | BYPLM.A33L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
2 | BJBAP.A32L3 |  |  | BLECF_ | 601 | 0x | 0259 | 6 | Arc | BYPLM.A32L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | bad | ok | doesn-t reset remotely
3 | BJBAP.A31L3 |  |  | BLECF_ | 727 | 0x | 02D7 | 6 | Arc | BYPLM.A31L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
4 | BJBAP.A30L3 |  |  | BLECF_ | 705 | 0x | 02C1 | 6 | Arc | BYPLM.A30L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
5 | BJBAP.A29L3 |  |  | BLECF_ | 646 | 0x | 0286 | 6 | Arc | BYPLM.A29L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
6 | BJBAP.A28L3 |  |  | BLECF_ | 702 | 0x | 02BE | 6 | Arc | BYPLM.A28L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
7 | BJBAP.A27L3 |  |  | BLECF_ | 607 | 0x | 025F | 6 | Arc | BYPLM.A27L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
8 | BJBAP.A26L3 |  |  | BLECF_ | 623 | 0x | 026F | 6 | Arc | BYPLM.A26L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
9 | BJBAP.A25L3 |  |  | BLECF_ | 717 | 0x | 02CD | 6 | Arc | BYPLM.A25L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
10 | BJBAP.A24L3 |  |  | BLECF_ | 319 | 0x | 013F | 6 | Arc | BYPLM.A24L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
11 | BJBAP.A23L3 |  |  | BLECF_ | 745 | 0x | 02E9 | 6 | Arc | BYPLM.A23L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
12 | BJBAP.A22L3 |  |  | BLECF_ | 645 | 0x | 0285 | 6 | Arc | BYPLM.A22L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
13 | BJBAP.A21L3 |  |  | BLECF_ | 741 | 0x | 02E5 | 6 | Arc | BYPLM.A21L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
14 | BJBAP.A20L3 |  |  | BLECF_ | 614 | 0x | 0266 | 6 | Arc | BYPLM.A20L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
15 | BJBAP.A19L3 |  |  | BLECF_ | 620 | 0x | 026C | 6 | Arc | BYPLM.A19L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
16 | BJBAP.A18L3 |  |  | BLECF_ | 720 | 0x | 02D0 | 6 | Arc | BYPLM.A18L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
17 | BJBAP.A17L3 |  |  | BLECF_ | 668 | 0x | 029C | 6 | Arc | BYPLM.A17L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
18 | BJBAP.A16L3 |  |  | BLECF_ | 604 | 0x | 025C | 6 | Arc | BYPLM.A16L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
19 | BJBAP.A15L3 |  |  | BLECF_ | 622 | 0x | 026E | 6 | Arc | BYPLM.A15L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
20 | BJBAP.A14L3 |  |  | BLECF_ | 305 | 0x | 0131 | 6 | Arc | BYPLM.A14L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
21 | BJBAP.A13L3 |  |  | BLECF_ | 600 | 0x | 0258 | 10 | Arc | BYPLM.A13L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
22 | BJBAP.A12L3 |  |  | BLECF_ | 703 | 0x | 02BF | 10 | Arc | BYPLM.A12L3 | 3 | L |  |  |  | 7-8 | SR3 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
23 | BJBAP.B11L3 |  |  | BLECF_ | 657 | 0x | 0291 | 7 | UJ33 | BY02 | 3 | L | 2 | 8 | C | 11-12 | SR3 | BY05 | C | 3 | 11-12 | BY02 | L | 1 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
24 | BJBAP.A11L3 |  |  | BLECF_ | 654 | 0x | 028E | 12 | UJ33 | BY02 | 3 | L | 2 | 7 | C | 9-10 | SR3 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
25 | BJBAP.A10L3 |  |  | BLECF_ | 574 | 0x | 023E | 11 | UJ33 | BY02 | 3 | L | 2 | 6 | C | 7-8 | SR3 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
26 | BJBAP.B9L3 |  |  | BLECF_ | 674 | 0x | 02A2 | 11 | UJ33 | BY02 | 3 | L | 2 | 5 | C | 5-6 | SR3 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
27 | BJBAP.A9L3 |  |  | BLECF_ | 609 | 0x | 0261 | 10 | UJ33 | BY02 | 3 | L | 2 | 4 | C | 3-4 | SR3 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
28 | BJBAP.B8L3 |  |  | BLECF_ | 571 | 0x | 023B | 6 | UJ33 | BY02 | 3 | L | 2 | 3 | C | 1-2 | SR3 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
29 | BJBAP.A8L3 |  |  | BLECF_ | 103 | 0x | 0067 | 6 | UJ33 | BY02 | 3 | L | 2 | 2 | B | 23-24 | SR3 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
30 | BJBAP.B7L3 |  |  | BLECF_ | 576 | 0x | 0240 | 7 | UJ33 | BY02 | 3 | L | 2 | 1 | B | 21-22 | SR3 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
31 | BJSAP.A7L3 |  |  | BLECF_ | 637 | 0x | 027D | 1 | UJ33 | BY02 | 3 | L | 1 | 10 | B | 19-20 | SR3 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
32 | BJBAP.C6L3 |  |  | BLECF_ | 211 | 0x | 00D3 | 6 | UJ33 | BY02 | 3 | L | 1 | 9 | B | 17-18 | SR3 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
33 | BJSAP.B6L3 |  |  | BLECF_ | 629 | 0x | 0275 | 4 | UJ33 | BY02 | 3 | L | 1 | 8 | B | 15-16 | SR3 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
34 | BJBAP.A6L3 |  |  | BLECF_ | 536 | 0x | 0218 | 4 | UJ33 | BY02 | 3 | L | 1 | 7 | B | 13-14 | SR3 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
35 | BJSAP.D5L3 |  |  | BLECF_ | 175 | 0x | 00AF | 2 | UJ33 | BY02 | 3 | L | 1 | 6 | B | 11-12 | SR3 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
36 | BJBAP.C5L3 |  |  | BLECF_ | 577 | 0x | 0241 | 8 | UJ33 | BY02 | 3 | L | 1 | 5 | B | 9-10 | SR3 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
37 | BJSAP.B5L3 |  |  | BLECF_ | 575 | 0x | 023F | 6 | UJ33 | BY02 | 3 | L | 1 | 4 | B | 7-8 | SR3 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
38 | BJBAP.A5L3 |  |  | BLECF_ | 602 | 0x | 025A | 6 | UJ33 | BY02 | 3 | L | 1 | 3 | B | 5-6 | SR3 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
39 | BJSAP.B4L3 |  |  | BLECF_ | 711 | 0x | 02C7 | 2 | UJ33 | BY02 | 3 | L | 1 | 2 | B | 3-4 | SR3 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
40 | BJBAP.A4L3 |  |  | BLECF_ | 729 | 0x | 02D9 | 8 | UJ33 | BY02 | 3 | L | 1 | 1 | B | 1-2 | SR3 | BY05 | C | 1 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
41 | BJBAP.A4R3 |  |  | BLECF_ | 543 | 0x | 021F | 8 | UJ33 | BY03 | 3 | R | 1 | 1 | C | 13-14 | SR3 | BY05 | C | 4 | 1-2 | BY01 | R | 3 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
42 | BJSAP.B4R3 |  |  | BLECF_ | 677 | 0x | 02A5 | 2 | UJ33 | BY03 | 3 | R | 1 | 2 | C | 15-16 | SR3 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
43 | BJBAP.A5R3 |  |  | BLECF_ | 182 | 0x | 00B6 | 6 | UJ33 | BY03 | 3 | R | 1 | 3 | C | 17-18 | SR3 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
44 | BJSAP.B5R3 |  |  | BLECF_ | 326 | 0x | 0146 | 6 | UJ33 | BY03 | 3 | R | 1 | 4 | C | 19-20 | SR3 | BY05 | C | 4 | 7-8 | BY02 | C | 2 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
45 | BJBAP.C5R3 |  |  | BLECF_ | 487 | 0x | 01E7 | 8 | UJ33 | BY03 | 3 | R | 1 | 5 | C | 21-22 | SR3 | BY05 | C | 4 | 9-10 | BY01 | R | 3 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
46 | BJSAP.D5R3 |  |  | BLECF_ | 393 | 0x | 0189 | 2 | UJ33 | BY03 | 3 | R | 1 | 6 | C | 23-24 | SR3 | BY05 | C | 4 | 11-12 | BY02 | C | 2 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
47 | BJBAP.A6R3 |  |  | BLECF_ | 541 | 0x | 021D | 4 | UJ33 | BY03 | 3 | R | 1 | 7 | D | 1-2 | SR3 | BY05 | C | 5 | 1-2 | BY02 | C | 2 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
48 | BJSAP.B6R3 |  |  | BLECF_ | 341 | 0x | 0155 | 4 | UJ33 | BY03 | 3 | R | 1 | 8 | D | 3-4 | SR3 | BY05 | C | 5 | 3-4 | BY02 | C | 2 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
49 | BJBAP.C6R3 |  |  | BLECF_ | 339 | 0x | 0153 | 6 | UJ33 | BY03 | 3 | R | 1 | 9 | D | 5-6 | SR3 | BY05 | C | 5 | 5-6 | BY01 | R | 3 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
50 | BJSAP.A7R3 |  |  | BLECF_ | 367 | 0x | 016F | 1 | UJ33 | BY03 | 3 | R | 1 | 10 | D | 7-8 | SR3 | BY05 | C | 5 | 7-8 | BY02 | C | 2 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
51 | BJBAP.B7R3 |  |  | BLECF_ | 709 | 0x | 02C5 | 7 | UJ33 | BY03 | 3 | R | 2 | 1 | D | 9-10 | SR3 | BY05 | C | 5 | 9-10 | BY01 | R | 3 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
52 | BJBAP.A8R3 |  |  | BLECF_ | 502 | 0x | 01F6 | 6 | UJ33 | BY03 | 3 | R | 2 | 2 | D | 11-12 | SR3 | BY05 | C | 5 | 11-12 | BY02 | C | 2 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
53 | BJBAP.B8R3 |  |  | BLECF_ | 183 | 0x | 00B7 | 6 | UJ33 | BY03 | 3 | R | 2 | 3 | D | 13-14 | SR3 | BY05 | C | 6 | 1-2 | BY01 | R | 3 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
54 | BJBAP.A9R3 |  |  | BLECF_ | 375 | 0x | 0177 | 10 | UJ33 | BY03 | 3 | R | 2 | 4 | D | 15-16 | SR3 | BY05 | C | 6 | 3-4 | BY02 | C | 2 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
55 | BJBAP.B9R3 |  | NES18 | BLECF_ | 410 | 0x | 019A | 11 | UJ33 | BY03 | 3 | R | 2 | 5 | D | 17-18 | SR3 | BY05 | C | 6 | 5-6 | BY01 | R | 3 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
56 | BJBAP.A10R3 |  |  | BLECF_ | 186 | 0x | 00BA | 11 | UJ33 | BY03 | 3 | R | 2 | 6 | D | 19-20 | SR3 | BY05 | C | 6 | 7-8 | BY01 | R | 3 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
57 | BJBAP.A11R3 |  |  | BLECF_ | 562 | 0x | 0232 | 12 | UJ33 | BY03 | 3 | R | 2 | 7 | D | 21-22 | SR3 | BY05 | C | 6 | 9-10 | BY02 | C | 2 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
58 | BJBAP.B11R3 |  | NES18 | BLECF_ | 459 | 0x | 01CB | 7 | UJ33 | BY03 | 3 | R | 2 | 8 | D | 23-24 | SR3 | BY05 | C | 6 | 11-12 | BY01 | R | 3 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
59 | BJBAP.A12R3 |  |  | BLECF_ | 68 | 0x | 0044 | 10 | Arc | BYPLM.A12R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 8 | 7-8 | BY01 | R | 3 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
60 | BJBAP.A13R3 |  |  | BLECF_ | 527 | 0x | 020F | 10 | Arc | BYPLM.A13R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 9 | 7-8 | BY01 | R | 3 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
61 | BJBAP.A14R3 |  |  | BLECF_ | 504 | 0x | 01F8 | 6 | Arc | BYPLM.A14R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 10 | 7-8 | BY01 | R | 3 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
62 | BJBAP.A15R3 |  |  | BLECF_ | 499 | 0x | 01F3 | 6 | Arc | BYPLM.A15R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 11 | 7-8 | BY01 | R | 3 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
63 | BJBAP.A16R3 |  |  | BLECF_ | 514 | 0x | 0202 | 6 | Arc | BYPLM.A16R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | F | 12 | 7-8 | BY01 | R | 3 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
64 | BJBAP.A17R3 |  |  | BLECF_ | 493 | 0x | 01ED | 6 | Arc | BYPLM.A17R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 7 | 7-8 | BY01 | R | 3 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
65 | BJBAP.A18R3 |  |  | BLECF_ | 496 | 0x | 01F0 | 6 | Arc | BYPLM.A18R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 8 | 7-8 | BY01 | R | 3 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
66 | BJBAP.A19R3 |  |  | BLECF_ | 314 | 0x | 013A | 6 | Arc | BYPLM.A19R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 9 | 7-8 | BY01 | R | 3 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
67 | BJBAP.A20R3 |  |  | BLECF_ | 523 | 0x | 020B | 6 | Arc | BYPLM.A20R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 10 | 7-8 | BY01 | R | 3 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
68 | BJBAP.A21R3 |  |  | BLECF_ | 185 | 0x | 00B9 | 6 | Arc | BYPLM.A21R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 11 | 7-8 | BY01 | R | 3 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
69 | BJBAP.A22R3 |  |  | BLECF_ | 387 | 0x | 0183 | 6 | Arc | BYPLM.A22R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | G | 12 | 7-8 | BY01 | R | 3 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
70 | BJBAP.A23R3 |  |  | BLECF_ | 485 | 0x | 01E5 | 6 | Arc | BYPLM.A23R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 7 | 7-8 | BY01 | R | 3 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
71 | BJBAP.A24R3 |  |  | BLECF_ | 309 | 0x | 0135 | 6 | Arc | BYPLM.A24R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 8 | 7-8 | BY01 | R | 3 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
72 | BJBAP.A25R3 |  |  | BLECF_ | 184 | 0x | 00B8 | 6 | Arc | BYPLM.A25R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 9 | 7-8 | BY01 | R | 3 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
73 | BJBAP.A26R3 |  |  | BLECF_ | 337 | 0x | 0151 | 6 | Arc | BYPLM.A26R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 10 | 7-8 | BY01 | R | 3 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
74 | BJBAP.A27R3 |  |  | BLECF_ | 519 | 0x | 0207 | 6 | Arc | BYPLM.A27R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 11 | 7-8 | BY01 | R | 3 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
75 | BJBAP.A28R3 |  |  | BLECF_ | 448 | 0x | 01C0 | 6 | Arc | BYPLM.A28R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | H | 12 | 7-8 | BY01 | R | 3 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
76 | BJBAP.A29R3 |  |  | BLECF_ | 489 | 0x | 01E9 | 6 | Arc | BYPLM.A29R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 7 | 7-8 | BY01 | R | 3 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
77 | BJBAP.A30R3 |  |  | BLECF_ | 492 | 0x | 01EC | 6 | Arc | BYPLM.A30R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 8 | 7-8 | BY01 | R | 3 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
78 | BJBAP.A31R3 |  |  | BLECF_ | 488 | 0x | 01E8 | 6 | Arc | BYPLM.A31R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 9 | 7-8 | BY01 | R | 3 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
79 | BJBAP.A32R3 |  |  | BLECF_ | 407 | 0x | 0197 | 6 | Arc | BYPLM.A32R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 10 | 7-8 | BY01 | R | 3 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
80 | BJBAP.A33R3 |  |  | BLECF_ | 453 | 0x | 01C5 | 6 | Arc | BYPLM.A33R3 | 3 | R |  |  |  | 7-8 | SR3 | BY05 | I | 11 | 7-8 | BY01 | R | 3 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
81 | BJBAP.A34R3 |  |  | BLECF_ | 109 | 0x | 006D | 6 | Arc | BYPLM.A34R3 | 3 | R |  |  |  | 6-7 | SR3 | BY05 | I | 12 | 6-7 | BY01 | R | 3 | 20 | 16 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
Door
[28 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[28 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[8 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[8 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[8 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[28 empty] | O.K. | No problems observed

## Point 4
| Tunnel - position | [15 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | NES18 | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 1024 | HV | Comments
Pos. | position |  |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBAP.A33L4 |  |  | BLECF_ | 452 | 0x | 01C4 | 6 | Arc | BYPLM.A33L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
2 | BJBAP.A32L4 |  |  | BLECF_ | 748 | 0x | 02EC | 6 | Arc | BYPLM.A32L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
3 | BJBAP.A31L4 |  |  | BLECF_ | 616 | 0x | 0268 | 6 | Arc | BYPLM.A31L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
4 | BJBAP.A30L4 |  |  | BLECF_ | 662 | 0x | 0296 | 6 | Arc | BYPLM.A30L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
5 | BJBAP.A29L4 |  |  | BLECF_ | 38 | 0x | 0026 | 6 | Arc | BYPLM.A29L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
6 | BJBAP.A28L4 |  |  | BLECF_ | 76 | 0x | 004C | 6 | Arc | BYPLM.A28L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
7 | BJBAP.A27L4 |  |  | BLECF_ | 572 | 0x | 023C | 6 | Arc | BYPLM.A27L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
8 | BJBAP.A26L4 |  |  | BLECF_ | 21 | 0x | 0015 | 6 | Arc | BYPLM.A26L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
9 | BJBAP.A25L4 |  |  | BLECF_ | 613 | 0x | 0265 | 6 | Arc | BYPLM.A25L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
10 | BJBAP.A24L4 |  |  | BLECF_ | 655 | 0x | 028F | 6 | Arc | BYPLM.A24L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
11 | BJBAP.A23L4 |  |  | BLECF_ | 631 | 0x | 0277 | 6 | Arc | BYPLM.A23L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
12 | BJBAP.A22L4 |  |  | BLECF_ | 743 | 0x | 02E7 | 6 | Arc | BYPLM.A22L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
13 | BJBAP.A21L4 |  |  | BLECF_ | 664 | 0x | 0298 | 6 | Arc | BYPLM.A21L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
14 | BJBAP.A20L4 |  |  | BLECF_ | 471 | 0x | 01D7 | 6 | Arc | BYPLM.A20L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
15 | BJBAP.A19L4 |  |  | BLECF_ | 552 | 0x | 0228 | 6 | Arc | BYPLM.A19L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
16 | BJBAP.A18L4 |  |  | BLECF_ | 395 | 0x | 018B | 6 | Arc | BYPLM.A18L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
17 | BJBAP.A17L4 |  |  | BLECF_ | 615 | 0x | 0267 | 6 | Arc | BYPLM.A17L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
18 | BJBAP.A16L4 |  |  | BLECF_ | 591 | 0x | 024F | 6 | Arc | BYPLM.A16L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
19 | BJBAP.A15L4 |  |  | BLECF_ | 642 | 0x | 0282 | 6 | Arc | BYPLM.A15L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
20 | BJBAP.A14L4 |  |  | BLECF_ | 670 | 0x | 029E | 6 | Arc | BYPLM.A14L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
21 | BJBAP.A13L4 |  |  | BLECF_ | 599 | 0x | 0257 | 6 | Arc | BYPLM.A13L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
22 | BJBAP.A12L4 |  |  | BLECF_ | 659 | 0x | 0293 | 6 | Arc | BYPLM.A12L4 | 4 | L |  |  |  | 7-8 | SX4 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
23 | BJBAP.A5L4 |  | NES18 | BLECF_ |  | 0x | 0000 | 0 | UA43 | BY12 | 4 | CL | 1 | 9 | A | 17-18 | SX4 | BY04 | C | 2 | 5-6 | BY02 | C | 2 | 4 | 1 | 9-16
24 | BJBAP.A11L4 |  |  | BLECF_ | 643 | 0x | 0283 | 6 | UA43 | BY12 | 4 | CL | 1 | 8 | A | 15-16 | SX4 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
25 | BJBAP.A10L4 |  | NES18 | BLECF_ | 578 | 0x | 0242 | 6 | UA43 | BY12 | 4 | CL | 1 | 7 | A | 13-14 | SX4 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
26 | BJBAP.A9L4 |  |  | BLECF_ | 408 | 0x | 0198 | 6 | UA43 | BY12 | 4 | CL | 1 | 6 | A | 11-12 | SX4 | BY05 | C | 1 | 11-12 | BY02 | L | 1 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
27 | BJBAP.A8L4 |  |  | BLECF_ | 651 | 0x | 028B | 6 | UA43 | BY12 | 4 | CL | 1 | 5 | A | 9-10 | SX4 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
28 | BJBAP.A7L4 |  |  | BLECF_ | 386 | 0x | 0182 | 6 | UA43 | BY12 | 4 | CL | 1 | 4 | A | 7-8 | SX4 | BY05 | C | 1 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
29 | BJBAP.A6L4 |  |  | BLECF_ | 157 | 0x | 009D | 6 | UA43 | BY12 | 4 | CL | 1 | 3 | A | 5-6 | SX4 | BY05 | C | 1 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
30 | BJBAP.C5L4 |  |  | BLECF_ | 744 | 0x | 02E8 | 8 | UA43 | BY12 | 4 | CL | 1 | 2 | A | 3-4 | SX4 | BY05 | C | 1 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
31 | BJBAP.B5L4 |  |  | BLECF_ | 624 | 0x | 0270 | 2 | UA43 | BY12 | 4 | CL | 1 | 1 | A | 1-2 | SX4 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
32 | BJBAP.B5R4 |  |  | BLECF_ | 8 | 0x | 0008 | 2 | UA47 | BY09 | 4 | CR | 1 | 1 | A | 1-2 | SX4 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
33 | BJBAP.C5R4 |  | NES18 | BLECF_ | 5 | 0x | 0005 | 8 | UA47 | BY09 | 4 | CR | 1 | 2 | A | 3-4 | SX4 | BY05 | C | 12 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
34 | BJBAP.A6R4 |  |  | BLECF_ | 316 | 0x | 013C | 6 | UA47 | BY09 | 4 | CR | 1 | 3 | A | 5-6 | SX4 | BY05 | C | 12 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
35 | BJBAP.A7R4 |  | NES18 | BLECF_ | 235 | 0x | 00EB | 6 | UA47 | BY09 | 4 | CR | 1 | 4 | A | 7-8 | SX4 | BY05 | C | 12 | 7-8 | BY02 | R | 3 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
36 | BJBAP.A8R4 |  |  | BLECF_ | 584 | 0x | 0248 | 6 | UA47 | BY09 | 4 | CR | 1 | 5 | A | 9-10 | SX4 | BY05 | C | 12 | 9-10 | BY02 | R | 3 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
37 | BJBAP.A9R4 |  |  | BLECF_ | 509 | 0x | 01FD | 6 | UA47 | BY09 | 4 | CR | 1 | 6 | A | 11-12 | SX4 | BY05 | C | 12 | 11-12 | BY02 | R | 3 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
38 | BJBAP.A10R4 |  |  | BLECF_ | 7 | 0x | 0007 | 6 | UA47 | BY09 | 4 | CR | 1 | 7 | A | 13-14 | SX4 | BY05 | C | 11 | 1-2 | BY02 | R | 3 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
39 | BJBAP.A11R4 |  |  | BLECF_ | 10 | 0x | 000A | 6 | UA47 | BY09 | 4 | CR | 1 | 8 | A | 15-16 | SX4 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
40 | BJBAP.A5R4 |  |  | BLECF_ | 119 | 0x | 0077 | 3 | UA47 | BY09 | 4 | CR | 1 | 9 | A | 17-18 | SX4 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
41 | BJBAP.A12R4 |  |  | BLECF_ | 217 | 0x | 00D9 | 6 | Arc | BYPLM.A12R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
42 | BJBAP.A13R4 |  |  | BLECF_ | 531 | 0x | 0213 | 6 | Arc | BYPLM.A13R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
43 | BJBAP.A14R4 |  |  | BLECF_ | 568 | 0x | 0238 | 6 | Arc | BYPLM.A14R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
44 | BJBAP.A15R4 |  |  | BLECF_ | 57 | 0x | 0039 | 6 | Arc | BYPLM.A15R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
45 | BJBAP.A16R4 |  |  | BLECF_ | 134 | 0x | 0086 | 6 | Arc | BYPLM.A16R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
46 | BJBAP.A17R4 |  |  | BLECF_ | 403 | 0x | 0193 | 6 | Arc | BYPLM.A17R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
47 | BJBAP.A18R4 |  |  | BLECF_ | 299 | 0x | 012B | 6 | Arc | BYPLM.A18R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
48 | BJBAP.A19R4 |  |  | BLECF_ | 304 | 0x | 0130 | 6 | Arc | BYPLM.A19R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
49 | BJBAP.A20R4 |  |  | BLECF_ | 585 | 0x | 0249 | 6 | Arc | BYPLM.A20R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
50 | BJBAP.A21R4 |  |  | BLECF_ | 315 | 0x | 013B | 6 | Arc | BYPLM.A21R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
51 | BJBAP.A22R4 |  |  | BLECF_ | 340 | 0x | 0154 | 6 | Arc | BYPLM.A22R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
52 | BJBAP.A23R4 |  |  | BLECF_ | 537 | 0x | 0219 | 6 | Arc | BYPLM.A23R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
53 | BJBAP.A24R4 |  |  | BLECF_ | 405 | 0x | 0195 | 6 | Arc | BYPLM.A24R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
54 | BJBAP.A25R4 |  |  | BLECF_ | 483 | 0x | 01E3 | 6 | Arc | BYPLM.A25R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
55 | BJBAP.A26R4 |  |  | BLECF_ | 438 | 0x | 01B6 | 6 | Arc | BYPLM.A26R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
56 | BJBAP.A27R4 |  |  | BLECF_ | 547 | 0x | 0223 | 6 | Arc | BYPLM.A27R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
57 | BJBAP.A28R4 |  |  | BLECF_ | 551 | 0x | 0227 | 6 | Arc | BYPLM.A28R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
58 | BJBAP.A29R4 |  |  | BLECF_ | 83 | 0x | 0053 | 6 | Arc | BYPLM.A29R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
59 | BJBAP.A30R4 |  |  | BLECF_ | 121 | 0x | 0079 | 6 | Arc | BYPLM.A30R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
60 | BJBAP.A31R4 |  |  | BLECF_ | 40 | 0x | 0028 | 6 | Arc | BYPLM.A31R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
61 | BJBAP.A32R4 |  |  | BLECF_ | 160 | 0x | 00A0 | 6 | Arc | BYPLM.A32R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
62 | BJBAP.A33R4 |  |  | BLECF_ | 582 | 0x | 0246 | 6 | Arc | BYPLM.A33R4 | 4 | R |  |  |  | 7-8 | SX4 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
63 | BJBAP.A34R4 |  |  | BLECF_ | 389 | 0x | 0185 | 6 | Arc | BYPLM.A34R4 | 4 | R |  |  |  | 7-9 | SX4 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
Door
[28 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[28 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[8 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[8 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[8 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[28 empty] | O.K. | No problems observed
Note on F66: Ewald Effinger: Number from the calibration list of Ion

## Point 5
| Tunnel - position | [15 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | NSE18 | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 1024 | HV | Comments
Pos. | position |  |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBAP.A33L5 |  |  | BLECF_ | 335 | 0x | 014F | 6 | Arc | BYPLM.A33L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
2 | BJBAP.A32L5 |  |  | BLECF_ | 330 | 0x | 014A | 6 | Arc | BYPLM.A32L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
3 | BJBAP.A31L5 |  |  | BLECF_ | 650 | 0x | 028A | 6 | Arc | BYPLM.A31L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
4 | BJBAP.A30L5 |  |  | BLECF_ | 414 | 0x | 019E | 6 | Arc | BYPLM.A30L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
5 | BJBAP.A29L5 |  |  | BLECF_ | 307 | 0x | 0133 | 6 | Arc | BYPLM.A29L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
6 | BJBAP.A28L5 |  |  | BLECF_ | 447 | 0x | 01BF | 6 | Arc | BYPLM.A28L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
7 | BJBAP.A27L5 |  |  | BLECF_ | 226 | 0x | 00E2 | 6 | Arc | BYPLM.A27L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
8 | BJBAP.A26L5 |  |  | BLECF_ | 644 | 0x | 0284 | 6 | Arc | BYPLM.A26L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
9 | BJBAP.A25L5 |  |  | BLECF_ | 424 | 0x | 01A8 | 6 | Arc | BYPLM.A25L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
10 | BJBAP.A24L5 |  |  | BLECF_ | 266 | 0x | 010A | 6 | Arc | BYPLM.A24L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
11 | BJBAP.A23L5 |  |  | BLECF_ | 34 | 0x | 0022 | 6 | Arc | BYPLM.A23L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
12 | BJBAP.A22L5 |  |  | BLECF_ | 625 | 0x | 0271 | 6 | Arc | BYPLM.A22L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
13 | BJBAP.A21L5 |  |  | BLECF_ | 506 | 0x | 01FA | 6 | Arc | BYPLM.A21L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
14 | BJBAP.A20L5 |  |  | BLECF_ | 589 | 0x | 024D | 6 | Arc | BYPLM.A20L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
15 | BJBAP.A19L5 |  |  | BLECF_ | 546 | 0x | 0222 | 6 | Arc | BYPLM.A19L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
16 | BJBAP.A18L5 |  |  | BLECF_ | 473 | 0x | 01D9 | 6 | Arc | BYPLM.A18L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
17 | BJBAP.A17L5 |  |  | BLECF_ | 475 | 0x | 01DB | 6 | Arc | BYPLM.A17L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
18 | BJBAP.A16L5 |  |  | BLECF_ | 196 | 0x | 00C4 | 6 | Arc | BYPLM.A16L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
19 | BJBAP.A15L5 |  |  | BLECF_ | 346 | 0x | 015A | 6 | Arc | BYPLM.A15L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
20 | BJBAP.A14L5 |  |  | BLECF_ | 48 | 0x | 0030 | 6 | Arc | BYPLM.A14L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
21 | BJBAP.A13L5 |  |  | BLECF_ | 640 | 0x | 0280 | 6 | Arc | BYPLM.A13L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
22 | BJBAP.A12L5 |  |  | BLECF_ | 477 | 0x | 01DD | 6 | Arc | BYPLM.A12L5 | 5 | L |  |  |  | 7-8 | SR5 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
23 | BJBAP.B11L5 |  |  | BLECF_ | 753 | 0x | 02F1 | 6 | RR53 | BY02 | 5 | CL | 2 | 9 | B | 13-14 | SR5 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
24 | BJBAP.A11L5 |  |  | BLECF_ | 581 | 0x | 0245 | 7 | RR53 | BY02 | 5 | CL | 2 | 8 | B | 11-12 | SR5 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
25 | BJBAP.B10L5 |  |  | BLECF_ | 14 | 0x | 000E | 6 | RR53 | BY02 | 5 | CL | 2 | 7 | B | 9-10 | SR5 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
26 | BJBAP.A10L5 |  |  | BLECF_ | 124 | 0x | 007C | 2 | RR53 | BY02 | 5 | CL | 2 | 6 | B | 7-8 | SR5 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
27 | BJBAP.A9L5 |  |  | BLECF_ | 757 | 0x | 02F5 | 6 | RR53 | BY02 | 5 | CL | 2 | 5 | B | 5-6 | SR5 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
28 | BJBAP.B8L5 |  |  | BLECF_ | 190 | 0x | 00BE | 6 | RR53 | BY02 | 5 | CL | 2 | 4 | B | 3-4 | SR5 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
29 | BJBAP.A8L5 |  | NES18 | BLECF_ | 9 | 0x | 0009 | 6 | RR53 | BY02 | 5 | CL | 2 | 3 | B | 1-2 | SR5 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
30 | BJBAP.A7L5 |  |  | BLECF_ | 525 | 0x | 020D | 6 | RR53 | BY02 | 5 | CL | 2 | 2 | A | 23-24 | SR5 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
31 | BJBAP.B6L5 |  | NES18 | BLECF_ | 632 | 0x | 0278 | 7 | RR53 | BY02 | 5 | CL | 2 | 1 | A | 21-22 | SR5 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
32 | BJBAP.A6L5 |  |  | BLECF_ | 535 | 0x | 0217 | 1 | RR53 | BY02 | 5 | CL | 1 | 10 | A | 19-20 | SR5 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
33 | BJBAP.B5L5 |  | NES18 | BLECF_ | 313 | 0x | 0139 | 6 | RR53 | BY02 | 5 | CL | 1 | 9 | A | 17-18 | SR5 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
34 | BJSAP.A5L5 |  |  | BLECF_ | 521 | 0x | 0209 | 2 | RR53 | BY02 | 5 | CL | 1 | 8 | A | 15-16 | SR5 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
35 | BJBAP.C4L5 |  | NES18 | BLECF_ | 18 | 0x | 0012 | 6 | RR53 | BY02 | 5 | CL | 1 | 7 | A | 13-14 | SR5 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
36 | BJSAP.B4L5 |  |  | BLECF_ | 370 | 0x | 0172 | 7 | RR53 | BY02 | 5 | CL | 1 | 6 | A | 11-12 | SR5 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
37 | BJBAP.A4L5 |  | NES18 | BLECF_ | 3 | 0x | 0003 | 2 | RR53 | BY02 | 5 | CL | 1 | 5 | A | 9-10 | SR5 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
38 | BJBAP.A3L5 |  | NES18 | BLECF_ | 72 | 0x | 0048 | 5 | RR53 | BY02 | 5 | CL | 1 | 4 | A | 7-8 | SR5 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
39 | BJBAP.A2L5 |  | NES18 | BLECF_ | 288 | 0x | 0120 | 8 | RR53 | BY02 | 5 | CL | 1 | 3 | A | 5-6 | SR5 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
40 | BJBAP.B1L5 |  |  | BLECF_ | 586 | 0x | 024A | 4 | RR53 | BY02 | 5 | CL | 1 | 2 | A | 3-4 | SR5 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
41 | BJSAP.A1L5 | yes |  | BLECF_ | 579 | 0x | 0243 | 2 | RR53 | BY02 | 5 | CL | 1 | 1 | A | 1-2 | SR5 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
42 | BJSAP.A1R5 | yes |  | BLECF_ | 176 | 0x | 00B0 | 2 | RR57 | BY02 | 5 | CR | 1 | 1 | A | 1-2 | SR5 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
43 | BJBAP.B1R5 |  | NES18 | BLECF_ | 749 | 0x | 02ED | 6 | RR57 | BY02 | 5 | CR | 1 | 2 | A | 3-4 | SR5 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
44 | BJBAP.A2R5 |  |  | BLECF_ | 671 | 0x | 029F | 6 | RR57 | BY02 | 5 | CR | 1 | 3 | A | 5-6 | SR5 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
45 | BJBAP.A3R5 |  |  | BLECF_ | 750 | 0x | 02EE | 6 | RR57 | BY02 | 5 | CR | 1 | 4 | A | 7-8 | SR5 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
46 | BJBAP.A4R5 |  |  | BLECF_ | 726 | 0x | 02D6 | 2 | RR57 | BY02 | 5 | CR | 1 | 5 | A | 9-10 | SR5 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
47 | BJSAP.B4R5 |  |  | BLECF_ | 663 | 0x | 0297 | 8 | RR57 | BY02 | 5 | CR | 1 | 6 | A | 11-12 | SR5 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
48 | BJBAP.C4R5 |  |  | BLECF_ | 710 | 0x | 02C6 | 6 | RR57 | BY02 | 5 | CR | 1 | 7 | A | 13-14 | SR5 | BY05 | C | 11 | 1-2 | BY01 | R | 3 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
49 | BJSAP.A5R5 |  |  | BLECF_ | 598 | 0x | 0256 | 2 | RR57 | BY02 | 5 | CR | 1 | 8 | A | 15-16 | SR5 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
50 | BJBAP.B5R5 |  |  | BLECF_ | 697 | 0x | 02B9 | 6 | RR57 | BY02 | 5 | CR | 1 | 9 | A | 17-18 | SR5 | BY05 | C | 11 | 5-6 | BY01 | R | 3 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
51 | BJBAP.A6R5 |  |  | BLECF_ | 648 | 0x | 0288 | 2 | RR57 | BY02 | 5 | CR | 1 | 10 | A | 19-20 | SR5 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
52 | BJBAP.B6R5 |  | NES18 | BLECF_ | 656 | 0x | 0290 | 6 | RR57 | BY02 | 5 | CR | 2 | 1 | A | 21-22 | SR5 | BY05 | C | 11 | 9-10 | BY01 | R | 3 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
53 | BJBAP.A7R5 |  |  | BLECF_ | 383 | 0x | 017F | 6 | RR57 | BY02 | 5 | CR | 2 | 2 | A | 23-24 | SR5 | BY05 | C | 11 | 11-12 | BY01 | R | 3 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
54 | BJBAP.A8R5 |  |  | BLECF_ | 596 | 0x | 0254 | 6 | RR57 | BY02 | 5 | CR | 2 | 3 | B | 1-2 | SR5 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
55 | BJBAP.B8R5 |  |  | BLECF_ | 660 | 0x | 0294 | 6 | RR57 | BY02 | 5 | CR | 2 | 4 | B | 3-4 | SR5 | BY05 | C | 10 | 3-4 | BY01 | R | 3 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
56 | BJBAP.A9R5 |  |  | BLECF_ | 149 | 0x | 0095 | 6 | RR57 | BY02 | 5 | CR | 2 | 5 | B | 5-6 | SR5 | BY05 | C | 10 | 5-6 | BY01 | R | 3 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
57 | BJBAP.A10R5 |  |  | BLECF_ | 605 | 0x | 025D | 2 | RR57 | BY02 | 5 | CR | 2 | 6 | B | 7-8 | SR5 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
58 | BJBAP.B10R5 |  | NES18 | BLECF_ | 673 | 0x | 02A1 | 6 | RR57 | BY02 | 5 | CR | 2 | 7 | B | 9-10 | SR5 | BY05 | C | 10 | 9-10 | BY01 | R | 3 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
59 | BJBAP.A11R5 |  |  | BLECF_ | 621 | 0x | 026D | 8 | RR57 | BY02 | 5 | CR | 2 | 8 | B | 11-12 | SR5 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
60 | BJBAP.B11R5 |  |  | BLECF_ | 603 | 0x | 025B | 6 | RR57 | BY02 | 5 | CR | 2 | 9 | B | 13-14 | SR5 | BY05 | C | 9 | 1-2 | BY01 | R | 3 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
61 | BJBAP.A12R5 |  |  | BLECF_ | 699 | 0x | 02BB | 6 | Arc | BYPLM.A12R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 8 | 7-8 | BY01 | R | 3 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
62 | BJBAP.A13R5 |  |  | BLECF_ | 728 | 0x | 02D8 | 6 | Arc | BYPLM.A13R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 9 | 7-8 | BY01 | R | 3 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
63 | BJBAP.A14R5 |  |  | BLECF_ | 698 | 0x | 02BA | 6 | Arc | BYPLM.A14R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 10 | 7-8 | BY01 | R | 3 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
64 | BJBAP.A15R5 |  |  | BLECF_ | 737 | 0x | 02E1 | 6 | Arc | BYPLM.A15R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 11 | 7-8 | BY01 | R | 3 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
65 | BJBAP.A16R5 |  |  | BLECF_ | 733 | 0x | 02DD | 6 | Arc | BYPLM.A16R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | F | 12 | 7-8 | BY01 | R | 3 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
66 | BJBAP.A17R5 |  |  | BLECF_ | 634 | 0x | 027A | 6 | Arc | BYPLM.A17R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 7 | 7-8 | BY01 | R | 3 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
67 | BJBAP.A18R5 |  |  | BLECF_ | 732 | 0x | 02DC | 6 | Arc | BYPLM.A18R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 8 | 7-8 | BY01 | R | 3 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
68 | BJBAP.A19R5 |  |  | BLECF_ | 742 | 0x | 02E6 | 6 | Arc | BYPLM.A19R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 9 | 7-8 | BY01 | R | 3 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
69 | BJBAP.A20R5 |  |  | BLECF_ | 695 | 0x | 02B7 | 6 | Arc | BYPLM.A20R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 10 | 7-8 | BY01 | R | 3 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
70 | BJBAP.A21R5 |  |  | BLECF_ | 688 | 0x | 02B0 | 6 | Arc | BYPLM.A21R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 11 | 7-8 | BY01 | R | 3 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
71 | BJBAP.A22R5 |  |  | BLECF_ | 739 | 0x | 02E3 | 6 | Arc | BYPLM.A22R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | G | 12 | 7-8 | BY01 | R | 3 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
72 | BJBAP.A23R5 |  |  | BLECF_ | 691 | 0x | 02B3 | 6 | Arc | BYPLM.A23R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 7 | 7-8 | BY01 | R | 3 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
73 | BJBAP.A24R5 |  |  | BLECF_ | 617 | 0x | 0269 | 6 | Arc | BYPLM.A24R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 8 | 7-8 | BY01 | R | 3 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
74 | BJBAP.A25R5 |  |  | BLECF_ | 736 | 0x | 02E0 | 6 | Arc | BYPLM.A25R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 9 | 7-8 | BY01 | R | 3 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
75 | BJBAP.A26R5 |  |  | BLECF_ | 683 | 0x | 02AB | 6 | Arc | BYPLM.A26R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 10 | 7-8 | BY01 | R | 3 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
76 | BJBAP.A27R5 |  |  | BLECF_ | 490 | 0x | 01EA | 6 | Arc | BYPLM.A27R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 11 | 7-8 | BY01 | R | 3 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
77 | BJBAP.A28R5 |  |  | BLECF_ | 50 | 0x | 0032 | 6 | Arc | BYPLM.A28R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | H | 12 | 7-8 | BY01 | R | 3 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
78 | BJBAP.A29R5 |  |  | BLECF_ | 701 | 0x | 02BD | 6 | Arc | BYPLM.A29R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 7 | 7-8 | BY01 | R | 3 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
79 | BJBAP.A30R5 |  |  | BLECF_ | 704 | 0x | 02C0 | 6 | Arc | BYPLM.A30R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 8 | 7-8 | BY01 | R | 3 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
80 | BJBAP.A31R5 |  |  | BLECF_ | 693 | 0x | 02B5 | 6 | Arc | BYPLM.A31R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 9 | 7-8 | BY01 | R | 3 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
81 | BJBAP.A32R5 |  |  | BLECF_ | 735 | 0x | 02DF | 6 | Arc | BYPLM.A32R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 10 | 7-8 | BY01 | R | 3 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
82 | BJBAP.A33R5 |  |  | BLECF_ | 731 | 0x | 02DB | 6 | Arc | BYPLM.A33R5 | 5 | R |  |  |  | 7-8 | SR5 | BY05 | I | 11 | 7-8 | BY01 | R | 3 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | doesn't reset remotely
83 | BJBAP.A34R5 |  |  | BLECF_ | 661 | 0x | 0295 | 6 | Arc | BYPLM.A34R5 | 5 | R |  |  |  | 7-9 | SR5 | BY05 | I | 12 | 7-8 | BY01 | R | 3 | 20 | 16 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
84 | cryoblm |  |  | BLECF_ | 168 | 0x | 00A8 | 6 |  |  | 5 | R |  |  |  |  | SR6 | BY06 | I |  |  | BY02 | R | 3 | 20 | 16 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | card 16 9-16 present, from cryoblm
Door
[27 empty] | ↑↑↑ | ↑↑↑ | CRC Error consantly fast increasing | [7 empty] | Level on the channel o.k
[28 empty] | ↑↑ | CRC Error sometimes fast increasing | [7 empty] | Problem on the channel (e.g. negative offset current)
[8 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[8 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [5 empty] | 0 | Adjustment of 10pA potentiometer needed
[8 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [5 empty] | FF | Problem with one channel (e.g. negative offset current)
[28 empty] | O.K. | No problems observed
Note on Q37: Ewald Effinger: C - 4 - 3 to B - 15
Note on V37: Ewald Effinger: C - 4 - 3 to B - 15

## Point 6
| Tunnel - position | [15 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | NES18 | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 1024 | HV | Comments
Pos. | position |  |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBAP.A33L6 |  |  | BLECF_ | 533 | 0x | 0215 | 6 | Arc | BYPLM.A33L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
2 | BJBAP.A32L6 |  |  | BLECF_ | 117 | 0x | 0075 | 6 | Arc | BYPLM.A32L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
3 | BJBAP.A31L6 |  |  | BLECF_ | 232 | 0x | 00E8 | 6 | Arc | BYPLM.A31L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
4 | BJBAP.A30L6 |  |  | BLECF_ | 93 | 0x | 005D | 6 | Arc | BYPLM.A30L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
5 | BJBAP.A29L6 |  |  | BLECF_ | 163 | 0x | 00A3 | 6 | Arc | BYPLM.A29L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
6 | BJBAP.A28L6 |  |  | BLECF_ | 91 | 0x | 005B | 6 | Arc | BYPLM.A28L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
7 | BJBAP.A27L6 |  |  | BLECF_ | 166 | 0x | 00A6 | 6 | Arc | BYPLM.A27L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
8 | BJBAP.A26L6 |  |  | BLECF_ | 848 | 0x | 0350 | 6 | Arc | BYPLM.A26L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
9 | BJBAP.A25L6 |  |  | BLECF_ | 162 | 0x | 00A2 | 6 | Arc | BYPLM.A25L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
10 | BJBAP.A24L6 |  |  | BLECF_ | 549 | 0x | 0225 | 6 | Arc | BYPLM.A24L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
11 | BJBAP.A23L6 |  |  | BLECF_ | 177 | 0x | 00B1 | 6 | Arc | BYPLM.A23L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
12 | BJBAP.A22L6 |  |  | BLECF_ | 197 | 0x | 00C5 | 6 | Arc | BYPLM.A22L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
13 | BJBAP.A21L6 |  |  | BLECF_ | 830 | 0x | 033E | 6 | Arc | BYPLM.A21L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
14 | BJBAP.A20L6 |  |  | BLECF_ | 179 | 0x | 00B3 | 6 | Arc | BYPLM.A20L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
15 | BJBAP.A19L6 |  |  | BLECF_ | 843 | 0x | 034B | 6 | Arc | BYPLM.A19L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
16 | BJBAP.A18L6 |  |  | BLECF_ | 207 | 0x | 00CF | 6 | Arc | BYPLM.A18L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
17 | BJBAP.A17L6 |  |  | BLECF_ | 813 | 0x | 032D | 6 | Arc | BYPLM.A17L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
18 | BJBAP.A16L6 |  |  | BLECF_ | 708 | 0x | 02C4 | 6 | Arc | BYPLM.A16L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
19 | BJBAP.A15L6 |  |  | BLECF_ | 230 | 0x | 00E6 | 6 | Arc | BYPLM.A15L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
20 | BJBAP.A14L6 |  |  | BLECF_ | 281 | 0x | 0119 | 6 | Arc | BYPLM.A14L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
21 | BJBAP.A13L6 |  |  | BLECF_ | 268 | 0x | 010C | 6 | Arc | BYPLM.A13L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
22 | BJBAP.A12L6 |  |  | BLECF_ | 172 | 0x | 00AC | 6 | Arc | BYPLM.A12L6 | 6 | L |  |  |  | 7-8 | SR6 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
23 | UD62 |  |  | BLECF_ | 187 | 0x | 00BB | 8 | UJ63 | BY03 | 6 | CL | 3 | 1 | B | 17-18 | SR6 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
24 | TD62.2 | yes |  | BLECF_ | 255 | 0x | 00FF | 8 | UJ63 | BY03 | 6 | CL | 2 | 10 | B | 15-16 | SR6 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
25 | TD62.1 | yes |  | BLECF_ | 511 | 0x | 01FF | 8 | UJ63 | BY03 | 6 | CL | 2 | 9 | B | 13-14 | SR6 | BY05 | C | 4 | 1-2 | BY02 | C | 2 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
26 | BJBAP.A11L6 |  |  | BLECF_ | 108 | 0x | 006C | 6 | UJ63 | BY03 | 6 | CL | 2 | 8 | B | 11-12 | SR6 | BY05 | C | 3 | 11-12 | BY02 | L | 1 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
27 | BJBAP.A10L6 |  | NES18 | BLECF_ | 866 | 0x | 0362 | 6 | UJ63 | BY03 | 6 | CL | 2 | 7 | B | 9-10 | SR6 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
28 | BJBAP.A9L6 |  |  | BLECF_ | 161 | 0x | 00A1 | 6 | UJ63 | BY03 | 6 | CL | 2 | 6 | B | 7-8 | SR6 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
29 | BJBAP.A8L6 |  | NES18 | BLECF_ | 115 | 0x | 0073 | 6 | UJ63 | BY03 | 6 | CL | 2 | 5 | B | 5-6 | SR6 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
30 | BJBAP.C5L6 |  |  | BLECF_ | 49 | 0x | 0031 | 1 | UJ63 | BY03 | 6 | CL | 2 | 4 | B | 3-4 | SR6 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
31 | BJBAP.B5L6 |  |  | BLECF_ | 429 | 0x | 01AD | 8 | UJ63 | BY03 | 6 | CL | 2 | 3 | B | 1-2 | SR6 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
32 | BJSAP.A5L6 |  |  | BLECF_ | 107 | 0x | 006B | 3 | UJ63 | BY03 | 6 | CL | 2 | 2 | A | 23-24 | SR6 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
33 | BJBAP.K4L6 |  |  | BLECF_ | 223 | 0x | 00DF | 8 | UJ63 | BY03 | 6 | CL | 2 | 1 | A | 21-22 | SR6 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
34 | BJSAP.J4L6 | yes |  | BLECF_ | 548 | 0x | 0224 | 5 | UJ63 | BY03 | 6 | CL | 1 | 10 | A | 19-20 | SR6 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
35 | BJBAP.I4L6 |  |  | BLECF_ | 508 | 0x | 01FC | 5 | UJ63 | BY03 | 6 | CL | 1 | 9 | A | 17-18 | SR6 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
36 | BJBAP.H4L6 |  |  | BLECF_ | 336 | 0x | 0150 | 3 | UJ63 | BY03 | 6 | CL | 1 | 8 | A | 15-16 | SR6 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
37 | BJBAP.G4L6 |  |  | BLECF_ | 66 | 0x | 0042 | 3 | UJ63 | BY03 | 6 | CL | 1 | 7 | A | 13-14 | SR6 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
38 | BJSAP.F4L6 |  |  | BLECF_ | 610 | 0x | 0262 | 2 | UJ63 | BY03 | 6 | CL | 1 | 6 | A | 11-12 | SR6 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
39 | BJBAP.E4L6 |  | NES18 | BLECF_ | 392 | 0x | 0188 | 2 | UJ63 | BY03 | 6 | CL | 1 | 5 | A | 9-10 | SR6 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
40 | BJSAP.D4L6 |  |  | BLECF_ | 612 | 0x | 0264 | 8 | UJ63 | BY03 | 6 | CL | 1 | 4 | A | 7-8 | SR6 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
41 | BJBAP.C4L6 |  | NES18 | BLECF_ | 12 | 0x | 000C | 8 | UJ63 | BY03 | 6 | CL | 1 | 3 | A | 5-6 | SR6 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
42 | BJSAP.B4L6 |  |  | BLECF_ | 436 | 0x | 01B4 | 7 | UJ63 | BY03 | 6 | CL | 1 | 2 | A | 3-4 | SR6 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
43 | BJBAP.A4L6 | yes |  | BLECF_ | 144 | 0x | 0090 | 7 | UJ63 | BY03 | 6 | CL | 1 | 1 | A | 1-2 | SR6 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
44 | BJBAP.A4R6 | yes |  | BLECF_ | 102 | 0x | 0066 | 7 | UJ67 | BY03 | 6 | CR | 1 | 1 | A | 1-2 | SR6 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
45 | BJSAP.B4R6 |  |  | BLECF_ | 561 | 0x | 0231 | 7 | UJ67 | BY03 | 6 | CR | 1 | 2 | A | 3-4 | SR6 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
46 | BJBAP.C4R6 | yes |  | BLECF_ | 156 | 0x | 009C | 8 | UJ67 | BY03 | 6 | CR | 1 | 3 | A | 5-6 | SR6 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
47 | BJSAP.D4R6 |  |  | BLECF_ | 301 | 0x | 012D | 8 | UJ67 | BY03 | 6 | CR | 1 | 4 | A | 7-8 | SR6 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
48 | BJBAP.E4R6 | yes |  | BLECF_ | 167 | 0x | 00A7 | 2 | UJ67 | BY03 | 6 | CR | 1 | 5 | A | 9-10 | SR6 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
49 | BJSAP.F4R6 |  |  | BLECF_ | 565 | 0x | 0235 | 2 | UJ67 | BY03 | 6 | CR | 1 | 6 | A | 11-12 | SR6 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
50 | BJBAP.G4R6 |  |  | BLECF_ | 289 | 0x | 0121 | 3 | UJ67 | BY03 | 6 | CR | 1 | 7 | A | 13-14 | SR6 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
51 | BJBAP.H4R6 |  |  | BLECF_ | 45 | 0x | 002D | 3 | UJ67 | BY03 | 6 | CR | 1 | 8 | A | 15-16 | SR6 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
52 | BJBAP.I4R6 | yes |  | BLECF_ | 173 | 0x | 00AD | 5 | UJ67 | BY03 | 6 | CR | 1 | 9 | A | 17-18 | SR6 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
53 | BJSAP.J4R6 | yes |  | BLECF_ | 262 | 0x | 0106 | 4 | UJ67 | BY03 | 6 | CR | 1 | 10 | A | 19-20 | SR6 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
54 | BJBAP.K4R6 |  |  | BLECF_ | 198 | 0x | 00C6 | 8 | UJ67 | BY03 | 6 | CR | 2 | 1 | A | 21-22 | SR6 | BY05 | C | 11 | 9-10 | BY01 | R | 3 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
55 | BJSAP.A5R6 |  |  | BLECF_ | 219 | 0x | 00DB | 3 | UJ67 | BY03 | 6 | CR | 2 | 2 | A | 23-24 | SR6 | BY05 | C | 11 | 11-12 | BY02 | C | 2 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
56 | BJBAP.B5R6 |  |  | BLECF_ | 165 | 0x | 00A5 | 8 | UJ67 | BY03 | 6 | CR | 2 | 3 | B | 1-2 | SR6 | BY05 | C | 10 | 1-2 | BY01 | R | 3 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
57 | BJBAP.C5R6 |  |  | BLECF_ | 88 | 0x | 0058 | 1 | UJ67 | BY03 | 6 | CR | 2 | 4 | B | 3-4 | SR6 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
58 | BJBAP.A8R6 |  |  | BLECF_ | 545 | 0x | 0221 | 6 | UJ67 | BY03 | 6 | CR | 2 | 5 | B | 5-6 | SR6 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
59 | BJBAP.A9R6 |  |  | BLECF_ | 542 | 0x | 021E | 6 | UJ67 | BY03 | 6 | CR | 2 | 6 | B | 7-8 | SR6 | BY05 | C | 10 | 7-8 | BY01 | R | 3 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
60 | BJBAP.A10R6 |  |  | BLECF_ | 164 | 0x | 00A4 | 6 | UJ67 | BY03 | 6 | CR | 2 | 7 | B | 9-10 | SR6 | BY05 | C | 10 | 9-10 | BY01 | R | 3 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
61 | BJBAP.A11R6 |  |  | BLECF_ | 844 | 0x | 034C | 6 | UJ67 | BY03 | 6 | CR | 2 | 8 | B | 11-12 | SR6 | BY05 | C | 10 | 11-12 | BY01 | R | 3 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
62 | TD68.1 | yes |  | BLECF_ | 391 | 0x | 0187 | 8 | UJ67 | BY03 | 6 | CR | 2 | 9 | B | 13-14 | SR6 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
63 | TD68.2 | yes |  | BLECF_ | 56 | 0x | 0038 | 8 | UJ67 | BY03 | 6 | CR | 2 | 10 | B | 15-16 | SR6 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
64 | UD68 |  |  | BLECF_ | 480 | 0x | 01E0 | 8 | UJ67 | BY03 | 6 | CR | 3 | 1 | B | 17-18 | SR6 | BY05 | C | 9 | 5-6 | BY02 | C | 2 | 20 | 16 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
65 | BJBAP.A12R6 |  |  | BLECF_ | 611 | 0x | 0263 | 6 | Arc | BYPLM.A12R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 8 | 7-8 | BY01 | R | 3 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
66 | BJBAP.A13R6 |  |  | BLECF_ | 730 | 0x | 02DA | 6 | Arc | BYPLM.A13R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 9 | 7-8 | BY01 | R | 3 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
67 | BJBAP.A14R6 |  |  | BLECF_ | 202 | 0x | 00CA | 6 | Arc | BYPLM.A14R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 10 | 7-8 | BY01 | R | 3 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
68 | BJBAP.A15R6 |  |  | BLECF_ | 569 | 0x | 0239 | 6 | Arc | BYPLM.A15R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 11 | 7-8 | BY01 | R | 3 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
69 | BJBAP.A16R6 |  |  | BLECF_ | 274 | 0x | 0112 | 6 | Arc | BYPLM.A16R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | F | 12 | 7-8 | BY01 | R | 3 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
70 | BJBAP.A17R6 |  |  | BLECF_ | 125 | 0x | 007D | 6 | Arc | BYPLM.A17R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 7 | 7-8 | BY01 | R | 3 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
71 | BJBAP.A18R6 |  |  | BLECF_ | 19 | 0x | 0013 | 6 | Arc | BYPLM.A18R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 8 | 7-8 | BY01 | R | 3 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
72 | BJBAP.A19R6 |  |  | BLECF_ | 563 | 0x | 0233 | 6 | Arc | BYPLM.A19R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 9 | 7-8 | BY01 | R | 3 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
73 | BJBAP.A20R6 |  |  | BLECF_ | 566 | 0x | 0236 | 6 | Arc | BYPLM.A20R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 10 | 7-8 | BY01 | R | 3 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
74 | BJBAP.A21R6 |  |  | BLECF_ | 295 | 0x | 0127 | 6 | Arc | BYPLM.A21R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 11 | 7-8 | BY01 | R | 3 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
75 | BJBAP.A22R6 |  |  | BLECF_ | 242 | 0x | 00F2 | 6 | Arc | BYPLM.A22R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | G | 12 | 7-8 | BY01 | R | 3 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
76 | BJBAP.A23R6 |  |  | BLECF_ | 540 | 0x | 021C | 6 | Arc | BYPLM.A23R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 7 | 7-8 | BY01 | R | 3 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
77 | BJBAP.A24R6 |  |  | BLECF_ | 234 | 0x | 00EA | 6 | Arc | BYPLM.A24R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 8 | 7-8 | BY01 | R | 3 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
78 | BJBAP.A25R6 |  |  | BLECF_ | 555 | 0x | 022B | 6 | Arc | BYPLM.A25R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 9 | 7-8 | BY01 | R | 3 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
79 | BJBAP.A26R6 |  |  | BLECF_ | 680 | 0x | 02A8 | 6 | Arc | BYPLM.A26R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 10 | 7-8 | BY01 | R | 3 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
80 | BJBAP.A27R6 |  |  | BLECF_ | 557 | 0x | 022D | 6 | Arc | BYPLM.A27R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 11 | 7-8 | BY01 | R | 3 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
81 | BJBAP.A28R6 |  |  | BLECF_ | 63 | 0x | 003F | 6 | Arc | BYPLM.A28R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | H | 12 | 7-8 | BY01 | R | 3 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
82 | BJBAP.A29R6 |  |  | BLECF_ | 218 | 0x | 00DA | 6 | Arc | BYPLM.A29R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 7 | 7-8 | BY01 | R | 3 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
83 | BJBAP.A30R6 |  |  | BLECF_ | 853 | 0x | 0355 | 6 | Arc | BYPLM.A30R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 8 | 7-8 | BY01 | R | 3 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
84 | BJBAP.A31R6 |  |  | BLECF_ | 210 | 0x | 00D2 | 6 | Arc | BYPLM.A31R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 9 | 7-8 | BY01 | R | 3 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
85 | BJBAP.A32R6 |  |  | BLECF_ | 154 | 0x | 009A | 6 | Arc | BYPLM.A32R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 10 | 7-8 | BY01 | R | 3 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
86 | BJBAP.A33R6 |  |  | BLECF_ | 356 | 0x | 0164 | 6 | Arc | BYPLM.A33R6 | 6 | R |  |  |  | 7-8 | SR6 | BY05 | I | 11 | 7-8 | BY01 | R | 3 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
87 | BJBAP.A34R6 |  |  | BLECF_ | 857 | 0x | 0359 | 6 | Arc | BYPLM.A34R6 | 6 | R |  |  |  | 7-9 | SR6 | BY05 | I | 12 | 7-8 | BY01 | R | 3 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
Door
|  |  |  | BLECF_ | 176 |  | CFC card repaired after HV issue | [20 empty] | ↑↑↑ | CRC Error consantly fast increasing | [7 empty] | Level on the channel o.k
[28 empty] | ↑↑ | CRC Error sometimes fast increasing | [7 empty] | Problem on the channel (e.g. negative offset current)
[8 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[8 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [5 empty] | 0 | Adjustment of 10pA potentiometer needed
[8 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [5 empty] | FF | Problem with one channel (e.g. negative offset current)
[28 empty] | O.K. | No problems observed
Note on C29: Ewald Effinger: All Channels 1uF/100k
Note on C30: Ewald Effinger: All Channels 1uF/100k
Note on C68: Ewald Effinger: All Channels 1uF/100k
Note on C69: Ewald Effinger: All Channels 1uF/100k

## Point 7
| Tunnel - position | [15 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | NES18 | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 1024 | HV | Comments
Pos. | position |  |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBAP.A33L7 |  |  | BLECF_ | 69 | 0x | 0045 | 6 | Arc | BYPLM.A33L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
2 | BJBAP.A32L7 |  |  | BLECF_ | 122 | 0x | 007A | 6 | Arc | BYPLM.A32L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
3 | BJBAP.A31L7 |  |  | BLECF_ | 426 | 0x | 01AA | 6 | Arc | BYPLM.A31L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
4 | BJBAP.A30L7 |  |  | BLECF_ | 580 | 0x | 0244 | 6 | Arc | BYPLM.A30L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
5 | BJBAP.A29L7 |  |  | BLECF_ | 123 | 0x | 007B | 6 | Arc | BYPLM.A29L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
6 | BJBAP.A28L7 |  |  | BLECF_ | 20 | 0x | 0014 | 6 | Arc | BYPLM.A28L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
7 | BJBAP.A27L7 |  |  | BLECF_ | 140 | 0x | 008C | 6 | Arc | BYPLM.A27L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
8 | BJBAP.A26L7 |  |  | BLECF_ | 126 | 0x | 007E | 6 | Arc | BYPLM.A26L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
9 | BJBAP.A25L7 |  |  | BLECF_ | 43 | 0x | 002B | 6 | Arc | BYPLM.A25L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
10 | BJBAP.A24L7 |  |  | BLECF_ | 454 | 0x | 01C6 | 6 | Arc | BYPLM.A24L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
11 | BJBAP.A23L7 |  |  | BLECF_ | 17 | 0x | 0011 | 6 | Arc | BYPLM.A23L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
12 | BJBAP.A22L7 |  |  | BLECF_ | 47 | 0x | 002F | 6 | Arc | BYPLM.A22L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
13 | BJBAP.A21L7 |  |  | BLECF_ | 87 | 0x | 0057 | 6 | Arc | BYPLM.A21L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
14 | BJBAP.A20L7 |  |  | BLECF_ | 74 | 0x | 004A | 6 | Arc | BYPLM.A20L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
15 | BJBAP.A19L7 |  |  | BLECF_ | 466 | 0x | 01D2 | 6 | Arc | BYPLM.A19L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
16 | BJBAP.A18L7 |  |  | BLECF_ | 208 | 0x | 00D0 | 6 | Arc | BYPLM.A18L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
17 | BJBAP.A17L7 |  |  | BLECF_ | 334 | 0x | 014E | 6 | Arc | BYPLM.A17L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
18 | BJBAP.A16L7 |  |  | BLECF_ | 23 | 0x | 0017 | 6 | Arc | BYPLM.A16L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
19 | BJBAP.A15L7 |  |  | BLECF_ | 534 | 0x | 0216 | 6 | Arc | BYPLM.A15L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
20 | BJBAP.A14L7 |  |  | BLECF_ | 461 | 0x | 01CD | 6 | Arc | BYPLM.A14L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
21 | BJBAP.A13L7 |  |  | BLECF_ | 449 | 0x | 01C1 | 6 | Arc | BYPLM.A13L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
22 | BJBAP.A12L7 |  |  | BLECF_ | 472 | 0x | 01D8 | 6 | Arc | BYPLM.A12L7 | 7 | L |  |  |  | 7-8 | SR7 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
23 | BJBAP.B11L7 |  |  | BLECF_ | 469 | 0x | 01D5 | 8 | RR73 | BY01 | 7 | CL | 3 | 8 | F | 1-2 | SR7 | BY05 | C | 5 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
24 | BJBAP.A11L7 |  |  | BLECF_ | 679 | 0x | 02A7 | 8 | RR73 | BY01 | 7 | CL | 3 | 7 | E | 5-6 | SR7 | BY05 | C | 5 | 5-6 | BY01 | E | 4 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
25 | BJBAP.A10L7 |  |  | BLECF_ | 544 | 0x | 0220 | 8 | RR73 | BY01 | 7 | CL | 3 | 6 | E | 3-4 | SR7 | BY05 | C | 5 | 3-4 | BY02 | L | 1 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
26 | BJBAP.B9L7 |  |  | BLECF_ | 678 | 0x | 02A6 | 8 | RR73 | BY01 | 7 | CL | 3 | 5 | E | 1-2 | SR7 | BY05 | C | 5 | 1-2 | BY02 | L | 1 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
27 | BJBAP.A9L7 |  |  | BLECF_ | 559 | 0x | 022F | 8 | RR73 | BY01 | 7 | CL | 3 | 4 | D | 23-24 | SR7 | BY05 | C | 4 | 11-12 | BY01 | E | 4 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
28 | BJBAP.B8L7 |  |  | BLECF_ | 649 | 0x | 0289 | 6 | RR73 | BY01 | 7 | CL | 3 | 3 | D | 21-22 | SR7 | BY05 | C | 4 | 9-10 | BY02 | L | 1 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
29 | BJBAP.A8L7 |  |  | BLECF_ | 462 | 0x | 01CE | 3 | RR73 | BY01 | 7 | CL | 3 | 2 | D | 19-20 | SR7 | BY05 | C | 4 | 7-8 | BY01 | E | 4 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
30 | BJBAP.B7L7 |  |  | BLECF_ | 513 | 0x | 0201 | 7 | RR73 | BY01 | 7 | CL | 3 | 1 | D | 17-18 | SR7 | BY05 | C | 4 | 5-6 | BY02 | L | 1 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
31 | BJSAP.A7L7 |  |  | BLECF_ | 411 | 0x | 019B | 2 | RR73 | BY01 | 7 | CL | 2 | 10 | D | 15-16 | SR7 | BY05 | C | 4 | 3-4 | BY01 | E | 4 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
32 | BJBAP.I6L7 |  |  | BLECF_ | 587 | 0x | 024B | 7 | RR73 | BY01 | 7 | CL | 2 | 9 | D | 13-14 | SR7 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
33 | BJSAP.H6L7 |  |  | BLECF_ | 118 | 0x | 0076 | 4 | RR73 | BY01 | 7 | CL | 2 | 8 | D | 11-12 | SR7 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
34 | BJBAP.G6L7 |  |  | BLECF_ | 191 | 0x | 00BF | 4 | RR73 | BY01 | 7 | CL | 2 | 7 | D | 9-10 | SR7 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
35 | BJSAP.F6L7 |  |  | BLECF_ | 434 | 0x | 01B2 | 7 | RR73 | BY01 | 7 | CL | 2 | 6 | D | 7-8 | SR7 | BY05 | C | 3 | 7-8 | BY02 | C | 2 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
36 | BJBAP.E6L7 |  |  | BLECF_ | 638 | 0x | 027E | 7 | RR73 | BY01 | 7 | CL | 2 | 5 | D | 5-6 | SR7 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
37 | BJBAP.D6L7 |  |  | BLECF_ | 364 | 0x | 016C | 3 | RR73 | BY01 | 7 | CL | 2 | 4 | D | 3-4 | SR7 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
38 | BJSAP.C6L7 |  |  | BLECF_ | 360 | 0x | 0168 | 3 | RR73 | BY01 | 7 | CL | 2 | 3 | D | 1-2 | SR7 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
39 | BJSAP.B6L7 |  |  | BLECF_ | 682 | 0x | 02AA | 7 | RR73 | BY01 | 7 | CL | 2 | 2 | C | 23-24 | SR7 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
40 | BJBAP.A6L7 |  |  | BLECF_ | 265 | 0x | 0109 | 7 | RR73 | BY01 | 7 | CL | 2 | 1 | C | 21-22 | SR7 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
41 | BJBAP.E5L7 |  | NES18 | BLECF_ | 96 | 0x | 0060 | 6 | RR73 | BY01 | 7 | CL | 1 | 10 | C | 19-20 | SR7 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
42 | BJSAP.D5L7 |  |  | BLECF_ | 94 | 0x | 005E | 6 | RR73 | BY01 | 7 | CL | 1 | 9 | C | 17-18 | SR7 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
43 | BJBAP.C5L7 |  | NES18 | BLECF_ | 273 | 0x | 0111 | 6 | RR73 | BY01 | 7 | CL | 1 | 8 | C | 15-16 | SR7 | BY05 | C | 2 | 3-4 | BY02 | C | 2 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
44 | BJSAP.B5L7 |  |  | BLECF_ | 590 | 0x | 024E | 8 | RR73 | BY01 | 7 | CL | 1 | 7 | C | 13-14 | SR7 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
45 | BJBAP.A5L7 |  |  | BLECF_ | 244 | 0x | 00F4 | 8 | RR73 | BY01 | 7 | CL | 1 | 6 | C | 11-12 | SR7 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
46 | BJBAP.E4L7 |  |  | BLECF_ | 372 | 0x | 0174 | 6 | RR73 | BY01 | 7 | CL | 1 | 5 | C | 9-10 | SR7 | BY05 | C | 1 | 9-10 | BY02 | L | 1 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
47 | BJSAP.D4L7 |  |  | BLECF_ | 445 | 0x | 01BD | 6 | RR73 | BY01 | 7 | CL | 1 | 4 | C | 7-8 | SR7 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
48 | BJBAP.C4L7 |  |  | BLECF_ | 369 | 0x | 0171 | 6 | RR73 | BY01 | 7 | CL | 1 | 3 | C | 5-6 | SR7 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
49 | BJSAP.B4L7 |  |  | BLECF_ | 633 | 0x | 0279 | 6 | RR73 | BY01 | 7 | CL | 1 | 2 | C | 3-4 | SR7 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
50 | BJBAP.A4L7 |  |  | BLECF_ | 308 | 0x | 0134 | 6 | RR73 | BY01 | 7 | CL | 1 | 1 | C | 1-2 | SR7 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
51 | BJBAP.A4R7 |  |  | BLECF_ | 443 | 0x | 01BB | 6 | RR77 | BY02 | 7 | CR | 1 | 1 | C | 1-2 | SR7 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
52 | BJSAP.B4R7 |  |  | BLECF_ | 455 | 0x | 01C7 | 6 | RR77 | BY02 | 7 | CR | 1 | 2 | C | 3-4 | SR7 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
53 | BJBAP.C4R7 |  |  | BLECF_ | 351 | 0x | 015F | 6 | RR77 | BY02 | 7 | CR | 1 | 3 | C | 5-6 | SR7 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
54 | BJSAP.D4R7 |  |  | BLECF_ | 431 | 0x | 01AF | 6 | RR77 | BY02 | 7 | CR | 1 | 4 | C | 7-8 | SR7 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
55 | BJBAP.E4R7 |  |  | BLECF_ | 420 | 0x | 01A4 | 6 | RR77 | BY02 | 7 | CR | 1 | 5 | C | 9-10 | SR7 | BY05 | C | 12 | 9-10 | BY01 | R | 3 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
56 | BJBAP.A5R7 |  |  | BLECF_ | 713 | 0x | 02C9 | 8 | RR77 | BY02 | 7 | CR | 1 | 6 | C | 11-12 | SR7 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
57 | BJSAP.B5R7 |  |  | BLECF_ | 446 | 0x | 01BE | 8 | RR77 | BY02 | 7 | CR | 1 | 7 | C | 13-14 | SR7 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
58 | BJBAP.C5R7 |  | NES18 | BLECF_ | 421 | 0x | 01A5 | 6 | RR77 | BY02 | 7 | CR | 1 | 8 | C | 15-16 | SR7 | BY05 | C | 11 | 3-4 | BY02 | C | 2 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
59 | BJSAP.D5R7 |  |  | BLECF_ | 415 | 0x | 019F | 6 | RR77 | BY02 | 7 | CR | 1 | 9 | C | 17-18 | SR7 | BY05 | C | 11 | 5-6 | BY02 | C | 2 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
60 | BJBAP.E5R7 |  |  | BLECF_ | 368 | 0x | 0170 | 6 | RR77 | BY02 | 7 | CR | 1 | 10 | C | 19-20 | SR7 | BY05 | C | 11 | 7-8 | BY01 | R | 3 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
61 | BJBAP.A6R7 |  |  | BLECF_ | 192 | 0x | 00C0 | 4 | RR77 | BY02 | 7 | CR | 2 | 1 | C | 21-22 | SR7 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
62 | BJSAP.B6R7 |  |  | BLECF_ | 221 | 0x | 00DD | 4 | RR77 | BY02 | 7 | CR | 2 | 2 | C | 23-24 | SR7 | BY05 | C | 11 | 11-12 | BY02 | C | 2 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
63 | BJBAP.C6R7 |  |  | BLECF_ | 228 | 0x | 00E4 | 5 | RR77 | BY02 | 7 | CR | 2 | 3 | D | 1-2 | SR7 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
64 | BJSAP.D6R7 |  |  | BLECF_ | 245 | 0x | 00F5 | 5 | RR77 | BY02 | 7 | CR | 2 | 4 | D | 3-4 | SR7 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
65 | BJBAP.E6R7 |  |  | BLECF_ | 31 | 0x | 001F | 8 | RR77 | BY02 | 7 | CR | 2 | 5 | D | 5-6 | SR7 | BY05 | C | 10 | 5-6 | BY02 | C | 2 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
66 | BJSAP.F6R7 |  |  | BLECF_ | 13 | 0x | 000D | 8 | RR77 | BY02 | 7 | CR | 2 | 6 | D | 7-8 | SR7 | BY05 | C | 10 | 7-8 | BY02 | C | 2 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
67 | BJBAP.G6R7 |  |  | BLECF_ | 433 | 0x | 01B1 | 4 | RR77 | BY02 | 7 | CR | 2 | 7 | D | 9-10 | SR7 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 20 | 16 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
68 | BJSAP.H6R7 |  |  | BLECF_ | 482 | 0x | 01E2 | 4 | RR77 | BY02 | 7 | CR | 2 | 8 | D | 11-12 | SR7 | BY05 | C | 10 | 11-12 | BY02 | C | 2 | 20 | 16 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
69 | BJBAP.I6R7 |  |  | BLECF_ | 464 | 0x | 01D0 | 7 | RR77 | BY02 | 7 | CR | 2 | 9 | D | 13-14 | SR7 | BY05 | C | 9 | 1-2 | BY01 | R | 3 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
70 | BJSAP.A7R7 |  |  | BLECF_ | 359 | 0x | 0167 | 2 | RR77 | BY02 | 7 | CR | 2 | 10 | D | 15-16 | SR7 | BY05 | C | 9 | 3-4 | BY01 | E | 4 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
71 | BJBAP.B7R7 |  |  | BLECF_ | 470 | 0x | 01D6 | 7 | RR77 | BY02 | 7 | CR | 3 | 1 | D | 17-18 | SR7 | BY05 | C | 9 | 5-6 | BY01 | R | 3 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
72 | BJBAP.A8R7 |  |  | BLECF_ | 700 | 0x | 02BC | 3 | RR77 | BY02 | 7 | CR | 3 | 2 | D | 19-20 | SR7 | BY05 | C | 9 | 7-8 | BY01 | E | 4 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
73 | BJBAP.B8R7 |  |  | BLECF_ | 358 | 0x | 0166 | 6 | RR77 | BY02 | 7 | CR | 3 | 3 | D | 21-22 | SR7 | BY05 | C | 9 | 9-10 | BY01 | R | 3 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
74 | BJBAP.A9R7 |  |  | BLECF_ | 505 | 0x | 01F9 | 8 | RR77 | BY02 | 7 | CR | 3 | 4 | D | 23-24 | SR7 | BY05 | C | 9 | 11-12 | BY01 | E | 4 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
75 | BJBAP.B9R7 |  |  | BLECF_ | 345 | 0x | 0159 | 8 | RR77 | BY02 | 7 | CR | 3 | 5 | E | 1-2 | SR7 | BY05 | C | 8 | 1-2 | BY01 | R | 3 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
76 | BJBAP.A10R7 |  |  | BLECF_ | 381 | 0x | 017D | 8 | RR77 | BY02 | 7 | CR | 3 | 6 | E | 3-4 | SR7 | BY05 | C | 8 | 3-4 | BY01 | R | 3 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
77 | BJBAP.A11R7 |  |  | BLECF_ | 428 | 0x | 01AC | 8 | RR77 | BY02 | 7 | CR | 3 | 7 | E | 5-6 | SR7 | BY05 | C | 8 | 5-6 | BY01 | E | 4 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
78 | BJBAP.B11R7 |  | NES18 | BLECF_ | 441 | 0x | 01B9 | 7 | RR77 | BY02 | 7 | CR | 3 | 8 | F | 1-2 | SR7 | BY05 | C | 7 | 1-2 | BY01 | R | 3 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
79 | cryoblm |  |  | BLECF_ | 465 | 0x | 01D1 |  | RR77 | BY03 | 8 | CR | 3 | 9 |  |  | SR7 | BY05 |  |  |  | BY01 | R | 3 | 20 | 16 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
79 | BJBAP.A12R7 |  |  | BLECF_ | 283 | 0x | 011B | 6 | Arc | BYPLM.A12R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 8 | 7-8 | BY01 | R | 3 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
80 | BJBAP.A13R7 |  |  | BLECF_ | 716 | 0x | 02CC | 6 | Arc | BYPLM.A13R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 9 | 7-8 | BY01 | R | 3 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
81 | BJBAP.A14R7 |  |  | BLECF_ | 44 | 0x | 002C | 6 | Arc | BYPLM.A14R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 10 | 7-8 | BY01 | R | 3 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
82 | BJBAP.A15R7 |  |  | BLECF_ | 264 | 0x | 0108 | 6 | Arc | BYPLM.A15R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 11 | 7-8 | BY01 | R | 3 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
83 | BJBAP.A16R7 |  |  | BLECF_ | 413 | 0x | 019D | 6 | Arc | BYPLM.A16R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | F | 12 | 7-8 | BY01 | R | 3 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
84 | BJBAP.A17R7 |  |  | BLECF_ | 467 | 0x | 01D3 | 6 | Arc | BYPLM.A17R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 7 | 7-8 | BY01 | R | 3 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
85 | BJBAP.A18R7 |  |  | BLECF_ | 263 | 0x | 0107 | 6 | Arc | BYPLM.A18R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 8 | 7-8 | BY01 | R | 3 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
86 | BJBAP.A19R7 |  |  | BLECF_ | 291 | 0x | 0123 | 6 | Arc | BYPLM.A19R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 9 | 7-8 | BY01 | R | 3 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
87 | BJBAP.A20R7 |  |  | BLECF_ | 236 | 0x | 00EC | 6 | Arc | BYPLM.A20R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 10 | 7-8 | BY01 | R | 3 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
88 | BJBAP.A21R7 |  |  | BLECF_ | 158 | 0x | 009E | 6 | Arc | BYPLM.A21R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 11 | 7-8 | BY01 | R | 3 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
89 | BJBAP.A22R7 |  |  | BLECF_ | 256 | 0x | 0100 | 6 | Arc | BYPLM.A22R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | G | 12 | 7-8 | BY01 | R | 3 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
90 | BJBAP.A23R7 |  |  | BLECF_ | 65 | 0x | 0041 | 6 | Arc | BYPLM.A23R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 7 | 7-8 | BY01 | R | 3 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
91 | BJBAP.A24R7 |  |  | BLECF_ | 152 | 0x | 0098 | 6 | Arc | BYPLM.A24R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 8 | 7-8 | BY01 | R | 3 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
92 | BJBAP.A25R7 |  |  | BLECF_ | 220 | 0x | 00DC | 6 | Arc | BYPLM.A25R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 9 | 7-8 | BY01 | R | 3 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
93 | BJBAP.A26R7 |  |  | BLECF_ | 133 | 0x | 0085 | 6 | Arc | BYPLM.A26R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 10 | 7-8 | BY01 | R | 3 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
94 | BJBAP.A27R7 |  |  | BLECF_ | 138 | 0x | 008A | 6 | Arc | BYPLM.A27R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 11 | 7-8 | BY01 | R | 3 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
95 | BJBAP.A28R7 |  |  | BLECF_ | 238 | 0x | 00EE | 6 | Arc | BYPLM.A28R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | H | 12 | 7-8 | BY01 | R | 3 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
96 | BJBAP.A29R7 |  |  | BLECF_ | 216 | 0x | 00D8 | 6 | Arc | BYPLM.A29R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 7 | 7-8 | BY01 | R | 3 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
97 | BJBAP.A30R7 |  |  | BLECF_ | 174 | 0x | 00AE | 6 | Arc | BYPLM.A30R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 8 | 7-8 | BY01 | R | 3 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
98 | BJBAP.A31R7 |  |  | BLECF_ | 132 | 0x | 0084 | 6 | Arc | BYPLM.A31R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 9 | 7-8 | BY01 | R | 3 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
99 | BJBAP.A32R7 |  |  | BLECF_ | 397 | 0x | 018D | 6 | Arc | BYPLM.A32R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 10 | 7-8 | BY01 | R | 3 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
100 | BJBAP.A33R7 |  |  | BLECF_ | 2 | 0x | 0002 | 6 | Arc | BYPLM.A33R7 | 7 | R |  |  |  | 7-8 | SR7 | BY05 | I | 11 | 7-8 | BY01 | R | 3 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
101 | BJBAP.A34R7 |  |  | BLECF_ | 486 | 0x | 01E6 | 6 | Arc | BYPLM.A34R7 | 7 | R |  |  |  | 7-9 | SR7 | BY05 | I | 12 | 7-8 | BY01 | R | 3 | 20 | 16 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
Door
[28 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[28 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[8 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[8 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[8 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[28 empty] | O.K. | No problems observed

## Point 8
| Tunnel - position | [15 empty] | Surface position | [18 empty] | LEVEL
| HV box | Filter | NES18 | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 1024 | HV | Comments
Pos. | position |  |  | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBAP.A33L8 |  |  | BLECF_ | 229 | 0x | 00E5 | 6 | Arc | BYPLM.A33L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
2 | BJBAP.A32L8 |  |  | BLECF_ | 237 | 0x | 00ED | 6 | Arc | BYPLM.A32L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
3 | BJBAP.A31L8 |  |  | BLECF_ | 292 | 0x | 0124 | 6 | Arc | BYPLM.A31L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
4 | BJBAP.A30L8 |  |  | BLECF_ | 257 | 0x | 0101 | 6 | Arc | BYPLM.A30L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
5 | BJBAP.A29L8 |  |  | BLECF_ | 215 | 0x | 00D7 | 6 | Arc | BYPLM.A29L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
6 | BJBAP.A28L8 |  |  | BLECF_ | 213 | 0x | 00D5 | 6 | Arc | BYPLM.A28L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
7 | BJBAP.A27L8 |  |  | BLECF_ | 275 | 0x | 0113 | 6 | Arc | BYPLM.A27L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
8 | BJBAP.A26L8 |  |  | BLECF_ | 277 | 0x | 0115 | 6 | Arc | BYPLM.A26L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
9 | BJBAP.A25L8 |  |  | BLECF_ | 247 | 0x | 00F7 | 6 | Arc | BYPLM.A25L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
10 | BJBAP.A24L8 |  |  | BLECF_ | 137 | 0x | 0089 | 6 | Arc | BYPLM.A24L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
11 | BJBAP.A23L8 |  |  | BLECF_ | 145 | 0x | 0091 | 6 | Arc | BYPLM.A23L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
12 | BJBAP.A22L8 |  |  | BLECF_ | 129 | 0x | 0081 | 6 | Arc | BYPLM.A22L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
13 | BJBAP.A21L8 |  |  | BLECF_ | 112 | 0x | 0070 | 6 | Arc | BYPLM.A21L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
14 | BJBAP.A20L8 |  |  | BLECF_ | 100 | 0x | 0064 | 6 | Arc | BYPLM.A20L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
15 | BJBAP.A19L8 |  |  | BLECF_ | 114 | 0x | 0072 | 6 | Arc | BYPLM.A19L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
16 | BJBAP.A18L8 |  |  | BLECF_ | 293 | 0x | 0125 | 6 | Arc | BYPLM.A18L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
17 | BJBAP.A17L8 |  |  | BLECF_ | 150 | 0x | 0096 | 6 | Arc | BYPLM.A17L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
18 | BJBAP.A16L8 |  |  | BLECF_ | 130 | 0x | 0082 | 6 | Arc | BYPLM.A16L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
19 | BJBAP.A15L8 |  |  | BLECF_ | 101 | 0x | 0065 | 6 | Arc | BYPLM.A15L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 4 | 7-8 | BY02 | L | 1 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
20 | BJBAP.A14L8 |  |  | BLECF_ | 105 | 0x | 0069 | 6 | Arc | BYPLM.A14L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
21 | BJBAP.A13L8 |  |  | BLECF_ | 99 | 0x | 0063 | 6 | Arc | BYPLM.A13L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
22 | BJBAP.A12L8 |  |  | BLECF_ | 712 | 0x | 02C8 | 6 | Arc | BYPLM.A12L8 | 8 | L |  |  |  | 7-8 | SR8 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
23 | BJBAP.C11L8 |  |  | BLECF_ | 95 | 0x | 005F | 6 | UA83 | BY05 | 8 | CL | 2 | 9 | C | 13-14 | SR8 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
24 | BJBAP.B11L8 |  |  | BLECF_ | 92 | 0x | 005C | 6 | UA83 | BY05 | 8 | CL | 2 | 8 | C | 11-12 | SR8 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
25 | BJSAP.A11L8 | yes |  | BLECF_ | 86 | 0x | 0056 | 6 | UA83 | BY05 | 8 | CL | 2 | 7 | C | 9-10 | SR8 | BY05 | C | 3 | 9-10 | BY02 | C | 2 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
26 | BJBAP.A10L8 |  |  | BLECF_ | 67 | 0x | 0043 | 6 | UA83 | BY05 | 8 | CL | 2 | 6 | C | 7-8 | SR8 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
27 | BJBAP.A9L8 |  |  | BLECF_ | 33 | 0x | 0021 | 6 | UA83 | BY05 | 8 | CL | 2 | 5 | C | 5-6 | SR8 | BY05 | C | 3 | 5-6 | BY02 | L | 1 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
28 | BJBAP.C8L8 |  |  | BLECF_ | 224 | 0x | 00E0 | 6 | UA83 | BY05 | 8 | CL | 2 | 4 | C | 3-4 | SR8 | BY05 | C | 3 | 3-4 | BY02 | L | 1 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
29 | BJBAP.B8L8 |  |  | BLECF_ | 205 | 0x | 00CD | 6 | UA83 | BY05 | 8 | CL | 2 | 3 | C | 1-2 | SR8 | BY05 | C | 3 | 1-2 | BY02 | C | 2 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
30 | BJSAP.A8L8 | yes |  | BLECF_ | 151 | 0x | 0097 | 6 | UA83 | BY05 | 8 | CL | 2 | 2 | B | 23-24 | SR8 | BY05 | C | 2 | 11-12 | BY02 | C | 2 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
31 | BJBAP.A7L8 |  |  | BLECF_ | 743 | 0x | 02E7 | 6 | UA83 | BY05 | 8 | CL | 2 | 1 | B | 21-22 | SR8 | BY05 | C | 2 | 9-10 | BY02 | L | 1 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
32 | BJBAP.B6L8 | yes |  | BLECF_ | 239 | 0x | 00EF | 7 | UA83 | BY05 | 8 | CL | 1 | 10 | B | 19-20 | SR8 | BY05 | C | 2 | 7-8 | BY02 | L | 1 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
33 | BJSAP.A6L8 | yes |  | BLECF_ | 231 | 0x | 00E7 | 1 | UA83 | BY05 | 8 | CL | 1 | 9 | B | 17-18 | SR8 | BY05 | C | 2 | 5-6 | BY02 | C | 2 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
34 | BJBAP.A5L8 |  | NES18 | BLECF_ | 240 | 0x | 00F0 | 6 | UA83 | BY05 | 8 | CL | 1 | 8 | B | 15-16 | SR8 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
35 | BJBAP.C4L8 |  |  | BLECF_ | 246 | 0x | 00F6 | 8 | UA83 | BY05 | 8 | CL | 1 | 7 | B | 13-14 | SR8 | BY05 | C | 2 | 1-2 | BY02 | L | 1 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
36 | BJSAP.B4L8 | yes |  | BLECF_ | 248 | 0x | 00F8 | 1 | UA83 | BY05 | 8 | CL | 1 | 6 | B | 11-12 | SR8 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
37 | BJSAP.A4L8 | yes |  | BLECF_ | 243 | 0x | 00F3 | 2 | UA83 | BY05 | 8 | CL | 1 | 5 | B | 9-10 | SR8 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
38 | BJBAP.A3L8 |  |  | BLECF_ | 250 | 0x | 00FA | 8 | UA83 | BY05 | 8 | CL | 1 | 4 | B | 7-8 | SR8 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
39 | BJBAP.A2L8 |  |  | BLECF_ | 261 | 0x | 0105 | 8 | UA83 | BY05 | 8 | CL | 1 | 3 | B | 5-6 | SR8 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
40 | BJBAP.B1L8 |  |  | BLECF_ | 269 | 0x | 010D | 5 | UA83 | BY05 | 8 | CL | 1 | 2 | B | 3-4 | SR8 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
41 | BJSAP.A1L8 | yes |  | BLECF_ | 180 | 0x | 00B4 | 1 | UA83 | BY05 | 8 | CL | 1 | 1 | B | 1-2 | SR8 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
42 | BJSAP.A1R8 |  |  | BLECF_ | 78 | 0x | 004E | 1 | UA87 | BY06 | 8 | CR | 1 | 1 | B | 1-2 | SR8 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
43 | BJBAP.B1R8 |  |  | BLECF_ | 75 | 0x | 004B | 5 | UA87 | BY06 | 8 | CR | 1 | 2 | B | 3-4 | SR8 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
44 | BJBAP.A2R8 |  |  | BLECF_ | 6 | 0x | 0006 | 8 | UA87 | BY06 | 8 | CR | 1 | 3 | B | 5-6 | SR8 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
45 | BJBAP.A3R8 |  |  | BLECF_ | 528 | 0x | 0210 | 4 | UA87 | BY06 | 8 | CR | 1 | 4 | B | 7-8 | SR8 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
| BJBAP.A4R8 |  |  | BLECF_ | 71 | 0x | 0047 | 7 | UA87 | BY06 | 8 | CR | 1 | 5 | B | 9-10 | SR8 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8
46 | BJBAP.A4R8 | yes | NES18 | BLECF_ | 71 | 0x | 0047 | 7 | UA87 | BY06 | 8 | CR | 1 | 5 | B | 9-10 | SR8 | BY05 | C | 12 | 9-10 | BY01 | I | 4 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
47 | BJSAP.B4R8 | yes |  | BLECF_ | 287 | 0x | 011F | 4 | UA87 | BY06 | 8 | CR | 1 | 6 | B | 11-12 | SR8 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
48 | BJSAP.C4R8 | yes |  | BLECF_ | 278 | 0x | 0116 | 1 | UA87 | BY06 | 8 | CR | 1 | 7 | B | 13-14 | SR8 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
49 | BJBAP.D4R8 | yes |  | BLECF_ | 209 | 0x | 00D1 | 8 | UA87 | BY06 | 8 | CR | 1 | 8 | B | 15-16 | SR8 | BY05 | C | 11 | 3-4 | BY01 | R | 3 | 4 | 1 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
50 | BJBAP.A5R8 |  |  | BLECF_ | 233 | 0x | 00E9 | 6 | UA87 | BY06 | 8 | CR | 1 | 9 | B | 17-18 | SR8 | BY05 | C | 11 | 5-6 | BY01 | R | 3 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
51 | BJBAP.A6R8 |  |  | BLECF_ | 267 | 0x | 010B | 6 | UA87 | BY06 | 8 | CR | 1 | 10 | B | 19-20 | SR8 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16
51 | BJBAP.A6R8 | yes |  | BLECF_ | 267 | 0x | 010B | 6 | UA87 | BY06 | 8 | CR | 1 | 10 | B | 19-20 | SR8 | BY05 | C | 11 | 7-8 | BY01 | I | 4 | 4 | 1 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
52 | BJSAP.B6R8 | yes |  | BLECF_ | 29 | 0x | 001D | 6 | UA87 | BY06 | 8 | CR | 2 | 1 | B | 21-22 | SR8 | BY05 | C | 11 | 9-10 | BY02 | C | 2 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
53 | BJBAP.C6R8 | yes | NES18 | BLECF_ | 27 | 0x | 001B | 6 | UA87 | BY06 | 8 | CR | 2 | 2 | B | 23-24 | SR8 | BY05 | C | 11 | 11-12 | BY01 | R | 3 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
54 | BJBAP.A7R8 | yes | NES18 | BLECF_ | 26 | 0x | 001A | 6 | UA87 | BY06 | 8 | CR | 2 | 3 | C | 1-2 | SR8 | BY05 | C | 10 | 1-2 | BY01 | R | 3 | 5 | 2 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
55 | BJBAP.A8R8 |  |  | BLECF_ | 24 | 0x | 0018 | 6 | UA87 | BY06 | 8 | CR | 2 | 4 | C | 3-4 | SR8 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
| BJBAP.B8R8 |  |  | BLECF_ | 25 | 0x | 0019 | 6 | UA87 | BY06 | 8 | CR | 2 | 5 | C | 5-6 | SR8 | BY05 | C | 10 | 5-6 | BY02 | C | 2 | 18 | 14 | 1-8
56 | BJSAP.B8R8 | yes |  | BLECF_ | 25 | 0x | 0019 | 6 | UA87 | BY06 | 8 | CR | 2 | 5 | C | 5-6 | SR8 | BY05 | C | 10 | 5-6 | BY01 | I | 4 | 5 | 2 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
57 | BJBAP.C8R8 | yes |  | BLECF_ | 127 | 0x | 007F | 6 | UA87 | BY06 | 8 | CR | 2 | 6 | C | 7-8 | SR8 | BY05 | C | 10 | 7-8 | BY01 | R | 3 | 6 | 3 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
58 | BJBAP.A9R8 |  |  | BLECF_ | 80 | 0x | 0050 | 6 | UA87 | BY06 | 8 | CR | 2 | 7 | C | 9-10 | SR8 | BY05 | C | 10 | 9-10 | BY01 | R | 3 | 6 | 3 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
59 | BJBAP.A10R8 |  |  | BLECF_ | 353 | 0x | 0161 | 6 | UA87 | BY06 | 8 | CR | 2 | 8 | C | 11-12 | SR8 | BY05 | C | 10 | 11-12 | BY01 | R | 3 | 7 | 4 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
60 | BJBAP.A11R8 |  |  | BLECF_ | 39 | 0x | 0027 | 6 | UA87 | BY06 | 8 | CR | 2 | 9 | C | 13-14 | SR8 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
61 | BJSAP.B11R8 | yes |  | BLECF_ | 113 | 0x | 0071 | 6 | UA87 | BY06 | 8 | CR | 2 | 10 | C | 15-16 | SR8 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
62 | BJBAP.C11R8 |  | NES18 | BLECF_ | 300 | 0x | 012C | 6 | UA87 | BY06 | 8 | CR | 3 | 1 | C | 17-18 | SR8 | BY05 | C | 9 | 5-6 | BY01 | R | 3 | 7 | 4 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
63 | BJBAP.A12R8 |  |  | BLECF_ | 272 | 0x | 0110 | 6 | Arc | BYPLM.A12R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 8 | 7-8 | BY01 | R | 3 | 8 | 5 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
64 | BJBAP.A13R8 |  |  | BLECF_ | 51 | 0x | 0033 | 6 | Arc | BYPLM.A13R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 9 | 7-8 | BY01 | R | 3 | 8 | 5 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
65 | BJBAP.A14R8 |  |  | BLECF_ | 54 | 0x | 0036 | 6 | Arc | BYPLM.A14R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 10 | 7-8 | BY01 | R | 3 | 9 | 6 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
66 | BJBAP.A15R8 |  |  | BLECF_ | 30 | 0x | 001E | 6 | Arc | BYPLM.A15R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 11 | 7-8 | BY01 | R | 3 | 9 | 6 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
67 | BJBAP.A16R8 |  |  | BLECF_ | 142 | 0x | 008E | 6 | Arc | BYPLM.A16R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | F | 12 | 7-8 | BY01 | R | 3 | 10 | 7 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
68 | BJBAP.A17R8 |  |  | BLECF_ | 61 | 0x | 003D | 6 | Arc | BYPLM.A17R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 7 | 7-8 | BY01 | R | 3 | 10 | 7 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
69 | BJBAP.A18R8 |  |  | BLECF_ | 148 | 0x | 0094 | 6 | Arc | BYPLM.A18R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 8 | 7-8 | BY01 | R | 3 | 11 | 8 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
70 | BJBAP.A19R8 |  |  | BLECF_ | 15 | 0x | 000F | 6 | Arc | BYPLM.A19R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 9 | 7-8 | BY01 | R | 3 | 11 | 8 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
71 | BJBAP.A20R8 |  |  | BLECF_ | 294 | 0x | 0126 | 6 | Arc | BYPLM.A20R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 10 | 7-8 | BY01 | R | 3 | 13 | 9 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
72 | BJBAP.A21R8 |  |  | BLECF_ | 286 | 0x | 011E | 6 | Arc | BYPLM.A21R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 11 | 7-8 | BY01 | R | 3 | 13 | 9 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
73 | BJBAP.A22R8 |  |  | BLECF_ | 206 | 0x | 00CE | 6 | Arc | BYPLM.A22R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | G | 12 | 7-8 | BY01 | R | 3 | 14 | 10 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
74 | BJBAP.A23R8 |  |  | BLECF_ | 270 | 0x | 010E | 6 | Arc | BYPLM.A23R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 7 | 7-8 | BY01 | R | 3 | 14 | 10 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
75 | BJBAP.A24R8 |  |  | BLECF_ | 131 | 0x | 0083 | 6 | Arc | BYPLM.A24R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 8 | 7-8 | BY01 | R | 3 | 15 | 11 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
76 | BJBAP.A25R8 |  |  | BLECF_ | 227 | 0x | 00E3 | 6 | Arc | BYPLM.A25R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 9 | 7-8 | BY01 | R | 3 | 15 | 11 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
77 | BJBAP.A26R8 |  |  | BLECF_ | 253 | 0x | 00FD | 6 | Arc | BYPLM.A26R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 10 | 7-8 | BY01 | R | 3 | 16 | 12 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
78 | BJBAP.A27R8 |  |  | BLECF_ | 62 | 0x | 003E | 6 | Arc | BYPLM.A27R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 11 | 7-8 | BY01 | R | 3 | 16 | 12 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
79 | BJBAP.A28R8 |  |  | BLECF_ | 153 | 0x | 0099 | 6 | Arc | BYPLM.A28R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | H | 12 | 7-8 | BY01 | R | 3 | 17 | 13 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
80 | BJBAP.A29R8 |  |  | BLECF_ | 106 | 0x | 006A | 6 | Arc | BYPLM.A29R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 7 | 7-8 | BY01 | R | 3 | 17 | 13 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
81 | BJBAP.A30R8 |  |  | BLECF_ | 252 | 0x | 00FC | 6 | Arc | BYPLM.A30R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 8 | 7-8 | BY01 | R | 3 | 18 | 14 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
82 | BJBAP.A31R8 |  |  | BLECF_ | 104 | 0x | 0068 | 6 | Arc | BYPLM.A31R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 9 | 7-8 | BY01 | R | 3 | 18 | 14 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
83 | BJBAP.A32R8 |  |  | BLECF_ | 271 | 0x | 010F | 6 | Arc | BYPLM.A32R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 10 | 7-8 | BY01 | R | 3 | 19 | 15 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
84 | BJBAP.A33R8 |  |  | BLECF_ | 225 | 0x | 00E1 | 6 | Arc | BYPLM.A33R8 | 8 | R |  |  |  | 7-8 | SR8 | BY05 | I | 11 | 7-8 | BY01 | R | 3 | 19 | 15 | 9-16 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
85 | BJBAP.A34R8 |  |  | BLECF_ | 297 | 0x | 0129 | 6 | Arc | BYPLM.A34R8 | 8 | R |  |  |  | 7-9 | SR8 | BY05 | I | 12 | 7-8 | BY01 | R | 3 | 20 | 16 | 1-8 | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok | ok
Door
[28 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[28 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[8 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[8 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[8 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[28 empty] | O.K. | No problems observed
Note on Q53: effinger: Fibre 2 : C22
Note on V53: effinger: Fibre 2: C13 10

## Vertical Slice
| Tunnel - position | [13 empty] | Surface position | [18 empty] | LEVEL
| HV box | CFC-Card |  |  |  |  | CFC - Position | [5 empty] | Patch |  | Optical patch position |  |  |  |  | VME-TC position | [5 empty] | ERRA | ERRB | LF A | LF B | ERRC | CID | FID | DAC | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 1024 | HV | Comments
Pos. | position | No. |  | Hex. |  | Ch | Place | Rack | P. | Pos. | Cr. | Slot |  | No. | Bldg. | Crate |  |  | No. | Rack | VME | Crate | Slot | Card | Dec.
1 | BJBHT.A33L2 | BLECF_ | 399 | 0x | 018F | 6 | Arc | BYPLM.A33L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 5 | 7-8 | BY02 | L | 1 | 19 | 15 | 9-16
2 | BJBHT.A32L2 | BLECF_ | 348 | 0x | 015C | 6 | Arc | BYPLM.A32L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 4 | 7-8 | BY02 | L | 1 | 19 | 15 | 1-8
3 | BJBHT.A31L2 | BLECF_ | 497 | 0x | 01F1 | 6 | Arc | BYPLM.A31L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 3 | 7-8 | BY02 | L | 1 | 18 | 14 | 9-16
4 | BJBHT.A30L2 | BLECF_ | 432 | 0x | 01B0 | 6 | Arc | BYPLM.A30L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 2 | 7-8 | BY02 | L | 1 | 18 | 14 | 1-8
5 | BJBHT.A29L2 | BLECF_ | 520 | 0x | 0208 | 6 | Arc | BYPLM.A29L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | I | 1 | 7-8 | BY02 | L | 1 | 17 | 13 | 9-16
6 | BJBHT.A28L2 | BLECF_ | 439 | 0x | 01B7 | 6 | Arc | BYPLM.A28L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 6 | 7-8 | BY02 | L | 1 | 17 | 13 | 1-8
7 | BJBHT.A27L2 | BLECF_ | 342 | 0x | 0156 | 6 | Arc | BYPLM.A27L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 5 | 7-8 | BY02 | L | 1 | 16 | 12 | 9-16
8 | BJBHT.A26L2 | BLECF_ | 460 | 0x | 01CC | 6 | Arc | BYPLM.A26L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 4 | 7-8 | BY02 | L | 1 | 16 | 12 | 1-8
9 | BJBHT.A25L2 | BLECF_ | 388 | 0x | 0184 | 6 | Arc | BYPLM.A25L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 3 | 7-8 | BY02 | L | 1 | 15 | 11 | 9-16
10 | BJBHT.A24L2 | BLECF_ | 440 | 0x | 01B8 | 6 | Arc | BYPLM.A24L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 2 | 7-8 | BY02 | L | 1 | 15 | 11 | 1-8
11 | BJBHT.A23L2 | BLECF_ | 468 | 0x | 01D4 | 6 | Arc | BYPLM.A23L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | H | 1 | 7-8 | BY02 | L | 1 | 14 | 10 | 9-16
12 | BJBHT.A22L2 | BLECF_ | 456 | 0x | 01C8 | 6 | Arc | BYPLM.A22L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 6 | 7-8 | BY02 | L | 1 | 14 | 10 | 1-8
13 | BJBHT.A21L2 | BLECF_ | 111 | 0x | 006F | 6 | Arc | BYPLM.A21L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 5 | 7-8 | BY02 | L | 1 | 13 | 9 | 9-16
14 | BJBHT.A20L2 | BLECF_ | 11 | 0x | 000B | 6 | Arc | BYPLM.A20L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 4 | 7-8 | BY02 | L | 1 | 13 | 9 | 1-8
15 | BJBHT.A19L2 | BLECF_ | 715 | 0x | 02CB | 6 | Arc | BYPLM.A19L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 3 | 7-8 | BY02 | L | 1 | 11 | 8 | 9-16
16 | BJBHT.A18L2 | BLECF_ | 366 | 0x | 016E | 6 | Arc | BYPLM.A18L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 2 | 7-8 | BY02 | L | 1 | 11 | 8 | 1-8
17 | BJBHT.A17L2 | BLECF_ | 427 | 0x | 01AB | 6 | Arc | BYPLM.A17L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | G | 1 | 7-8 | BY02 | L | 1 | 10 | 7 | 9-16
18 | BJBHT.A16L2 | BLECF_ | 385 | 0x | 0181 | 6 | Arc | BYPLM.A16L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 5 | 7-8 | BY02 | L | 1 | 10 | 7 | 1-8
19 | BJBHT.A15L2 | BLECF_ | 444 | 0x | 01BC | 6 | Arc | BYPLM.A15L2 | 2 | L |  |  |  | 5-6 | SR2 | BY05 | F | 4 | 5-6 | BY02 | L | 1 | 9 | 6 | 9-16
20 | BJBHT.A14L2 | BLECF_ | 380 | 0x | 017C | 6 | Arc | BYPLM.A14L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 3 | 7-8 | BY02 | L | 1 | 9 | 6 | 1-8
21 | BJBHT.A13L2 | BLECF_ | 347 | 0x | 015B | 6 | Arc | BYPLM.A13L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 2 | 7-8 | BY02 | L | 1 | 8 | 5 | 9-16
22 | BJBHT.A12L2 | BLECF_ | 200 | 0x | 00C8 | 8 | Arc | BYPLM.A12L2 | 2 | L |  |  |  | 7-8 | SR2 | BY05 | F | 1 | 7-8 | BY02 | L | 1 | 8 | 5 | 1-8
23 | BJBHT.C11L2 | BLECF_ | 583 | 0x | 0247 | 6 | UA23 | BY02 | 2 | CL | 3 | 2 | C | 19-20 | SR2 | BY05 | C | 4 | 7-8 | BY02 | L | 1 | 7 | 4 | 9-16
24 | BJBHT.B11L2 | BLECF_ | 343 | 0x | 0157 | 6 | UA23 | BY02 | 2 | CL | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 4 | 5-6 | BY02 | C | 2 | 10 | 7 | 9-16
25 | BJBHT.A11L2 | BLECF_ | 378 | 0x | 017A | 6 | UA23 | BY02 | 2 | CL | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 4 | 3-4 | BY02 | C | 2 | 10 | 7 | 1-8
26 | BJBHT.B10L2 | BLECF_ | 134 | 0x | 0086 | 6 | UA23 | BY02 | 2 | CL | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 4 | 1-2 | BY02 | L | 1 | 7 | 4 | 1-8
27 | BJBHT.A10L2 | BLECF_ | 50 | 0x | 0032 | 7 | UA23 | BY02 | 2 | CL | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 3 | 11-12 | BY02 | C | 2 | 9 | 6 | 9-16
28 | BJBHT.A9L2 | BLECF_ | 517 | 0x | 0205 | 6 | UA23 | BY02 | 2 | CL | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 3 | 9-10 | BY02 | L | 1 | 6 | 3 | 9-16
29 | BJBHT.C8L2 | BLECF_ | 357 | 0x | 0165 | 6 | UA23 | BY02 | 2 | CL | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 3 | 7-8 | BY02 | L | 1 | 6 | 3 | 1-8
30 | BJBHT.B8L2 | BLECF_ | 333 | 0x | 014D | 6 | UA23 | BY02 | 2 | CL | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 3 | 5-6 | BY02 | C | 2 | 9 | 6 | 1-8
31 | BJBHT.A8L2 | BLECF_ | 491 | 0x | 01EB | 6 | UA23 | BY02 | 2 | CL | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 3 | 3-4 | BY02 | C | 2 | 8 | 5 | 9-16
32 | BJBHT.A7L2 | BLECF_ | 362 | 0x | 016A | 6 | UA23 | BY02 | 2 | CL | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 3 | 1-2 | BY02 | L | 1 | 5 | 2 | 9-16
33 | BJBHT.C6L2 | BLECF_ | 189 | 0x | 00BD | 6 | UA23 | BY02 | 2 | CL | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 2 | 11-12 | BY02 | L | 1 | 5 | 2 | 1-8
34 | BJBHT.B6L2 | BLECF_ | 344 | 0x | 0158 | 6 | UA23 | BY02 | 2 | CL | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 2 | 9-10 | BY02 | C | 2 | 8 | 5 | 1-8
35 | BJBHT.A6L2 | BLECF_ | 474 | 0x | 01DA | 6 | UA23 | BY02 | 2 | CL | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 2 | 7-8 | BY02 | C | 2 | 7 | 4 | 9-16
36 | BJBHT.A5L2 | BLECF_ | 328 | 0x | 0148 | 6 | UA23 | BY02 | 2 | CL | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 2 | 5-6 | BY02 | L | 1 | 4 | 1 | 9-16
37 | BJBHT.D4L2 | BLECF_ | 371 | 0x | 0173 | 8 | UA23 | BY02 | 2 | CL | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 2 | 3-4 | BY02 | L | 1 | 4 | 1 | 1-8
X | BJBHT.TEST | BLECF_ | 208 | 0x | 00D0 | 2 | UA23 | BY02 | 2 | CL | 3 | 10 | C | 23-24 | SR2 | BY05 | C | 4 | 11-12 | BY02 | V | 1 | 6 | 3 | 1-8 | [18 empty] | dismounted
38 | BJBHT.C4L2 | BLECF_ | 181 | 0x | 00B5 | 1 | UA23 | BY02 | 2 | CL | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 2 | 1-2 | BY02 | C | 2 | 7 | 4 | 1-8
39 | BJBHT.B4L2 | BLECF_ | 320 | 0x | 0140 | 4 | UA23 | BY02 | 2 | CL | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 1 | 11-12 | BY02 | C | 2 | 6 | 3 | 9-16
40 | BJBHT.A4L2 | BLECF_ | 516 | 0x | 0204 | 7 | UA23 | BY02 | 2 | CL | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 1 | 9-10 | BY02 | C | 2 | 6 | 3 | 1-8
41 | BJBHT.A3L2 | BLECF_ | 324 | 0x | 0144 | 4 | UA23 | BY02 | 2 | CL | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 1 | 7-8 | BY02 | C | 2 | 5 | 2 | 9-16
42 | BJBHT.A2L2 | BLECF_ | 188 | 0x | 00BC | 8 | UA23 | BY02 | 2 | CL | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 1 | 5-6 | BY02 | C | 2 | 5 | 2 | 1-8
43 | BJBHT.B1L2 | BLECF_ | 510 | 0x | 01FE | 5 | UA23 | BY02 | 2 | CL | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 1 | 3-4 | BY02 | C | 2 | 4 | 1 | 9-16
44 | BJBHT.A1L2 | BLECF_ | 476 | 0x | 01DC | 1 | UA23 | BY02 | 2 | CL | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 1 | 1-2 | BY02 | C | 2 | 4 | 1 | 1-8
45 | BJBHT.A1R2 | BLECF_ | 706 | 0x | 02C2 | 1 | UA27 | BY03 | 2 | CR | 1 | 1 | B | 1-2 | SR2 | BY05 | C | 12 | 1-2 | BY02 | C | 2 | 13 | 9 | 1-8
46 | BJBHT.B1R2 | BLECF_ | 606 | 0x | 025E | 5 | UA27 | BY03 | 2 | CR | 1 | 2 | B | 3-4 | SR2 | BY05 | C | 12 | 3-4 | BY02 | C | 2 | 13 | 9 | 9-16
47 | BJBHT.A2R2 | BLECF_ | 675 | 0x | 02A3 | 8 | UA27 | BY03 | 2 | CR | 1 | 3 | B | 5-6 | SR2 | BY05 | C | 12 | 5-6 | BY02 | C | 2 | 14 | 10 | 1-8
48 | BJBHT.A3R2 | BLECF_ | 641 | 0x | 0281 | 4 | UA27 | BY03 | 2 | CR | 1 | 4 | B | 7-8 | SR2 | BY05 | C | 12 | 7-8 | BY02 | C | 2 | 14 | 10 | 9-16
49 | BJBHT.A4R2 | BLECF_ | 719 | 0x | 02CF | 5 | UA27 | BY03 | 2 | CR | 1 | 5 | B | 9-10 | SR2 | BY05 | C | 12 | 9-10 | BY02 | C | 2 | 15 | 11 | 1-8
50 | BJBHT.B4R2 | BLECF_ | 738 | 0x | 02E2 | 2 | UA27 | BY03 | 2 | CR | 1 | 6 | B | 11-12 | SR2 | BY05 | C | 12 | 11-12 | BY02 | C | 2 | 15 | 11 | 9-16
51 | BJBHT.C4R2 | BLECF_ | 652 | 0x | 028C | 1 | UA27 | BY03 | 2 | CR | 1 | 7 | B | 13-14 | SR2 | BY05 | C | 11 | 1-2 | BY02 | C | 2 | 16 | 12 | 1-8
52 | BJBHT.D4R2 | BLECF_ | 665 | 0x | 0299 | 8 | UA27 | BY03 | 2 | CR | 1 | 8 | B | 15-16 | SR2 | BY05 | C | 11 | 3-4 | BY02 | R | 3 | 4 | 1 | 1-8
53 | BJBHT.A5R2 | BLECF_ | 740 | 0x | 02E4 | 6 | UA27 | BY03 | 2 | CR | 1 | 9 | B | 17-18 | SR2 | BY05 | C | 11 | 5-6 | BY02 | R | 3 | 4 | 1 | 9-16
54 | BJBHT.A6R2 | BLECF_ | 666 | 0x | 029A | 1 | UA27 | BY03 | 2 | CR | 1 | 10 | B | 19-20 | SR2 | BY05 | C | 11 | 7-8 | BY02 | C | 2 | 16 | 12 | 9-16
55 | BJBHT.B6R2 | BLECF_ | 722 | 0x | 02D2 | 7 | UA27 | BY03 | 2 | CR | 2 | 1 | B | 21-22 | SR2 | BY05 | C | 11 | 9-10 | BY02 | R | 3 | 5 | 2 | 1-8
56 | BJBHT.A7R2 | BLECF_ | 658 | 0x | 0292 | 6 | UA27 | BY03 | 2 | CR | 2 | 2 | B | 23-24 | SR2 | BY05 | C | 11 | 11-12 | BY02 | R | 3 | 5 | 2 | 9-16
57 | BJBHT.A8R2 | BLECF_ | 626 | 0x | 0272 | 6 | UA27 | BY03 | 2 | CR | 2 | 3 | C | 1-2 | SR2 | BY05 | C | 10 | 1-2 | BY02 | C | 2 | 17 | 13 | 1-8
58 | BJBHT.B8R2 | BLECF_ | 406 | 0x | 0196 | 6 | UA27 | BY03 | 2 | CR | 2 | 4 | C | 3-4 | SR2 | BY05 | C | 10 | 3-4 | BY02 | C | 2 | 17 | 13 | 9-16
59 | BJBHT.C8R2 | BLECF_ | 4 | 0x | 0004 | 6 | UA27 | BY03 | 2 | CR | 2 | 5 | C | 5-6 | SR2 | BY05 | C | 10 | 5-6 | BY02 | R | 3 | 6 | 3 | 1-8
60 | BJBHT.A9R2 | BLECF_ | 279 | 0x | 0117 | 6 | UA27 | BY03 | 2 | CR | 2 | 6 | C | 7-8 | SR2 | BY05 | C | 10 | 7-8 | BY02 | R | 3 | 6 | 3 | 9-16
61 | BJBHT.A10R2 | BLECF_ | 97 | 0x | 0061 | 6 | UA27 | BY03 | 2 | CR | 2 | 7 | C | 9-10 | SR2 | BY05 | C | 10 | 9-10 | BY02 | C | 2 | 18 | 14 | 1-8
62 | BJBHT.B10R2 | BLECF_ | 721 | 0x | 02D1 | 6 | UA27 | BY03 | 2 | CR | 2 | 8 | C | 11-12 | SR2 | BY05 | C | 10 | 11-12 | BY02 | R | 3 | 7 | 4 | 1-8
63 | BJBHT.A11R2 | BLECF_ | 128 | 0x | 0080 | 6 | UA27 | BY03 | 2 | CR | 2 | 9 | C | 13-14 | SR2 | BY05 | C | 9 | 1-2 | BY02 | C | 2 | 18 | 14 | 9-16
64 | BJBHT.B11R2 | BLECF_ | 627 | 0x | 0273 | 6 | UA27 | BY03 | 2 | CR | 2 | 10 | C | 15-16 | SR2 | BY05 | C | 9 | 3-4 | BY02 | C | 2 | 19 | 15 | 1-8
65 | BJBHT.C11R2 | BLECF_ | 689 | 0x | 02B1 | 6 | UA27 | BY03 | 2 | CR | 3 | 1 | C | 17-18 | SR2 | BY05 | C | 9 | 5-6 | BY02 | R | 3 | 7 | 4 | 9-16
66 | BJBHT.A12R2 | BLECF_ | 681 | 0x | 02A9 | 8 | Arc | BYPLM.A12R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 8 | 7-8 | BY02 | R | 3 | 8 | 5 | 1-8
67 | BJBHT.A13R2 | BLECF_ | 687 | 0x | 02AF | 6 | Arc | BYPLM.A13R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 9 | 7-8 | BY02 | R | 3 | 8 | 5 | 9-16
68 | BJBHT.A14R2 | BLECF_ | 684 | 0x | 02AC | 6 | Arc | BYPLM.A14R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 10 | 7-8 | BY02 | R | 3 | 9 | 6 | 1-8
69 | BJBHT.A15R2 | BLECF_ | 484 | 0x | 01E4 | 6 | Arc | BYPLM.A15R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 11 | 7-8 | BY02 | R | 3 | 9 | 6 | 9-16
70 | BJBHT.A16R2 | BLECF_ | 686 | 0x | 02AE | 6 | Arc | BYPLM.A16R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | F | 12 | 7-8 | BY02 | R | 3 | 10 | 7 | 1-8
71 | BJBHT.A17R2 | BLECF_ | 143 | 0x | 008F | 6 | Arc | BYPLM.A17R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 7 | 7-8 | BY02 | R | 3 | 10 | 7 | 9-16
72 | BJBHT.A18R2 | BLECF_ | 463 | 0x | 01CF | 6 | Arc | BYPLM.A18R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 8 | 7-8 | BY02 | R | 3 | 11 | 8 | 1-8
73 | BJBHT.A19R2 | BLECF_ | 379 | 0x | 017B | 6 | Arc | BYPLM.A19R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 9 | 7-8 | BY02 | R | 3 | 11 | 8 | 9-16
74 | BJBHT.A20R2 | BLECF_ | 382 | 0x | 017E | 6 | Arc | BYPLM.A20R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 10 | 7-8 | BY02 | R | 3 | 13 | 9 | 1-8
75 | BJBHT.A21R2 | BLECF_ | 451 | 0x | 01C3 | 6 | Arc | BYPLM.A21R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 11 | 7-8 | BY02 | R | 3 | 13 | 9 | 9-16
76 | BJBHT.A22R2 | BLECF_ | 404 | 0x | 0194 | 6 | Arc | BYPLM.A22R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | G | 12 | 7-8 | BY02 | R | 3 | 14 | 10 | 1-8
77 | BJBHT.A23R2 | BLECF_ | 84 | 0x | 0054 | 6 | Arc | BYPLM.A23R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 7 | 7-8 | BY02 | R | 3 | 14 | 10 | 9-16
78 | BJBHT.A24R2 | BLECF_ | 418 | 0x | 01A2 | 6 | Arc | BYPLM.A24R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 8 | 7-8 | BY02 | R | 3 | 15 | 11 | 1-8
79 | BJBHT.A25R2 | BLECF_ | 390 | 0x | 0186 | 6 | Arc | BYPLM.A25R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 9 | 7-8 | BY02 | R | 3 | 15 | 11 | 9-16
80 | BJBHT.A26R2 | BLECF_ | 214 | 0x | 00D6 | 6 | Arc | BYPLM.A26R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 10 | 7-8 | BY02 | R | 3 | 16 | 12 | 1-8
81 | BJBHT.A27R2 | BLECF_ | 694 | 0x | 02B6 | 6 | Arc | BYPLM.A27R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 11 | 7-8 | BY02 | R | 3 | 16 | 12 | 9-16
82 | BJBHT.A28R2 | BLECF_ | 685 | 0x | 02AD | 6 | Arc | BYPLM.A28R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | H | 12 | 7-8 | BY02 | R | 3 | 17 | 13 | 1-8
83 | BJBHT.A29R2 | BLECF_ | 725 | 0x | 02D5 | 6 | Arc | BYPLM.A29R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 7 | 7-8 | BY02 | R | 3 | 17 | 13 | 9-16
84 | BJBHT.A30R2 | BLECF_ | 195 | 0x | 00C3 | 6 | Arc | BYPLM.A30R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 8 | 7-8 | BY02 | R | 3 | 18 | 14 | 1-8
85 | BJBHT.A31R2 | BLECF_ | 690 | 0x | 02B2 | 6 | Arc | BYPLM.A31R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 9 | 7-8 | BY02 | R | 3 | 18 | 14 | 9-16
86 | BJBHT.A32R2 | BLECF_ | 360 | 0x | 0168 | 6 | Arc | BYPLM.A32R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 10 | 7-8 | BY02 | R | 3 | 19 | 15 | 1-8
87 | BJBHT.A33R2 | BLECF_ | 23 | 0x | 0017 | 6 | Arc | BYPLM.A33R2 | 2 | R |  |  |  | 7-8 | SR2 | BY05 | I | 11 | 7-8 | BY02 | R | 3 | 19 | 15 | 9-16
88 | BJBHT.A34R2 | BLECF_ | 724 | 0x | 02D4 | 6 | Arc | BYPLM.A34R2 | 2 | R |  |  |  | 7-9 | SR2 | BY05 | I | 12 | 7-8 | BY02 | R | 3 | 20 | 16 | 1-8
Door
[26 empty] | ↑↑↑ | CRC Error consantly fast increasing | [8 empty] | Level on the channel o.k
[26 empty] | ↑↑ | CRC Error sometimes fast increasing | [8 empty] | Problem on the channel (e.g. negative offset current)
[6 empty] | Card number correct | [19 empty] | ↑ | CRC Errors slowly increasing
[6 empty] | Card number needs to be checked | [19 empty] | 1 | Number of CRC Errors (almost stable) | [6 empty] | 0 | Adjustment of 10pA potentiometer needed
[6 empty] | Card number wrong, number in the tunnel or fiber connection needed to be checked | [19 empty] | LF | Lost Frame | [6 empty] | FF | Problem with one channel (e.g. negative offset current)
[26 empty] | O.K. | No problems observed
Note on V23: effinger: Fibre 7 & 8 broken, spare fibre used
Note on H82: effinger: Checked-----> BLECF_0382 installed! Wrong number 0x017E (382) instead of 0x0188 (392) ; to be check in the tunnel; Card number checked: --> BLECF_0382 installed; sent number 382

## Overview
267 | BLECF_ | 457
551 | BLECF_ | #REF!
578 | BLECF_ | 757
435 | BLECF_ | 674
574 | BLECF_ | 753
596 | #REF! | 750
594 | BLECF_ | 749
492 | BLECF_ | 748
519 | BLECF_ | 744
502 | BLECF_ | 743
619 | BLECF_ | 742
422 | BLECF_ | 741
622 | BLECF_ | 739
615 | BLECF_ | 737
625 | BLECF_ | 736
632 | BLECF_ | 735
628 | BLECF_ | 50
616 | BLECF_ | 733
618 | BLECF_ | 732
633 | BLECF_ | 731
449 | BLECF_ | 729
613 | BLECF_ | 728
412 | BLECF_ | 727
597 | BLECF_ | 726
404 | BLECF_ | 725
409 | BLECF_ | 724
418 | BLECF_ | 717
448 | BLECF_ | 711
599 | BLECF_ | 710
64 | BLECF_ | 708
413 | BLECF_ | 705
630 | BLECF_ | 704
431 | BLECF_ | 703
415 | BLECF_ | 702
629 | BLECF_ | 701
612 | BLECF_ | 699
614 | BLECF_ | 698
601 | BLECF_ | 697
620 | BLECF_ | 695
402 | BLECF_ | 694
631 | BLECF_ | 693
623 | BLECF_ | 691
406 | BLECF_ | 690
621 | BLECF_ | 688
388 | BLECF_ | 687
391 | BLECF_ | 686
403 | BLECF_ | 685
389 | BLECF_ | 684
626 | BLECF_ | 683
651 | BLECF_ | 682
387 | BLECF_ | 681
559 | BLECF_ | 644
636 | #REF! | 679
638 | #REF! | 678
451 | BLECF_ | 677
285 | BLECF_ | 676
434 | BLECF_ | 574
609 | #REF! | 673
437 | BLECF_ | 571
595 | #REF! | 671
510 | BLECF_ | 670
426 | BLECF_ | 668
262 | BLECF_ | 667
444 | BLECF_ | 175
503 | BLECF_ | 664
598 | BLECF_ | 663
494 | BLECF_ | 662
634 | BLECF_ | 661
606 | BLECF_ | 660
512 | BLECF_ | 659
432 | BLECF_ | 657
603 | BLECF_ | 656
500 | #REF! | 655
433 | BLECF_ | 654
286 | BLECF_ | 653
516 | BLECF_ | 651
554 | BLECF_ | 650
640 | BLECF_ | 649
602 | #REF! | 648
294 | BLECF_ | 647
414 | BLECF_ | 646
421 | BLECF_ | 645
533 | BLECF_ | 134
513 | BLECF_ | 643
509 | BLECF_ | 642
572 | BLECF_ | 640
280 | BLECF_ | 251
440 | BLECF_ | 637
284 | BLECF_ | 636
279 | BLECF_ | 635
617 | BLECF_ | 634
661 | BLECF_ | 633
582 | BLECF_ | 632
501 | BLECF_ | 631
425 | BLECF_ | 720
442 | BLECF_ | 629
563 | BLECF_ | 625
417 | BLECF_ | 623
428 | BLECF_ | 622
610 | BLECF_ | 621
424 | BLECF_ | 620
410 | BLECF_ | 619
604 | BLECF_ | 383
624 | BLECF_ | 617
493 | BLECF_ | 616
507 | BLECF_ | 615
423 | BLECF_ | 614
499 | BLECF_ | 613
1 | BLECF_ | 612
469 | BLECF_ | 527
2 | BLECF_ | 610
436 | BLECF_ | 609
416 | BLECF_ | 607
608 | #REF! | 605
427 | BLECF_ | 604
611 | BLECF_ | 603
447 | BLECF_ | 602
411 | BLECF_ | 601
430 | BLECF_ | 600
511 | BLECF_ | 599
600 | BLECF_ | 598
273 | BLECF_ | 597
605 | BLECF_ | 596
276 | BLECF_ | 401
263 | BLECF_ | 594
277 | BLECF_ | 593
278 | BLECF_ | 592
508 | BLECF_ | 591
656 | BLECF_ | 590
565 | BLECF_ | 589
268 | BLECF_ | 588
644 | BLECF_ | 587
591 | BLECF_ | 586
537 | BLECF_ | 585
524 | BLECF_ | 584
549 | BLECF_ | 582
575 | BLECF_ | 581
120 | BLECF_ | 580
592 | BLECF_ | 579
514 | BLECF_ | 578
445 | BLECF_ | 577
439 | BLECF_ | 576
446 | BLECF_ | 575
438 | BLECF_ | 103
497 | BLECF_ | 572
3 | BLECF_ | 208
301 | BLECF_ | 570
4 | BLECF_ | 569
531 | BLECF_ | 568
297 | BLECF_ | 567
5 | BLECF_ | 566
6 | BLECF_ | 853
7 | BLECF_ | 680
8 | BLECF_ | 563
466 | BLECF_ | 562
9 | BLECF_ | 218
648 | BLECF_ | 638
639 | BLECF_ | 559
535 | BLECF_ | 299
10 | BLECF_ | 557
523 | BLECF_ | 235
11 | BLECF_ | 555
650 | BLECF_ | 360
505 | BLECF_ | 552
544 | #REF! | 551
265 | BLECF_ | 550
12 | BLECF_ | 91
13 | BLECF_ | 548
543 | BLECF_ | 547
566 | BLECF_ | 546
14 | BLECF_ | 565
637 | BLECF_ | 544
450 | #REF! | 543
15 | BLECF_ | 542
456 | #REF! | 541
16 | BLECF_ | 540
299 | BLECF_ | 539
17 | BLECF_ | 356
540 | BLECF_ | 537
443 | BLECF_ | 536
18 | BLECF_ | 534
19 | BLECF_ | 533
20 | BLECF_ | 202
530 | BLECF_ | 531
272 | BLECF_ | 530
21 | BLECF_ | 611
581 | BLECF_ | 525
476 | BLECF_ | 523
308 | BLECF_ | 522
585 | BLECF_ | 521
326 | BLECF_ | 520
483 | BLECF_ | 519
266 | BLECF_ | 518
349 | BLECF_ | 517
361 | BLECF_ | 516
305 | BLECF_ | 515
472 | BLECF_ | 514
642 | BLECF_ | 513
292 | BLECF_ | 512
23 | BLECF_ | 511
364 | BLECF_ | 510
525 | BLECF_ | 509
24 | BLECF_ | 508
564 | BLECF_ | 506
686 | BLECF_ | 505
470 | BLECF_ | 504
461 | #REF! | 502
283 | BLECF_ | 501
312 | BLECF_ | 500
25 | BLECF_ | #REF!
471 | BLECF_ | 499
324 | BLECF_ | 497
474 | BLECF_ | 496
318 | BLECF_ | 495
344 | BLECF_ | 583
473 | BLECF_ | 493
460 | BLECF_ | 709
352 | BLECF_ | 491
627 | BLECF_ | 490
485 | BLECF_ | 489
487 | BLECF_ | 488
454 | BLECF_ | 487
26 | BLECF_ | 486
479 | BLECF_ | 485
390 | BLECF_ | 484
680 | BLECF_ | 482
306 | BLECF_ | 481
465 | BLECF_ | 186
313 | BLECF_ | 479
314 | BLECF_ | 478
573 | BLECF_ | 477
365 | BLECF_ | 476
568 | BLECF_ | 475
356 | BLECF_ | 474
567 | BLECF_ | 473
468 | BLECF_ | 68
504 | BLECF_ | 471
683 | BLECF_ | 470
635 | BLECF_ | 469
332 | BLECF_ | 468
27 | BLECF_ | 467
29 | BLECF_ | 466
652 | #REF! | 265
681 | BLECF_ | 464
393 | BLECF_ | 463
641 | BLECF_ | 462
28 | BLECF_ | 461
329 | BLECF_ | 460
467 | BLECF_ | 459
333 | BLECF_ | 456
664 | BLECF_ | 455
30 | BLECF_ | 454
489 | BLECF_ | 453
491 | BLECF_ | 452
396 | BLECF_ | 451
31 | BLECF_ | 449
484 | BLECF_ | 448
557 | BLECF_ | 447
669 | BLECF_ | 446
659 | BLECF_ | 445
340 | BLECF_ | 444
663 | BLECF_ | 443
289 | BLECF_ | 442
690 | BLECF_ | 441
331 | BLECF_ | 440
327 | BLECF_ | 439
542 | BLECF_ | 438
288 | BLECF_ | 437
32 | BLECF_ | 436
546 | #REF! | 121
647 | BLECF_ | 434
679 | BLECF_ | 433
325 | BLECF_ | 432
666 | #REF! | 431
33 | BLECF_ | 334
34 | BLECF_ | 429
689 | BLECF_ | 428
338 | BLECF_ | 427
35 | BLECF_ | 426
36 | BLECF_ | #REF!
560 | BLECF_ | 424
310 | BLECF_ | 423
684 | BLECF_ | 700
670 | BLECF_ | 421
667 | BLECF_ | 420
561 | BLECF_ | 266
399 | BLECF_ | 418
320 | BLECF_ | 417
671 | BLECF_ | 415
555 | BLECF_ | 414
37 | BLECF_ | 413
668 | BLECF_ | 713
643 | BLECF_ | 411
464 | BLECF_ | 410
264 | BLECF_ | 409
515 | BLECF_ | 408
488 | BLECF_ | 407
541 | BLECF_ | 405
397 | BLECF_ | 404
534 | BLECF_ | 403
38 | BLECF_ | #REF!
486 | BLECF_ | 492
39 | BLECF_ | #REF!
322 | BLECF_ | 399
274 | BLECF_ | 398
40 | BLECF_ | 397
506 | BLECF_ | 395
677 | BLECF_ | 31
455 | BLECF_ | 393
41 | BLECF_ | 392
395 | BLECF_ | 382
42 | BLECF_ | 391
400 | BLECF_ | 390
550 | BLECF_ | 389
330 | BLECF_ | 388
478 | BLECF_ | 387
517 | BLECF_ | 386
339 | BLECF_ | 385
319 | BLECF_ | 384
290 | BLECF_ | 355
688 | BLECF_ | 381
341 | BLECF_ | 380
394 | BLECF_ | 379
346 | BLECF_ | 378
281 | BLECF_ | 377
287 | BLECF_ | 376
463 | BLECF_ | 375
275 | BLECF_ | 374
309 | BLECF_ | 373
658 | BLECF_ | 372
358 | BLECF_ | 371
587 | BLECF_ | 370
660 | BLECF_ | 369
672 | BLECF_ | 368
459 | BLECF_ | 367
337 | BLECF_ | 366
22 | BLECF_ | 234
348 | BLECF_ | 363
353 | BLECF_ | 362
300 | BLECF_ | 361
407 | BLECF_ | 82
682 | BLECF_ | 359
685 | BLECF_ | 358
350 | BLECF_ | 357
586 | #REF! | 18
649 | BLECF_ | 364
311 | BLECF_ | 352
665 | BLECF_ | 351
675 | BLECF_ | 228
298 | BLECF_ | 349
323 | BLECF_ | 348
342 | BLECF_ | 347
570 | BLECF_ | 346
687 | BLECF_ | 345
355 | BLECF_ | 344
345 | BLECF_ | 343
328 | BLECF_ | 342
457 | BLECF_ | 341
539 | BLECF_ | 340
458 | BLECF_ | 339
296 | BLECF_ | 338
482 | BLECF_ | 337
43 | BLECF_ | 336
552 | BLECF_ | 335
673 | BLECF_ | 192
351 | BLECF_ | 333
282 | BLECF_ | 331
553 | BLECF_ | 330
317 | BLECF_ | 329
357 | BLECF_ | 328
453 | BLECF_ | 326
293 | BLECF_ | 325
362 | BLECF_ | 324
291 | BLECF_ | 323
269 | BLECF_ | 322
321 | BLECF_ | 321
360 | BLECF_ | 320
419 | BLECF_ | 319
498 | BLECF_ | 21
307 | BLECF_ | 81
522 | BLECF_ | 316
538 | BLECF_ | 315
475 | BLECF_ | 314
584 | BLECF_ | 313
590 | BLECF_ | 288
303 | BLECF_ | 311
316 | BLECF_ | 310
480 | BLECF_ | 309
662 | BLECF_ | 308
556 | BLECF_ | 307
302 | BLECF_ | 306
429 | BLECF_ | 305
315 | BLECF_ | 317
304 | BLECF_ | 303
133 | BLECF_ | 125
44 | BLECF_ | 300
45 | BLECF_ | 230
46 | BLECF_ | #REF!
47 | BLECF_ | 297
48 | BLECF_ | 295
49 | BLECF_ | 294
50 | BLECF_ | 293
51 | BLECF_ | 292
52 | BLECF_ | 291
53 | BLECF_ | #REF!
54 | BLECF_ | 219
55 | BLECF_ | #REF!
56 | BLECF_ | 287
57 | BLECF_ | 286
58 | BLECF_ | 285
59 | BLECF_ | 284
60 | BLECF_ | 212
61 | BLECF_ | 282
62 | BLECF_ | 281
63 | BLECF_ | 280
65 | BLECF_ | 278
66 | BLECF_ | 277
67 | BLECF_ | 276
68 | BLECF_ | 275
69 | BLECF_ | 274
655 | #REF! | 273
70 | BLECF_ | 272
71 | BLECF_ | 271
72 | BLECF_ | 270
73 | BLECF_ | 269
74 | BLECF_ | 268
75 | BLECF_ | 267
76 | BLECF_ | 848
77 | BLECF_ | 180
78 | BLECF_ | 264
79 | BLECF_ | 263
80 | BLECF_ | 262
81 | BLECF_ | 261
82 | BLECF_ | #REF!
83 | BLECF_ | #REF!
84 | BLECF_ | 257
85 | BLECF_ | 256
86 | BLECF_ | 255
87 | BLECF_ | 253
88 | BLECF_ | 252
89 | BLECF_ | #REF!
90 | BLECF_ | 250
91 | BLECF_ | 249
92 | BLECF_ | 248
93 | BLECF_ | 247
94 | BLECF_ | 246
95 | BLECF_ | #REF!
96 | BLECF_ | #REF!
97 | BLECF_ | 243
98 | BLECF_ | #REF!
99 | BLECF_ | 241
100 | BLECF_ | 240
101 | BLECF_ | 239
102 | BLECF_ | 238
103 | BLECF_ | 237
104 | BLECF_ | 236
105 | BLECF_ | #REF!
106 | BLECF_ | 233
107 | BLECF_ | 232
108 | BLECF_ | 231
109 | BLECF_ | 730
110 | BLECF_ | 229
111 | BLECF_ | 528
112 | BLECF_ | 227
558 | BLECF_ | 226
113 | BLECF_ | 225
114 | BLECF_ | 224
115 | BLECF_ | 223
116 | BLECF_ | 743
117 | BLECF_ | 220
118 | BLECF_ | 289
119 | BLECF_ | 738
529 | BLECF_ | 217
121 | BLECF_ | 216
122 | BLECF_ | 215
401 | BLECF_ | 214
123 | BLECF_ | 213
441 | BLECF_ | 211
124 | BLECF_ | 210
125 | BLECF_ | 209
126 | BLECF_ | 62
127 | BLECF_ | 207
128 | BLECF_ | 206
129 | BLECF_ | 205
130 | BLECF_ | 204
131 | BLECF_ | 301
132 | BLECF_ | 201
343 | BLECF_ | 200
536 | BLECF_ | 304
134 | BLECF_ | 198
135 | BLECF_ | 197
569 | BLECF_ | 196
405 | BLECF_ | 195
270 | BLECF_ | 194
593 | #REF! | 176
548 | BLECF_ | 160
646 | BLECF_ | 191
579 | BLECF_ | 190
354 | BLECF_ | 189
363 | BLECF_ | 188
136 | BLECF_ | 187
520 | #REF! | 624
477 | BLECF_ | 185
481 | BLECF_ | 184
462 | BLECF_ | 183
452 | #REF! | 182
359 | BLECF_ | 181
137 | BLECF_ | 154
138 | BLECF_ | 179
139 | BLECF_ | 843
140 | BLECF_ | 177
141 | BLECF_ | 857
142 | BLECF_ | 174
143 | BLECF_ | 173
144 | BLECF_ | 172
145 | BLECF_ | #REF!
146 | BLECF_ | 830
147 | BLECF_ | #REF!
148 | BLECF_ | 549
149 | BLECF_ | 167
150 | BLECF_ | 166
151 | BLECF_ | 165
152 | BLECF_ | 164
153 | BLECF_ | 163
154 | BLECF_ | 162
155 | BLECF_ | 161
156 | BLECF_ | 283
157 | BLECF_ | 545
158 | BLECF_ | 158
518 | BLECF_ | 157
159 | BLECF_ | 156
160 | BLECF_ | 153
161 | BLECF_ | 152
162 | BLECF_ | 151
163 | BLECF_ | 150
607 | BLECF_ | 149
164 | BLECF_ | 148
165 | BLECF_ | 145
166 | BLECF_ | 144
392 | BLECF_ | 143
167 | BLECF_ | 142
168 | BLECF_ | 141
169 | BLECF_ | 140
170 | BLECF_ | #REF!
171 | BLECF_ | 138
172 | BLECF_ | 137
347 | BLECF_ | 98
173 | BLECF_ | 133
174 | BLECF_ | 132
175 | BLECF_ | 131
176 | BLECF_ | 130
177 | BLECF_ | 129
178 | BLECF_ | 127
179 | BLECF_ | 126
180 | BLECF_ | 813
577 | BLECF_ | 124
181 | BLECF_ | 123
182 | BLECF_ | 122
589 | BLECF_ | 72
183 | BLECF_ | 39
184 | BLECF_ | 716
645 | #REF! | 118
185 | BLECF_ | 117
186 | BLECF_ | 561
187 | BLECF_ | 115
188 | BLECF_ | 114
189 | BLECF_ | 113
190 | BLECF_ | 112
334 | BLECF_ | 111
191 | BLECF_ | #REF!
490 | BLECF_ | 109
192 | BLECF_ | 108
193 | BLECF_ | 107
194 | BLECF_ | 106
195 | BLECF_ | 105
196 | BLECF_ | 104
583 | BLECF_ | 535
197 | BLECF_ | 102
198 | BLECF_ | 101
199 | BLECF_ | 100
200 | BLECF_ | 99
201 | BLECF_ | 712
653 | #REF! | 96
202 | BLECF_ | 95
654 | BLECF_ | 94
203 | BLECF_ | 93
204 | BLECF_ | 92
205 | BLECF_ | 866
206 | BLECF_ | 88
207 | BLECF_ | 87
208 | BLECF_ | 86
295 | BLECF_ | 85
398 | BLECF_ | 55
545 | BLECF_ | 83
209 | BLECF_ | 480
676 | BLECF_ | 245
210 | BLECF_ | 80
211 | BLECF_ | 79
212 | BLECF_ | 78
213 | BLECF_ | 77
496 | BLECF_ | 76
214 | BLECF_ | 75
215 | BLECF_ | 74
216 | BLECF_ | 844
217 | BLECF_ | 71
218 | BLECF_ | 6
219 | BLECF_ | 69
220 | BLECF_ | 472
221 | BLECF_ | 67
222 | BLECF_ | 66
223 | BLECF_ | 65
224 | BLECF_ | 63
225 | BLECF_ | 61
420 | BLECF_ | 745
226 | BLECF_ | 59
227 | BLECF_ | 58
532 | BLECF_ | 57
228 | BLECF_ | 56
229 | BLECF_ | 54
230 | BLECF_ | #REF!
231 | BLECF_ | 64
232 | BLECF_ | 51
233 | BLECF_ | 49
571 | BLECF_ | 48
234 | BLECF_ | 47
235 | BLECF_ | #REF!
236 | BLECF_ | 45
237 | BLECF_ | 44
238 | BLECF_ | 43
562 | BLECF_ | 34
239 | BLECF_ | #REF!
547 | BLECF_ | 40
240 | BLECF_ | 23
495 | BLECF_ | 38
241 | BLECF_ | #REF!
242 | BLECF_ | 36
243 | BLECF_ | 35
244 | BLECF_ | 400
245 | BLECF_ | 33
674 | BLECF_ | 221
246 | BLECF_ | 30
247 | BLECF_ | 29
248 | BLECF_ | 242
249 | BLECF_ | 27
250 | BLECF_ | 26
251 | BLECF_ | 25
252 | BLECF_ | 24
408 | BLECF_ | 751
253 | BLECF_ | 22
254 | BLECF_ | 754
255 | BLECF_ | 20
256 | BLECF_ | 19
257 | BLECF_ | 353
258 | BLECF_ | 17
259 | BLECF_ | 15
576 | BLECF_ | 14
678 | BLECF_ | 13
260 | BLECF_ | 12
335 | BLECF_ | 11
527 | BLECF_ | 10
580 | BLECF_ | 9
526 | BLECF_ | 7
521 | BLECF_ | 5
588 | BLECF_ | 3
261 | BLECF_ | 2
657 | BLECF_ | 244
271 | BLECF_ | 89
336 | BLECF_ | 715
366 | BLECF_ | 706
367 | BLECF_ | 606
368 | BLECF_ | 675
369 | BLECF_ | 641
370 | BLECF_ | 719
371 | BLECF_ | 258
372 | BLECF_ | 652
373 | BLECF_ | 665
374 | BLECF_ | 740
375 | BLECF_ | 666
376 | BLECF_ | 722
377 | BLECF_ | 658
378 | BLECF_ | 626
379 | BLECF_ | 406
380 | BLECF_ | 4
381 | BLECF_ | 279
382 | BLECF_ | 97
383 | BLECF_ | 721
384 | BLECF_ | 128
385 | BLECF_ | 627
386 | BLECF_ | 689
528 | BLECF_ | 119