| LHC- - - 1999-09-22 |
| --- |
| MPS Commissioning Procedure |
| The commissioning of the LHC machine protection system MPS aspects of the Beam Loss Monitor system commissioning |
| Abstract This document describes the set of tests which will be carried-out to validate for operation the machine protection aspects of the LHC Beam Loss Monitor system. The area concerned by these tests extends over the whole LHC machine for each of the two LHC beams. These tests include the Hardware Commissioning, the machine check-out and the tests with beam. |
| Prepared by : Bernd Dehning Laurette Ponce Eva Barbara Holzer | Checked by : Gianluigi Arduini, Ralph Assmann, Roger Bailey, Brennan Goddard, Michel Jonker, Verena Kain, Mike Lamont, Stefano Redaelli, Benjamin Todd, Jan Uythoven, Jorg Wenninger, Markus Zerlauth | Approved by : Rüdiger Schmidt |
| History of Changes |
| Rev. No. | Date | Pages | Description of Changes |
| 0.1 0.2 | 2007-05-16 2008-09-02 | 12 | First draft for circulation. Second version…. |

Table of Contents

1.	Introduction	4

2.	Scope	4

3.	Purpose	4

4.	The layout	4

5.	Tests performed during the hardware commissioning	5

5.1	“BLM system tests”	5

6.	Link to other equipement	9

6.1	IntErfaceS with the Beam Interlock system	9

6.2	interface with the beam dumping system	10

7.	System tests during the machine checkout	10

7.1	Conditions required to perform tests	10

7.2	Description of the tests	10

7.3	Status of the system after the system tests	11

8.	Tests with beam.	11

8.1	Pilot of 5(109 p+ at 450 Gev	11

8.2	43 bunches of 4(1010 p+ at 450 Gev (below damage value)	11

# Introduction

The beam loss monitor system protects the equipment against damages by the energy deposition of the beam. It is the only system which allows to disable the beam permit in the loss duration range between 4 and 100 turns. For longer durations there is also the quench protection system able to initiate an abort trigger.

# Scope

This document covers the tests which will be carried out to validate the operation of the MPS relevant components of the Beam Loss Monitor (BLM) system and the relation to the machine protection system (MPS) for LHC beam 1 and 2. The area concerned by these tests is the whole LHC ring and the beam dump lines TD62 and TD68. The equipments concerned are the beam loss monitor electronics BLE (BLEB, BLECF, BLERX, BLETC, BLETX), and the monitors in the ring BLM (BLMEI, BLMES, BLMQI) and in the dump lines BLMD.

# Purpose

This document describes the procedures which will be applied for these tests and their sequence.

Each test in the test list has in front one of the following letters, defining at which interval or at which occasion the described test needs to be repeated (in the column labelled Repetition):

| N | Not to be repeated |
| --- | --- |
| S | To be repeated after every Shutdown |
| P | Periodical repetition required, like 1 x per month; details to be defined in text |
| O | To be repeated when LHC optics is changed |
| X | To be repeated when crossing scheme is changed |

This document is meant to be the reference document for the checklist which will be used during the commissioning of the MPS. Results of the tests will be documented in the MTF database.

# The layout

The BLM system is comprised of some 4000 monitors (3700 Ionisation Chambers and 280 Secondary Emission detectors) distributed around the machine. The principle of the system is to measure outside the cryostat or the vacuum chamber the secondary shower particles created by protons impacting on the vacuum chamber.  When the signal is over a pre-defined threshold, the BLM system triggers a beam dump to protect against damage of equipment and prevent quenches of magnet coils.

The derivation of the thresholds is based on a combination of several simulations (loss maps, secondary shower development, quench level, detector response) and verification measurements. These dependencies are schematically depicted in Figure 1.

Figure 1: Schematically view of the signal flow of the beam loss system including the beam permit signal.

To ensure the specified reliability level of the system, some of the tests presented here have to be done regularly (see G. Guaglio, PH-D thesis, CERN-LIB), namely before every fill.

# Tests performed during the hardware commissioning

This part describes the tests which are performed during the BLM related part of the machine protection system commissioning. Most of these tests verify the correct functioning of the “BLM System Tests” – they are tests of test procedures.

## “BLM system tests”

The “BLM System Tests” are described in detail in the EDMS document (LHC-BLM-TP-0001).  They test the correct implementation of the system and the system integrity. The frequency of these BLM tests is either A) before each start-up, B) before each fill or C) continuously during data taking.

The LHC beam loss monitors are foreseen to be hardware commissioned during the electronics installation phases. Independent tests of each monitor and each channel will be made by AB/BI/BLM to validate the connectivity and the topology of the system. All these tests are referenced in the MTF database and described in the EDMS document LHC-BLM-TP-0001.

For reference only: Some of the system test described in LHC-BLM-TP-0001

|  | Rep. | Action | Test condition | Group(s) Responsible |
| --- | --- | --- | --- | --- |
| 1 | N | See: LHC-BLM-TP-0001, 10.2.1. Tunnel Cables Test and Functional Test (MTF test 10-BLM Property6: LC(offset) w/cable&M-r): • Cable identification check (TS/EL sticker + AB/BI specific sticker) • Cable continuity check; particularly verify if there are no conductors swapped • Test of insulation, correct interconnection and current limits • The crate shall be equipped with a test card to record the measurements. • Measure the leakage current (offset) with the cable and the monitor connected. The measured current on all channels shall be below 0.5 pA. | 240 V | AB/BI |
| 2 | N | Test of insulation, correct interconnections and current limits (MTF test 10-BLM channel check (leakage, bias current) P7: Signal via source w/cable). Current source test: Check connection of the chamber up to the tunnel card | 240 V | AB/BI |

Tests which are performed before each fill are started by the ‘sequencer’ program or the SIS program. They verify that the results of the tests were “OK”. Then the programs inform the BLM system (combiner) of the test result. In case the result was “failed” the beam permit is taken away. The combiner card logs the test results in the logging database. The described functionality of the sequencer and the SIS still has to be implemented.

The BLM system itself (each combiner card) verifies that not too much time has passed in between these tests. After a time of T1 (T1=12h, currently implemented) passed, a flag goes in “warning”. This flag is updated at 1 Hz in the logging database. After a time T2 (T2=24h, currently implemented) passed the BLM system is not considered safe anymore and as soon as there is no beam, the beam permit is taken away.

### Conditions required to perform tests

Interconnection of cryostats performed in the sector

Installation of front-end mini-racks, crates, fan Unit supervised by AB/BI/BL.

Optical fibre terminations supervised by AB/BI/BL

BLM detector installation supervised by AB/BI/BL

‘daisy-chain’ high-voltage and signal cables supervised by AB/BI/BL

Installation of BLM electronics supervised by AB/BI/BL

### Conditions during the tests

The test conditions listed for each test separately in chapter 5.1.3 are defined as follows:

- 240 V power in tunnel (240 V)

- 240 V power in tunnel + optical fibre connection + Ethernet =  env I

- 240 V power in tunnel + optical fibre connection + Ethernet + closing the tunnel for source test = env II

- 240 V power in tunnel + optical fibre connection + Ethernet + connection to the beam interlock system = env III

- data base LSA + software infrastructure = env IV

### Description of the tests

The following tests are performed after a shutdown (S) in order to verify that the BLM test procedures are conform to specification. Any provoked error should be signalled by ‘Sequencer Program’.

|  | Rep. | Action | Test condition | Group(s) Responsible |
| --- | --- | --- | --- | --- |
| 1 | S | High Voltage Modulation test (MTF test BLM 40, checks the connectivity up to the surface card before every fill) Test the functionality of the T1 and T2 counting (warning created, beam permit taken away?) Test that the limits for “test OK” are correct; reduce and increase accepted min/max gain value locally (did test fail and take beam permit away?); test all channels of one crate at the same time; check that each monitor (!) goes to “failed” (check of firmware: is the comparison of the measured gain to the accepted gain working) | env I | AB/BI |
| 2 | S | 10pA Test (MTF 30-BLM Acquisition Chain Test via test signal, P2:10pA/84 sec run.sum) continuous test Lab test on one card with identical HW and firmware version (2 HW versions exist in tunnel), that the limits for “test OK” are correct; reduce and increase bias voltage below and above threshold (did test fail and take beam permit away?) |  | AB/BI |
| 3 | N | Double Optical Line comparison; Continuous check of the connectivity up to the surface card. Lab test on one card; force input through all the parameter space of the decision matrix and verify that the correct answer is given. |  | AB/BI |
| 4 | S | Acquisition Chain Test via test signal (MTF test 30-BLM, P1:100pA / 1.3 sec run.sum) Performed in sequence with the HV modulation test; The bias current is increased by 100 pA and the 1.3 second running sum is measured and checked. Bias current reduced to previous value.) Test the functionality of the T1 and T2 counting (warning created, beam permit taken away?) Lab-test on one card with identical firmware that the test can be failed (by artificially increasing the bias current) | env I | AB/BI AB/CO |
| 5 | S | For reference (see LHC-BLM-TP-0001 10.7.2): Radioactive Source Test; Functional test of full acquisition chain (MTF 50-BLM Acquisition Chain test via RA source): Create a signal on the ionisation chamber with the RA source Check the presence of signal in the corresponding DAB card | env II | AB/BI |
| 6 | S | Interlock test (MTF test 60-BLM) TC & CS vs DB comparison (MTF test 60-BLM Property1: Threshold/Ch Matrix Test). This comparison is done by “MCS check”. For each crate, test the functionality of the T1 and T2 counting (warning sent, beam permit taken away?) Lab test on one card, 16 channels, 12 integration time intervals, 32 energies (or on an random sample of the above). Automatic test procedure of the parameter space (thresholds, names, DCUM, etc., masking tables, connection tables). Change threshold value in BLETC (do not update data base) and observe the detection of the disagreement between BLETC value and data base (did test fail and take beam permit away?) For reference (see LHC-BLM-TP-0001 10.3.4): Remove Beam Permit (MTF test 60-BLM Property2: TTable User Beam Permit T) TC to CS transmission (Beam permit transmission from all threshold comparators to the last combiner card) (MTF test 60-BLM Property3: UBP Check BLETC to BLECS) For each crate, test the functionality of the T1 and T2 counting (warning set, beam permit taken away?) Observe visually (on the logging) the sequence (frequency ~ 0.5 Hz) of the test on one octant. For reference: CS to CIBU transmission – responsibility of BIS (Beam permit transmission from all the last combiner to the CIBU) (MTF test 60-BLM Property4: UBP Check BLECS to CIBU) Test functionality tested by BIS, EDMS 889281 | env IV env III | AB/BI |
| 7 | S | EMC test (MTF test 70-BLM) | env III | AB/BI |
| 8 | S | Beam Energy Reception test (MTF test 80-BLM) Unplug CISV input cable (observe error in toggle bit) and verify that energy is set to FFFF (max. energy) Unplug CISV output cable (observe lost packets) and verify that energy is set to FFFF (max. energy) | env III | AB/BI |

### Status of the system after tests

After these tests, the hardware is fully installed, cables are connected to the monitors and to electronics; fibre link and digital acquisition are operational. The logic of the interlocks has been checked and the system is ready to be tested with the beam interlock system.

# Link to other equipement

The interfaces listed in this paragraph concern only the ones in relation with the Machine Protection System; it does not describe any procedures to test the interfaces to protect individual equipment but only summarise the logic applied in the design.

## IntErfaceS with the Beam Interlock system

### SIGNALS between BLM SYSTEM and Beam INTERLOCK SYSTEM

The BLM system produces a beam permit from each monitor which are concentrated in 2 user beam permits (one for mask able and one for unmask able channels) and sent to 2 CIBU (mask able and unmask able) per octant.

The “Beam Info” signal given by the BIS is used to ensure that there is no beam before starting the periodic tests.

The beam energy is used to change the beam permit thresholds accordingly. The signal is received by the timing system via the CISV VME module (see test 5.1.3.8).

### Conditions and sequence for a beam dump

The beam user permit is changed to FALSE if:

- One of the 12 running sums for each channel is over the corresponding threshold.

- The comparison of the thresholds table loaded on the FPGA with the table stored in the database failed.

- Failure of one of the periodic or continuous tests (loss of communication) on any of the channels.

- Test mode of the BLMs.

Figure 2: Schematic view of beam permit, beam info and beam energy flow.

## interface with the beam dumping system

Direct interlocked monitors are the last protection in case of failure of the standard BLM system or of the BIS. There are two monitors per beam, one on the TCDQ and one on the TCSG. The threshold-signal comparison for the beam dump is done by comparing the voltage which is proportional to the losses with a reference voltage. The reference voltage adjustment can only be done with an access to the tunnel.

# System tests during the machine checkout

When the individual system tests, described in the previous sections, have been successfully completed, the integral system should be tested from the CERN Control Centre (CCC), simulating as much as possible future operations with beam.

## Conditions required to perform tests

Successful completion of the individual tests of BLM system sector by sector.

Logging system available (measurement and logging)

Setting data base (LSA) available

BIS operational

LBDS operational

Beam energy distribution operational

## Description of the tests

|  |  | Action | Group(s) Responsible |
| --- | --- | --- | --- |
| 1 | S | Test of transmission time of beam permit disabling Test 1 BLM per octant BLM takes away beam permit (by the toggle functionality of BLM system) Measure the delay – test result to go to MTF - (by time stamps) between the BLM initing and the dump system receiving the beam dump request | AB/CO AB/BT AB/BI |
| 2 | S | Test the change of the threshold value with the beam energy signal Ramp dipole in 4 sectors and observe logging of beam energy values | AB/BI |

## Status of the system after the system tests

After these tests, the complete acquisition system is commissioned for each octant.

# Tests with beam.

It is foreseen to make two different tests with beam, one to check the correct functioning of the beam permit disabling by depositing energy in the beam loss detectors and a second to check the correct setting of the threshold to prevent quenches. Below is the general description of the tests given, a more detailed description will be later appended to the document.

## Pilot of 5(109 p+ at 450 Gev

|  |  | Action | Group(s) Responsible |
| --- | --- | --- | --- |
| 1 | S | Test MPS functionnality of the BLM (without quenching): Decrease thresholds (trim aplication) to very low value Create local bump to force a beam dump request with very low level of losses Measure delay between signal over threshold and the beam dump : the whole chain from the chamber to the LBDS, via the LBIS, is tested | AB/CO AB/BI AB/BT AB/OP |
| 2 | S | Test interface of direct BLMs with the beam dumping system Reduce the voltage which sets the abort thresholds Dump beam on collimator TCDQ and TCSG of each ring Observe beam permit going away | AB/BI AB/BT |

## 43 bunches of 4(1010 p+ at 450 Gev (below damage value)

### Conditions required to perform tests

Stable and known optics : needed to determine the impact position and the lost proton density

Orbit corrected to 3 mm

BPM data logging available

### Conditions during the tests

Start with the pilot intensity and if not quenching, increasing step by step (to be defined) the intensity up to the quench level.

### DESCRIPTION of the tests

The test has to be done for one beam only.

|  |  | Action | Group(s) Responsible |
| --- | --- | --- | --- |
| 1 | N | Provoked quench (transient losses) Steer the beam in the chosen magnet Check level of losses and BLM readings at the quench level Repeat for MQM, MQTL, MQ and MB If no quench ocures the threshold DB values have to be reduced Test justification and described in Chamonix XV proceedings, 2006, Magnet Quenches with Beam (A. Koschik) | AB/CO AB/BI AB/OP |
| 2 | N | Provoked quench (steady state losses) Create a local bump |  |

# Text boxes

Beam_Info

LBDS

User_Permit

BIS

Beam Energy

From timing system

CISV

BLM